## REVISION NOTICE TABLE

| DATE | DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2025 | Updated channel names & formats in TOB files. GG |
| Feb. 2019 | Bottle spreadsheet converted to searchable BOT files. |

# PROCESSING NOTES

Cruise: 2012-03

Agency: PBS, Salmon and Freshwater Ecosystems, Nanaimo, BC

Location: Vancouver Island

Project: High Seas Salmon

Party Chief: Zubkowski T.

Platform: W.E. Ricker

Date: 1 March 2012 – 13 March 2012

Processed by: Germaine Gatien

Date of Processing: March 21, 2012 – 22 May, 2012

Number of original hex files: 	57		Number of CTD casts processed: 57

Number of original TSG files:     1		Number of TSG files processed: 1

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with an SBE43 Dissolved Oxygen sensor (#1119), an SBE pH sensor #0851 and a WetLabs ECO-FL Fluorometer #2215. The log indicates that the DO sensor was mounted on the primary pump. The salinometer used was a model 8400B Autosal, serial # 68572.

An SBE21 Thermosalinograph (serial #2488) sampled every 30s; there was no fluorometer mounted.

# SUMMARY OF QUALITY AND CONCERNS

The log books were in good order with an equipment list, a record of sampling done and notes about problems encountered.

There was one deep salinity calibration sample and many samples from ~10m. They were analyzed within 6 days of the end of the cruise. The comparison with the CTD data suggests that the secondary CTD salinity is likely slightly low, by about 0.002 to 0.004; the cruise that followed showed salinity low by from 0.003 to 0.007. Salinity was recalibrated by adding 0.004.

After its last factory service the SBE Dissolved Oxygen sensor was used on 3 cruises (including this one) with no calibration sampling. In April 2012 it was used on a Vector cruise that included calibration sampling, so the results from that cruise were used to calibrate these data.

It was discovered late in processing that the dissolved oxygen concentration had been calibrated incorrectly, so the data were re-derived with the right parameters and put through all the same steps as originally run. Because the temperature and salinity data had been edited and some records removed, the dissolved oxygen data were merged with the other channels at a stage after editing. This has no ill effects on any of the data, but the Dissolved Oxygen calibrations shown in the headers have incorrect values for Soc and Voffset; they should be 0.4442 and -0.4822, respectively. There is a note to this effect in the header comments of the CTD files.

The fit of ECO fluorescence versus extracted chlorophyll shows an offset of ~0.3ug/L. When the offset is subtracted from the readings, there is a reasonable fit, but we need more experience with these fluorometers over a wider range before attempting to recalibrate. We especially want to be able to relate the results to those with SeaPoint fluorometers because so much of that data has been collected. No recalibration was applied to these data, but that decision will be revisited if further studies lead to a reliable method.

The pH sensor malfunctioned; typically values rose throughout the cast; the channel has been removed from the data to be archived.

The thermosalinograph data compared well with data from CTD casts at the level of the intake. There was no intake temperature sensor, so a proxy for the intake temperature was derived by subtracting 0.065C° from the lab temperature, based on comparisons with the CTD data.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained. There were only a few minor problems noted.

Bottle data were provided in 2 spreadsheets, one with salinity, the other with a combination of extracted chlorophyll and nutrients. The dates of the most recent updates were added to the file names. Henceforth, an asterisk at the end of the file name refers to the date of the spreadsheet version.

The cruise summary sheet was completed.

The histories of the dissolved oxygen, conductivity and pressure sensors were obtained.

The configuration file did not change through the cruise. All entries were correct except the date of the fluorometer calibration. That was fixed and the file was then saved as 2012-03-ctd.xmlcon.

There was salinity, nutrient and CHL sampling at about 10m and pH sampling deeper in the casts.

##### Conversion of Raw Data

There was no DO sampling from this or any other previous cruise since this sensor was last recalibrated. The sensor was used for a few cruises in April that did include DO sampling, so processing was delayed to await those results; 2012-04 results were used to adjust the values for SOC and Voffset in the DO calibration parameters.

All data were converted using file 2012-03-ctd-new.xmlcon. (Note: The values of SOC and Voffset were later found to be wrong. They were corrected in file 2012-03-ctd-new2.xmlcon which was used to reprocess the Dissolved Oxygen concentration channel only.)

Plots were made for a few casts. The CTD stops at between 15 and 17db during the upcast. The Niskin bottle is usually mounted about 5m above the CTD, so data from 10 to 12db should be examined for the comparison with bottle samples. Checks were made of all the channels:

- The temperature and conductivity channels track well on the downcast but the upcasts show larger differences. These differences look as though they are due to alignment or pump problems because when the gradient reverses sign, so do the differences. There is also a significant vertical offset that is seen only in the upcasts. All sensors have some noisy features; it is not clear which sensor pair is better.

- Fluorescence looks noisy so it is hard to judge vertical offset between downcast and upcast..

- Dissolved Oxygen data look ok with the usual vertical offset between downcast and upcast.

- As usual for this project, the descent rate is noisy offshore and quiet in inlets, but it was kept fairly high most of the time.

- The pH signal looks bad with values rising through casts with a few exceptions. Near the surface it occasionally decreases slightly and then starts rising again, possibly due to shed wake corruption that is smoothed by the slow response rate.

##### pH tests

The pH data look bad, so a preliminary investigation was done to determine if the sensor needed to be sent to the factory.

First, the pH calibration parameters were checked and looked ok. Next, tests were done to see if the sensor might have been on a different external voltage than was entered in the configuration file, but there was no signal on any other voltage channels that looked like pH.

The pH rises throughout the cast with a few exceptions:

1. Near the surface on the way down it occasionally decreases slightly and then starts going up again. I suppose this could be due to shed wake corruption that is smoothed by the slow response.

2. At the end of cast #22 the CTD was taken to the surface and stayed there for more than a minute, then went down to 16db. The pH rose throughout the surface stop, but did go down as the CTD went down again, though it did not go down to the value it had at 16m on the full upcast.

3. During cast #67 it starts to decrease slightly below 170db below a local temperature maximum, and then increases throughout the upcast except at the end when the CTD was lowered a bit - it decreased somewhat then.

So it looks like the sensor is seeing some changes but they are mostly drowned out by the gradual rise.

The cap was not on – that was the first thing checked at sea when they noticed a problem. But could there have been some contamination of the sensor that gradually disappeared? The minimum values appear to rise through the cruise and I think the max does too, which might suggest again that something is coating the bit that actually does the measurement and maybe that coating was gradually washing off the sensor. However, it is just as likely that the range of pH is just different in the different regions. Seeing the pH samples should shed some light on that.

The pH calculation is temperature-dependent but there is no similar problem with the temperature traces.

The pH data will not be processed further.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### ALIGNCTD

Various advancements were tested to make DO and temperature traces have the same offset between downcast and upcast data. Since there are a few doubts about the upcast temperature, more weight than usual was put on matching distinctive features in the downcast, but what worked for the latter also provided the best offset between downcast and upcast traces. The best results were with an advance of +2s, which is unusually low. The membrane was replaced in March 2011, so settings from before then are not useful. During 2011-58 the primary temperature was often bad during upcasts, so results are very difficult to interpret.

ALIGNCTD was used to advance the DO signal relative to temperature by 2s.

##### CELLTM

Tests were run on 3 casts using a variety of settings and the best choice in each choice varied from feature to feature and cast to cast. Perhaps this is because the upcast data are very noisy and the temperature ranges small, so the corrections are so small. A setting of  (α=0.02, 1/β=7) looks ok for both channels and while not notably better than the other values tested, the results are better than applying no correction at all. The same results were found for 2011-58 and 2011-45

CELLTM was run using (0.02, 7) for the primary and the secondary conductivity for all casts. But tests were not run.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration; the tau correction was applied.

on cast #40 to look at differences between channel pairs. This was the only cast deeper than 250m.

##### Test Plots and Channel Check

There was only 1 cast deeper than 250m. The differences between the sensors were very noisy. The same sensors were used for 2011-58 (August) and 2011-45 (October) so data from two casts each from those cruises are included in the comparison to see if there are signficiant differences.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-58-0007 | 470 | +0.0001 | +0.00028 | +0.0030 | Mod, noisy |
| 2011-58-0013 | 470 | +0.0002 | +0.00028 | +0.0027 | High, noisy |
| 2011-45-0067 | 220 | +0.0005 | +0.00050 | +0.0050 | Mod, steady |
| 2011-45-0184 | 240 | +0.0008 | +0.00045 | +0.0045 | Mod, steady |
| 2012-03-0040 | 450 | +0.0007 | +0.00026 | +0.0023 | High, XNoisy |

The differences are similar to the earlier cruises.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. There are some negative pressure records, but they occur at the end of one cast when the CTD appears to have stayed right at the surface with the pumps running. Salinity drops quickly when pressure goes through 0, but there are low but “in-water” salinity values above that level. There are also some negative fluorescence values which will have to be investigated later.

A cross-reference list was produced and checked against log records. No errors were found.

The track plots were produced and look reasonable.

The surface report indicates that the average surface pressure was 2.2db which is shallow for Ricker work, but many of the casts were in protected inlets. The offshore casts start deeper a little deeper, with cast #40 starting at 3.6db.

##### SHIFT

Fluorescence

In previous uses of ECO Fluorometers it has been found necessary to advance the fluorescence by +48 records (2s) relative to pressure. A few casts were examined and the data are so noisy that it was difficult to test, so fluorescence was put through a filter for a few test casts. That suggests that at most a +12 record alignment is necessary. It is a little difficult to judge this because the upcast temperature is noisy. The two temperature channels are quite different, though the primary looks more like the downcast on a T-S surface, so the judgment was based on that channel. A few casts from cruise 2012-04 were examined since the same equipment was used (but in a different configuration because a rosette is present.) A small shift looks appropriate, ~12 records. The sensor was also used for two February cruises but the fluorescence signal was small and noisy and the upcast temperature noisy, so judgment was again difficult, though there did seem to be an offset.

SHIFT was run on the fluorescence channel to advance by 12 records.

Conductivity

Tests were run on both conductivity channels using a few casts with marginally stable features. For the primary conductivity the best setting was -0.5 records and for the secondary -0.7 records.

All casts were put through SHIFT using -0.5 records for the primary and -0.7s for the secondary conductivity.

Dissolved Oxygen

The SBE DO data have already been aligned, but a few casts were examined to see if further adjustments are required. While the vertical offset between upcast and downcast are slightly bigger than in the temperature traces, but it is a slight difference and matching features shows no difference in offsets. SHIFT will not be run on this channel.

pH

The pH data look bad with gradually increasing values, so alignment was not investigated.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: All warnings pertained to the upcast, so are of no concern for profiles.

##### COMPARISON WITH BOTTLES

The nutrient, chlorophyll and salinity bottle data were combined in spreadsheet 2012-03-bottles_plus_CTD_Data.xls.

Salinity

There was just one deep salinity bottle, but there were near-surface samples for most casts. The salinity samples were analyzed within 6 days of the end of the cruise.

REVERSE was run followed by DELETE and CLIP to 14db and Bin-Average with 0.5db bins. CTD data from 9, 9.5, 10, 10.5 and 11 db were exported to spreadsheets. The two salinity values for each level were combined in a spreadsheet with the bottle values. Differences were calculated and then averages, means and standard deviations in the differences. The standard deviations are lowest for the primary salinity at 10db, but there are many large differences. Data were sorted on the primary difference at 10db and data were excluded if the difference was >0.05. Using the reduced set of 43 bottles, the primary salinity at 10db has the lowest standard deviation and primary and secondary salinity were lower than bottles by an average of 0.007 and 0.003, respectively.

A plot of differences between bottles and 10db CTD data against a proxy for local gradient (CTD salinity at 11db – CTD salinity at 9db) shows that all the outliers are in high salinity gradients. The secondary salinity looks low by 0.004 at 9.5db and the standard deviation in the secondary is lowest at that level.

Earlier it was found that the two salinity channels differed by 0.0023 in a test cast, so the values of 0.007 and 0.003 are a little further apart than expected. The 9.5db difference of 0.004 for the secondary is more in line.

Next, the deep salinity sample from cast #40 was examined. The CTD dropped slowly at the bottom from 499db to 526db, so the level at which the bottle closed is unknown, but presumed to be fairly close to 5m above the deepest point. Differences were calculated based on salinity at 513, 517, 519 and 521db. The primary was found to be low by 0.0071, 0.0048, 0.0045, 0.0043 and the secondary by 0.0031, 0.0026, 0.0018 and 0.0018, respectively. The 513db result matches the shallow bottle differences best, but the others match the difference between primary and secondary better. The gradient near the bottom is small, so small errors in the depth estimate are not as significant as those at 10db.

The comparison is a rough one, given that the firing depth varies and the CTD often moved a lot during the stop. Moreover, for the surface samples we are at a level where there can be real changes in a short time and we are averaging over a half-metre in fairly high gradients. Most of these errors are not systematic, so it is likely that the primary is low by at least 0.004 and the secondary by at least 0.002.

The secondary salinity is closer to the bottles in most cases, and the level of noise is slightly less than in the primary, so that channel will be archived. Since the error is not large and the standard deviation in the comparison is large, this comparison alone does not justify recalibration.

Finally the data from the secondary channels from 10db were added to spreadsheet 2012-03-bottles_plus_CTD_data.xls.

Fluorescence

The spreadsheet with CTD and bottle data was used to do a comparison of the ECO fluorescence and extracted chlorophyll samples from 10db. The results showed a fairly good fit for values of CHL <2ug/L:

ECO Fluor = 1.06 * Extracted CHL +0.3255

The offset of 0.3255 was then subtracted from the ECO fluorescence and then the ratio of the adjusted fluorescence to extracted CHL was plotted against CHL. That ratio shows a pattern that has been seen on other cruises with this type of sensor, with the ratio starting very high for very low CHL, reducing steeply at first and then gradually approaching 1 as CHL rises. In this case there is only 1 bottle with CHL>1.5ug/L and for that the fluorometer reads too low.

| Extracted CHL | 0.03 ug/L | 0.1ug/L | 0.25 ug/L | 1.5 ug/L | 9.3 ug/L |
| --- | --- | --- | --- | --- | --- |
| Ratio adjusted FL/CHL | 13 | 4 | 2.5 | 1.2 | 0.8 |

When the same plot is made with (FL - 0.3255)/ CHL versus CHL, the results are much flatter, with values slightly high at the lowest CHl, but most ratios are between 0.5 and 1.5.

Further study of this sensor is required to see how the results compare with the SeaPoint fluorometers that have been used until recently. Not enough is known about this sensor to recalibrate.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

CTDEDIT was used to clean the secondary temperature and salinity data.

Most editing was light with removal of some surface records and records corrupted by shed wakes. Large spikes in salinity were smoothed where there were no spikes in temperature, and salinity was cleaned where small spikes appear to be caused by a mismatch of conductivity and temperature channels. All casts required some editing.

Special steps were needed for 2 casts:

  - Cast #1 is a problem because data acquisition did not come on until the CTD was at 33.7db. The 2nd drop to 15db makes using the upcast data complex. A special file was prepared by using a text editor to remove that 2nd drop from a reversed file and then running DELETE. This file was called 2012-03-9001 for the short-term. That file was put through DELETE. After editing it will be saved as 2012-03-0001.edt.

  - Cast #169 is another problem since the pumps were not turned on until the CTD was at 134.6db of the downcast, so the upcast was chosen for archiving. There was no 2nd drop to 20db, so no special steps were required. The REVDEL file was edited.

On-screen plots of descent rate and pump status were used to guide editing.

Note was made of the editing details in the headers of the relevant files.

The edited files were copied to *.EDT files.

##### Dissolved Oxygen Correction

At this stage it was discovered that the wrong parameters had been used for the derivation of oxygen concentration. Derive was run again with the correct parameters – these files were put in a separate folder to avoid confusion. The files were then converted to IOS Header format, put through CLEAN and DERIVE. For casts #1 and 169 steps were taken as described above before they were put through DELETE. The new files were put through CLEAN again to remove Sea-Bird headers so that they would not be repeated in the merged files.

MERGE was run with Pressure as the reference channel, choosing all channels except Oxygen:Dissolved:SBE  from the EDT files and only channel Oxygen:Dissolved:SBE  from the DEL3 files. These merged files were named EDZ.

##### Other Comparisons

Previous experience with these sensors –

- Both temperature and conductivity sensors have been used on 2 earlier cruises since their last factory service. There has been limited salinity calibration sampling and no recalibration was applied for either cruise. The cruise that followed this one, 2012-04, found both salinity channels to be low but there was scatter in the results and a lot of variability in the casts; an estimate was made that the secondary salinity was low by 0.005.

- The pressure sensor has been used 2 times since it was last recalibrated and no problem was found, though there is also little evidence that the pressure was ok. During 2012-04 the pressure was believed to be low by at least 1.2db.

- The dissolved oxygen sensor has been used for 2 previous cruises since its last factory service and for 2012-04 which followed this cruise. There was no calibration sampling for the earlier cruises, but there is sampling for 2012-04.

Historic ranges – Data were plotted with the 3-standard deviation plots of temperature and salinity ranges where local climatology was available. The only excursions from those ranges were some cases of salinity being low near the bottom for a few casts in inlets or close to shore. Similar results were found for the Strait of Georgia in February when different equipment was used, so this is likely a real variation. The 3-standard deviation range is too severe a criterion this close to shore.

##### Initial Recalibration

The bottle comparison suggests that the secondary salinity is low by from 0.002 to 0.004, but the standard deviation in the comparison does not lend great confidence in the comparison. During 2012-04 which followed this cruise the salinity was also found to be low. There were some problems with the comparison, but it was felt appropriate to add 0.005 to the secondary salinity. For this cruise the correction may be a little lower as the was likely drifting with time. The salinity will be adjusted by adding 0.004. Pressure is thought to be low due to results of 2012-04.

File 2012-03-recal1.ccf was prepared to add 1.2db to the pressure and 0.004 to the secondary salinity.

.

The dissolved oxygen data were recalibrated upon conversion using 2012-04 results. Fluorescence data will not be recalibrated at this time, but may be revisited after more research is done on this type of sensor.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 were clipped to 150db and processed in 2 ways for Dr. Peña, with a filter and without a filter, followed by bin averaging in both cases. The SHFC1 files were put through reverse. But the usual approach of clipping files to 17db does not work for these files because in most cases (but not all) the CTD came to the surface and then went down again. The files were recalibrated, but not clipped. If clipped files are needed, it will be necessary to use a graphical editor or text editor to extract the section that corresponds to the bottle firing.

The EDT files were put through a median filter, size 11, applied to the fluorescence channel only.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the filtered files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and look reasonable; there are a few unstable features, but they are in areas where they may be real and no instrumental cause was found.

Profile plots were made to check all channels and no problems were found.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE,  SBE:pH, Status:Pump, Descent_Rate and Flag.

A second CTD dissolved oxygen channel was derived with units umol/kg.

Dissolved Oxygen saturation was calculated and surface values were plotted. Most values were between 95% and 105%, with a few higher and lower values in inlets. The offshore values were between 95% and 105%. Values between 85% and 95% were found in Juan de Fuca Strait and Johnstone Strait and Queen Charlotte Sound and inlets. There were single casts with surface saturation of about 75% and 115% in inlets. These results suggest that the dissolved oxygen calibration is reasonable.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence data are nominal and unedited, except that some

records were removed in editing temperature and salinity.

There was no calibration sampling for dissolved oxygen, but the

results of cruise 2012-04 were used to derive DO concentration.

Recalibration of CTD salinity was based on many 10m bottles and 1

deep bottle from this cruise and from sampling during 2012-04. Both

comparisons were limited by great variability, but salinity data are

likely good to +/-0.003.

The SBE pH sensor malfunctioned throughout the cruise.

It was discovered late in processing that the dissolved oxygen concentration

had been calibrated incorrectly, so the data were re-derived with the

right parameters and put through all the same steps as originally run.

Because the temperature and salinity data had already been edited, the

corrected dissolved oxygen data were merged with the other channels at a

stage after editing. This has no ill effects on any of the data, but the

Dissolved Oxygen calibrations shown in the headers have incorrect values

for SOC and VOffset; they should be 0.4442 and -0.4822, respectively.

For details on the processing see processing report: 2012-03-proc.doc..

##### Producing final files

The Standards Check routine was run and no problems were found. The final files were named CTD.

A cross-reference listing was produced for CTD files.

HEADER CHECK was run and no problems were found.

A cruise track was plotted and no errors found.

The sensor history was updated for the CTD sensors.

##### Thermosalinograph

Data were provided in 1 hex file.

No loop samples are available and all bottle sampling was too deep to be compared with the TSG.

a.) Checking calibrations

The calibrations were checked and the only problem concerns the serial number of the conductivity sensor; this was changed to 2488 in the hdr, hex and calibration files.

After that change the CON file was saved as 2012-03-tsg.con. There was no intake temperature sensor.

b.) The file was converted to a CNV file using the configuration file mentioned above. It was then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Time-series plots were produced. There is a lot of variability in the salinity at times, but given that the ship went in and out of inlets, this is likely not noise.

c.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2012-03-ctd-tsg-comp.xls.

CTD data were also thinned to points within 0.5db of 3db and those were also added to the comparison spreadsheet. The intake level is believed to be closer to 3m than 4m, but there are often no useful data near 3m.

The ATC file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for salinity and temperature and the files were reduced to the times of CTD files (if CTD data was available from 4m). There were 56 matches.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude were all <0.00044° and in longitude <0.00025°. The median differences were both 0.0001°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature and salinity.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity and temperature channels for the CTD at 4m and the TSG. A plot of differences against standard deviations shows that for both variables the largest outliers are associated with high TSG variability. Since we are not matching the time perfectly and not allowing for time in the loop, this is expected. When outliers are excluded based on high standard deviation, the TSG lab temperature is found to be higher than the CTD by 0.065Cº if the 40 cases standard deviations <0.1 are included and by 0.064 Cº if the 43 cases with differences <0.1 are included. The standard deviations were <0.03Cº in both comparisons.

The TSG salinity is very noisy at times. When the 42 cases were removed for which standard deviations in the TSG salinity were >0.1; the median indicated that the TSG was lower than the CTD by 0.0029, but the standard deviation was 0.1456. When data were removed based on differences between TSG and CTD being >0.02 there remained 31 points of comparison and the TSG was found to be higher than the CTD by a median value of 0.0044 with a standard deviation of 0.0065. That looks more reliable. This comparison was run using the CTD data before recalibration. Using recalibrated salinity would bring the CTD and TSG into very close agreement.

Next, comparisons were made between the 3m data and the TSG. Using the same approach as for 4m, but with less data, the salinity was found to be high by a median of 0.0033 and standard deviation of 0.0075 and the temperature was found to be high by 0.0647 and standard deviation of 0.008. These results are close to those from 4m, but look better because the average and median are closer and the temperature standard deviation is significantly lower. The standard deviation in salinity is slightly higher. (This comparison was done before pressure was corrected, so the CTD data said to be from 3db and 4db are likely from ~4db and ~5db.)

- Calibration History

The TSG primary temperature and conductivity were recalibrated in March 2011 and there is no record of any other uses since that time.

Conclusions

1. The TSG clock appears to have worked well.

2. There was no flow meter, but it is clear that there was no flow at the beginning of the record, so that data will be removed. Through most of that period the ship was not moving.

3. The temperature in the lab is higher than the CTD by about 0.065Cº.

4. There was no TSG intake thermistor, so the lab temperature minus 0.065Cº should be used to derive a proxy for the intake temperature.

5. The TSG Salinity is very noisy as expected for a cruise with many excursions into inlets. The comparison with CTD casts shows the salinity to be high by approximately 0.004, which is very close especially given the fact that the CTD salinity is likely low by about the same amount. No recalibration will be applied.

6. The Position:New channel had zero values until late in the file even though the ship was clearly moving.

7. The 3m matches are better than those at 4m, as expected.

8. After the TSG processing was done, it was discovered that the CTD pressure was likely low by at least 1db, so the comparisons really from ~4m and 5m. So it is particularly unsurprising that the shallower results looked better.

f.) Editing

The ATC files were opened in CTDEDIT.

The T and S data were removed where there was clearly no flow. Position and time information were left so that a cruise plot is available.

Other data were examined but no further editing was applied.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

ADD CHANNEL was used to add channel Temperature:Lab after Temperature:Primary and that new channel was filled with values from Temperature:Primary.

File 2012-03-tsg-recal1.ccf was prepared to subtract 0.065 from the Temperature:Primary. A few values were checked to ensure it was applied correctly and it was.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary and Flag. Position:New was removed since there were very few 0 values, so it is not very useful.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

Particulars: (notes from log book)

40. deep salinity sample

64. anoxic

67. anoxic

154. Problem with sounder – did not get to bottom.

169. Pump not on until 135m on downcast

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 13Apr11 | Factory |  |  |
| Conductivity | 3394 | 29Mar11 | Factory |  |  |
| Secondary Temp. | 4883 | 31Mar11 | Factory |  |  |
| Secondary Cond. | 1763 | 29Mar11 | Factory |  |  |
| pH | 0851 | 21Apr11 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 29Mar2011 | Factory |  |  |
| WetLabs Eco Fluorometer | 2215 |  | IOS |  |  |
| Pressure Sensor | 69698 | 15Apr2011 | Factory |  |  |