## Notes
Analysis: | Hugh Maclean, 27-APRIL-12 using Water Properties' Guildline 8400B Salinometer Serial Number 68572. Data processing by Kenny Scozzafava
Chief Scientist: | Peter Chandler
2012-05-16 |  | KS
No duplicate samples were collected therefore no precision analysis could be conducted.
No comments as to how the analysis went so I am assuming there were no problems.
#77 was really #66. Sample #77 is from cast #20 at station ADCP and is not one intended to be sampled for salinity. It stood out as an outlier but looks fine if reassigned to #66.
The salinity analysis log confirms that it was just an inputting error by the analyst. Should have been station 69 - #66.
2012-06-12 |  | KS
Corrected Sample #56 event number. Was 13 but is really 14.
2015-03-30 |  | KS
Amended header in the QF worksheet to show that silicone inserts were used in place of the typical nylon inserts supplied by Guildline Instruments.

## QF 2012-04 SAL FINAL
Header: | Salinity samples were collected in 200 mL type ll glass bottles with reusable silicone inserts supplied by Xiamen Cuten Accessories Industry, Ltd. And
screw caps supplied by Ocean Scientific International Limited. They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer
standardized with IAPSO standard seawater 19 days after collection. For details including a duplicate analysis, see document QF2012-04sal.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Flag:Salinity Bottle | Salinity:Bottle | Comments
2012-04 | 2 | SI | 2 |  | 30.3023
2012-04 | 4 | 59 | 11 |  | 30.7682
2012-04 | 8 | 102 | 25 |  | 33.6245
2012-04 | 11 | 75 | 39 |  | 33.4884
2012-04 | 14 | 72 | 56 |  | 31.6839
2012-04 | 17 | 69 | 66 | 2 | 32.6994 | SAL: Mislabeled as #77 but corrected
2012-04 | 20 | ADCP | 78 |  | 31.4664
2012-04 | 24 | 65 | 88 |  | 30.6103
2012-04 | 27 | 62 | 99 |  | 30.5135
2012-04 | 32 | 56 | 109 |  | 30.2276
2012-04 | 42 | 46 | 123 |  | 30.1659
2012-04 | 46 | 42 | 135 |  | 30.7775
2012-04 | 51 | 39 | 151 |  | 30.9052
2012-04 | 54 | 27 | 168 |  | 30.8667
2012-04 | 58 | 2 | 184 |  | 30.8575
2012-04 | 61 | 3 | 200 |  | 30.9008
2012-04 | 63 | 6 | 216 |  | 30.5559
2012-04 | 67 | 9 | 229 |  | 30.548
2012-04 | 70 | 12 | 241 |  | 30.6154
2012-04 | 72 | 14 | 257 |  | 30.5976
2012-04 | 75 | 16 | 272 |  | 30.4599

## Salinity Data Processing
Sample ID | Event | Station | Bottle Label | DateTime | Bath T | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Flag | New Calculated Salinity | Comments
2 | 2 | SI | 2012-04#11 | 2012-04-27 10:59 | 24C | 0.87929 | 3.6109e-06 | 4.3988e-05 | 0.87934 | 30.3023 | 0.00014015 |  | #VALUE!
11 | 4 | 59 | 2012-04#12 | 2012-04-27 11:02 | 24C | 0.89138 | 3.6104e-06 | 4.4107e-05 | 0.89142 | 30.7682 | 0.00014187 |  | #VALUE!
25 | 8 | 102 | 2012-04#13 | 2012-04-27 11:04 | 24C | 0.9649 | 4.465e-06 | 4.3988e-05 | 0.96494 | 33.6245 | 0.00017125 |  | #VALUE!
39 | 11 | 75 | 2012-04#14 | 2012-04-27 11:07 | 24C | 0.96142 | 3.1534e-06 | 4.4107e-05 | 0.96146 | 33.4884 | 0.00012316 |  | #VALUE!
56 | 14 | 72 | 2012-04#15 | 2012-04-27 11:09 | 24C | 0.91506 | 2.8768e-06 | 4.4107e-05 | 0.9151 | 31.6839 | 0.00011178 |  | #VALUE!
66 | 17 | 69 | 2012-04#16 | 2012-04-27 11:12 | 24C | 0.9412 | 2.2405e-06 | 4.4107e-05 | 0.94124 | 32.6994 | 8.7866e-05 |  | #VALUE!
78 | 20 | ADCP | 2012-04#17 | 2012-04-27 11:14 | 24C | 0.90944 | 5.3054e-06 | 4.4107e-05 | 0.90949 | 31.4664 | 0.00020554 |  | #VALUE!
88 | 24 | 65 | 2012-04#18 | 2012-04-27 11:16 | 24C | 0.88729 | 2.2391e-06 | 4.4048e-05 | 0.88733 | 30.6103 | 8.9564e-05 |  | #VALUE!
99 | 27 | 62 | 2012-04#19 | 2012-04-27 11:19 | 24C | 0.88477 | 2.4776e-06 | 4.4048e-05 | 0.88482 | 30.5135 | 9.4459e-05 |  | #VALUE!
109 | 32 | 56 | 2012-04#20 | 2012-04-27 11:21 | 24C | 0.87735 | 3.4516e-06 | 4.3988e-05 | 0.8774 | 30.2276 | 0.00013158 |  | #VALUE!
123 | 42 | 46 | 2012-04#21 | 2012-04-27 11:24 | 24C | 0.87575 | 2.8897e-06 | 4.3988e-05 | 0.87579 | 30.1659 | 0.00011145 |  | #VALUE!
135 | 46 | 42 | 2012-04#22 | 2012-04-27 11:26 | 24C | 0.89162 | 2.4776e-06 | 4.3988e-05 | 0.89166 | 30.7775 | 9.7507e-05 |  | #VALUE!
151 | 51 | 39 | 2012-04#23 | 2012-04-27 11:28 | 24C | 0.89493 | 2.4912e-06 | 4.3988e-05 | 0.89497 | 30.9052 | 9.4885e-05 |  | #VALUE!
168 | 54 | 27 | 2012-04#24 | 2012-04-27 11:30 | 24C | 0.89393 | 3.4658e-06 | 4.4107e-05 | 0.89397 | 30.8666 | 0.00013637 |  | #VALUE!
184 | 58 | 2 | 2012-04#25 | 2012-04-27 11:33 | 24C | 0.89369 | 3.6104e-06 | 4.3988e-05 | 0.89374 | 30.8575 | 0.00013905 |  | #VALUE!
200 | 61 | 3 | 2012-04#26 | 2012-04-27 11:43 | 24C | 0.89481 | 2.1649e-06 | 4.4048e-05 | 0.89486 | 30.9008 | 8.4203e-05 |  | #VALUE!
216 | 63 | 6 | 2012-04#27 | 2012-04-27 11:46 | 24C | 0.88587 | 2.7426e-06 | 4.3988e-05 | 0.88592 | 30.5559 | 0.00010971 |  | #VALUE!
229 | 67 | 9 | 2012-04#28 | 2012-04-27 11:48 | 24C | 0.88567 | 4.4783e-06 | 4.4048e-05 | 0.88571 | 30.548 | 0.00017259 |  | #VALUE!
241 | 70 | 12 | 2012-04#29 | 2012-04-27 11:51 | 24C | 0.88742 | 3.974e-06 | 4.4048e-05 | 0.88746 | 30.6154 | 0.00015323 |  | #VALUE!
257 | 72 | 14 | 2012-04#30 | 2012-04-27 11:53 | 24C | 0.88696 | 4.1786e-06 | 4.4107e-05 | 0.887 | 30.5976 | 0.00016326 |  | #VALUE!
272 | 75 | 16 | 2012-04#31 | 2012-04-27 11:55 | 24C | 0.88338 | 9.1249e-06 | 4.3988e-05 | 0.88343 | 30.4599 | 0.00035138 |  | #VALUE!

## RAW Salinometer Output
Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 | K15 | 2012-04-27 09:25 | 24C | 0.9998 | 2.7425e-06 | 4.3094e-05 | 0.99984
Calibration#2 | K15 | 2012-04-27 09:26 | 24C | 0.9998 | 2.2391e-06 | 4.4107e-05 | 0.99984
2012-04#11 | SI | 2012-04-27 10:59 | 24C | 0.87929 | 3.6109e-06 | 4.3988e-05 | 0.87934 | 30.3023 | 0.00014015 |  | 2
2012-04#12 | 59 | 2012-04-27 11:02 | 24C | 0.89138 | 3.6104e-06 | 4.4107e-05 | 0.89142 | 30.7682 | 0.00014187 |  | 11
2012-04#13 | 102 | 2012-04-27 11:04 | 24C | 0.9649 | 4.465e-06 | 4.3988e-05 | 0.96494 | 33.6245 | 0.00017125 |  | 25
2012-04#14 | 75 | 2012-04-27 11:07 | 24C | 0.96142 | 3.1534e-06 | 4.4107e-05 | 0.96146 | 33.4884 | 0.00012316 |  | 39
2012-04#15 | 72 | 2012-04-27 11:09 | 24C | 0.91506 | 2.8768e-06 | 4.4107e-05 | 0.9151 | 31.6839 | 0.00011178 |  | 56
2012-04#16 | 69 | 2012-04-27 11:12 | 24C | 0.9412 | 2.2405e-06 | 4.4107e-05 | 0.94124 | 32.6994 | 8.7866e-05 |  | 77
2012-04#17 | ADCP | 2012-04-27 11:14 | 24C | 0.90944 | 5.3054e-06 | 4.4107e-05 | 0.90949 | 31.4664 | 0.00020554 |  | 78
2012-04#18 | 65 | 2012-04-27 11:16 | 24C | 0.88729 | 2.2391e-06 | 4.4048e-05 | 0.88733 | 30.6103 | 8.9564e-05 |  | 88
2012-04#19 | 62 | 2012-04-27 11:19 | 24C | 0.88477 | 2.4776e-06 | 4.4048e-05 | 0.88482 | 30.5135 | 9.4459e-05 |  | 99
2012-04#20 | 56 | 2012-04-27 11:21 | 24C | 0.87735 | 3.4516e-06 | 4.3988e-05 | 0.8774 | 30.2276 | 0.00013158 |  | 109
2012-04#21 | 46 | 2012-04-27 11:24 | 24C | 0.87575 | 2.8897e-06 | 4.3988e-05 | 0.87579 | 30.1659 | 0.00011145 |  | 123
2012-04#22 | 42 | 2012-04-27 11:26 | 24C | 0.89162 | 2.4776e-06 | 4.3988e-05 | 0.89166 | 30.7775 | 9.7507e-05 |  | 135
2012-04#23 | 39 | 2012-04-27 11:28 | 24C | 0.89493 | 2.4912e-06 | 4.3988e-05 | 0.89497 | 30.9052 | 9.4885e-05 |  | 151
2012-04#24 | 27 | 2012-04-27 11:30 | 24C | 0.89393 | 3.4658e-06 | 4.4107e-05 | 0.89397 | 30.8666 | 0.00013637 |  | 168
2012-04#25 | 2 | 2012-04-27 11:33 | 24C | 0.89369 | 3.6104e-06 | 4.3988e-05 | 0.89374 | 30.8575 | 0.00013905 |  | 184
2012-04#26 | 3 | 2012-04-27 11:43 | 24C | 0.89481 | 2.1649e-06 | 4.4048e-05 | 0.89486 | 30.9008 | 8.4203e-05 |  | 200
2012-04#27 | 6 | 2012-04-27 11:46 | 24C | 0.88587 | 2.7426e-06 | 4.3988e-05 | 0.88592 | 30.5559 | 0.00010971 |  | 216
2012-04#28 | 9 | 2012-04-27 11:48 | 24C | 0.88567 | 4.4783e-06 | 4.4048e-05 | 0.88571 | 30.548 | 0.00017259 |  | 229
2012-04#29 | 12 | 2012-04-27 11:51 | 24C | 0.88742 | 3.974e-06 | 4.4048e-05 | 0.88746 | 30.6154 | 0.00015323 |  | 241
2012-04#30 | 14 | 2012-04-27 11:53 | 24C | 0.88696 | 4.1786e-06 | 4.4107e-05 | 0.887 | 30.5976 | 0.00016326 |  | 257
2012-04#31 | 16 | 2012-04-27 11:55 | 24C | 0.88338 | 9.1249e-06 | 4.3988e-05 | 0.88343 | 30.4599 | 0.00035138 |  | 272