## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 May 2018 | Added DIC and Alkalinity data to 8 CHE casts. For details see document Carbon_Data_Addition.docx. |
| 31 March 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2012-05

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG/JdF Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: June 13, 2012 – June 18, 2012

Processed by: Germaine Gatien

Date of Processing: 5 September 2012 – 20 November 2012

Number of original HEX files: 75

Number of CTD files: 75

Number of bottle casts:	21

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1176), a SeaPoint Fluorometer (#2356) with a 3X cable, a Wet Labs Eco-AFL/FL Fluorometer (#2215), a Biospherical QSP-400 PAR sensor (#4601), a surface PAR (#16504), an SBE18 pH sensor (#0692) and an altimeter (no serial # available).

The deck unit was a Seabird model 11, serial #0425. The logging computer was #2.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, but the altimeter serial number was missing from the log and from the configuration file and there was no mention of the surface PAR. Since the SPAR is often in the equipment list but not actually mounted, it is helpful to have it listed in the log. The CTD and rosette logs were in good order with useful notes about problems encountered.

The surface PAR data look odd for many casts. Some of the early and later casts look reasonable and for night casts the lack of a signal is not significant. Where there is no signal, the “null” value varies. The data have been left in the files as some may be useful. A warning was placed in the header comments.

Two fluorometers were used for this cruise enabling further investigation of how they compare. The results look quite different from other cruises that used the same sensor types. The ECO and Seapoint were closer than usual and generally lower than extracted CHL. In the past the ECO has been found higher than the SeaPoint and both were higher than extracted CHL when CHL values are <5ug/L. No obvious explanation for this emerged. There does not appear to be a simple relationship between these sensors and extracted CHL.

The comparison between CTD and bottle salinity was noisier than usual, despite the fact that the samples had been analyzed shortly after collection. This was likely due to problems with the salinometer. The analyst commented on stability problems and technicians later found heavy algal growth in the cell and temperature bath. After these data were analyzed the salinometer was cleaned and appears to have performed well since then.

A change in the recommendation for how dissolved oxygen should be calibrated occurred while this processing job was ongoing. The bottle files were redone using the new technique, but since some editing had already occurred of profile files, corrections were made later by re-processing the dissolved oxygen channel and merging it with the edited files. For this reason the information in the header is incorrect as far as the Voffset and Slope parameters are concerned.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.4 mL/L from 0 to 50db

±0.15 mL/L from 50 to 150db

±0.10 mL/L below 150db (below 150db SBE DO values tend to be low)

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen and salinity. No CTD equipment problems were noted but there were some problems with bottles.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. All had been recalibrated in March or April 2011.

The calibration constants were checked for all instruments and no errors were found.

##### BOTTLE FILE PREPARATION

The ROS files were recreated using file 2012-05-ctd.xmlcon. They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. All data looked fine.

A preliminary header check turned up no problems.

The addsamp.csv file prepared in the DO calibration step was sorted on Event_Number and Sample_Number and then converted to CST files. The CST files will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2012-05-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2012-05chl.xls. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2012-05chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2012-05oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2012-05oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2012-05SAL.xls. The analysis was done within 4 to 10 days of collection. The files were simplified and saved as 2012-05sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2012-05nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers and the file was saved as 2012-05-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, ADD and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

Salinity

Compare was run with pressure as reference channel. Fits were made by excluding bottles from above 100db and those for which the difference between bottles and CTD >0.02. The CTD primary salinity was found to be low by an average of 0.0013 and the secondary high by 0.0004. The standard deviations in the fits were 0.005 for both sensor pairs, which is higher than usual. There was very slightly more pressure dependence in the secondary than in the primary. There is no evidence of significant salinity dependence in the results.

COMPARE was also run using Niskin Bottle # as the reference channel since there are some odd features to the main comparison. There is more noise at higher Niskin #, as expected, since those bottles are only used near the surface. There is no evidence of a bad bottle. (See 2012-05-Niskin.xls)

There is some time dependence but it is the same for the two. This may be due to variations in different regions, sampling in larger gradients at bottle levels in some areas and not in others, but there is also some concern about the analysis.

The analyst’s comments on erratic salinometer readings were examined in detail to determine which bottles might have been affected by this. Two of them intersect with bottles already noted as significant outliers with no associated noisy CTD data. Many are outliers but often the CTD data also looks noisy. Both sensor pairs have produced good data on average, but there are more outliers than usual. This is not likely to be due to evaporation since the samples were analyzed very close to collection. The salinometer problems are likely the problem.

No recalibration is required for either of these sensor pairs.

Outliers were identified and examined for evidence that the CTD may be responsible for the large differences, as shown by either high standard deviation in the CTD data or by large difference seen in one sensor pair but not the other. For the following cases it looks like the bottle is the problem and flag 3 was suggested with the comment “Significant outlier in COMPARE not explained by noise in CTD signal”:

Cast #1, 100db, sample #1, Niskin 1

Cast #13, 100db, sample #44, Niskin 7

Cast #23, 100db, sample #71, Niskin 8

Cast #23, 75db, sample #72, Niskin 9

Cast #25, 150db, sample #81, Niskin 2

Cast #51, 151db, sample #174, Niskin 6

Cast #51, 50db, sample #178, Niskin 10

(Kenny Scozzafava agreed to these flag and comment changes.)

For full details for the COMPARE run see file 2012-05-sal-comp1.xls and 2012-05-comp-Nisk.xls.

In November 2012 the salinometer was taken apart for maintenance and cleaning.  The technician reported that there was pretty heavy algae growth within the cell and quite a bit inside the temperature bath as well.  After cleaning it thoroughly and allowing it to reacclimatize overnight, it seemed to behave beautifully, holding the calibration very well over the analysis period.

This information supports the idea that the poor salinity comparison was largely due to salinometer problems.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. Following recent recommendations from SeaBird recalibration will be based on a fit of differences between bottles and CTD DO with the offset forced to equal 0. SeaBird warn that this will not fit well for values <2mL/L due to limitations in Winkler titrations at low values.

Outliers were excluded based on residuals. Most of the outliers are from near the surface. In the course of checking outliers against flagged samples, it was discovered that 1 dissolved oxygen cast had been misidentified in the DO spreadsheet – samples #226-238 are from event #70, not #68. Cast #68 was not a bottle cast, so the data were missed in the merge process; the merges were rerun and the comparison done again.

Using the fit of differences against bottle DO, and excluding outliers based on fit residuals, the following fit was found:

Bottle DO = 1.0704 * CTD DO

Removing more outliers had little effect on the slope, and this fit includes the full range of DO values. Using only CTD DO values >2mL/L had little effect on the fit.

The samples flagged by the analyst were examined to see if COMPARE suggested a need for change of flag, or if something should be added to the comments. In most cases COMPARE showed no problem, but there is sufficient noise in the fits that this does not mean flags should be removed. The following amendments were suggested and agreed to by Kenny Scozzafava who did the post-analysis checks:

Cast #1, sample #1 – major outlier in COMPARE, already flagged 4. Bottle low by 0.1mL/L in        COMPARE – add to comment.

Cast #1, sample #2 – major outlier in COMPARE, already flagged 2– change flag from 2 to 4 since spigot was open and it is high in COMPARE by 0.2m/L. Add to comment.

Cast #1, sample #7 – flagged 2 and is slightly high in COMPARE, flag 3.

Cast #41 in general – since values not trusted, change from flag 2 to flag 3. However, they were not outliers except for 148 and 150, but flag 3 is suitable for those as well...

Cast #41 – sample #148 was an outlier, but the CTD was noisy – change flag to 3.

Cast #41 - sample #150 was an outlier, high by about 0.1, but near surface that is not unusual. Flag 3.

Cast #73, sample #243 – outlier, off by 0.1, change flag from 3 to 4.

The following samples that were already flagged were slight outliers, but not by so much as to be rejected in the fit.

Cast #27, sample 102

Cast #37, sample 135

Cast #46, sample #168

Cast #64, sample #203

The outliers in COMPARE that were not already were flagged were then examined. Most were in the top 10db where the fits against CTD DO are often poor because of high gradients. The following is the only bottle that looks off by enough to justify adding a flag:

Cast #23, sample #74 – off by ~0.2mL/L - CTD is noisy but probably not by enough to cause this large an error –flag 3.

The SAM files were recalibrated using the fit found above and COMPARE was rerun. The results show values slightly low at the lower end of the range and slightly high at the higher end. When forced through the origin the fit was excellent. If that forcing is not done, the values look slightly low at the low end of the range and slightly high near the surface. Overall the recalibration worked well.

For full details of the comparisons see files 2012-05-dox-comp1.xls and 2012-05-dox.comp2.xls.

Fluorescence

COMPARE was run using the Wet Labs ECO CTD Fluorescence and the Extracted Chlorophyll from bottles. There were only 2 extracted CHL samples >4ug/L. For those 2 the ECO was lower than the CHL but came close and the SeaPoint was much lower. At lower CHL values both sensors were closer to each other and to the CHL than we have seen from other tests run. The SeaPoint was closest at the very lowest values of CHL. The sampling was shallower than for other cruises on which such tests have previously been run, with bottles at ~2m, 5m and 20m where we usually have 5m, 10m and 20m. This is a different SeaPoint sensor than in the other tests, but the SeaPoint behaves much as it usually does. The ECO is the same sensor as tested before but in previous tests it was found to read much higher at low CHL than is seen here. No notable outliers were found.

See 2012-05-chl-fluor-comp1.xls for the full COMPARE results.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

At this point the data from the MRGCLN2 files were exported to a spreadsheet for comparison with the rosette files to ensure no data were misplaced or missing. A few problems were discovered and fixed by rerunning the MERGE process. Data were exported to a spreadsheet again and the results look ok, though there are some bottles that are not reported on the rosette sheets.

##### Conversion of Full Files from Raw Data

All files were converted using 2012-05-ctd-new.con. Cast #58 could only be converted after fixing some headers that were on 1 line instead of 3 separate lines.

A few casts were examined and all expected channels are present. The descent rate varies greatly with some casts steady and others very noisy with obvious shed wake corruption.

The two temperature and conductivity channels are fairly close during the downcasts and, as usual, much farther apart during the upcasts.

Altimetry is sometimes spiky, but there appears to be a useful signal at the bottom, and fluorescence, PAR and transmissivity look normal.

The Surface PAR looks odd and will need examination later to see if it is useful.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on a few casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge because the temperature is so noisy on the upcast, but when downcasts only were compared, a setting of 3s produces the best match of T and DO features and that setting was also found appropriate for cruise 2012-57 which followed this cruise.

ALIGNCTD was used to advance the DO Voltage by 3.0s relative to the pressure.

##### CELLTM

The upcast data are noisy so the usual tests for CELLTM settings are difficult to interpret. The same equipment was used during six 2011 cruises. The tests for 2011-26 in June 2011 were reasonably clear and the best choice overall was found to be (α = 0.02, β=7) for the primary and (α = 0.03, β=9). Those results also appeared useful for the other 5 cruises when the same equipment was used. But for this cruise the results suggest the reverse is best, though differences were small. Perhaps there was a change of pumps and/or the way the equipment was mounted between 2011 and 2012. For 2012-57 which followed this cruise there are only 2 casts and they are noisy, but the results look similar to 2012-05.

CELLTM was run using (α = 0.03, β=9) for the primary conductivity and (α = 0.02, β=7) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

- on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. These are shallow casts, but a few were examined to ensure there were no obvious problems. For comparison one of the last casts from a September 2011 cruise is shown; the same T and C sensors were used then.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-17-0121 | 800 | -0.0004 | -0.00008 | -0.0005 | High, V Noisy |
| 2012-05-0028 | 315 | -0.0001 | +0.00015 | +0.0017 | Steady, high |
| 2012-05-0032 | 250 | 0.0001 | +0.00015 | +0.0018 | Steady, high |
| 2012-05-0039 | 350 | +0.0002 | +0.00014 | +0.0013 | Steady, high |

The differences are reasonably small, but the signs are different and differences are larger than for the earlier cruise. Given different pressure ranges, the change does not seem significant.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are a few negative values in pressure early in file #84, but they look like random spikes, not evidence of a pressure offset; they are associated with in-water values (salinity >29) and other pressure values near them are >4db. There are a few slightly negative values in the ECO fluorescence. Those are likely to disappear in processing.

Surface check was run and shows an average surface pressure for the cruise was 1.8db which looks a little low, especially for the casts in Juan de Fuca where conditions were rough. Examination of the bottle files suggests a problem because many of the surface samples have CTD pressure at ~1db, where we would expect 1.5db. A few other casts were checked and show similar surface pressures. One provided good evidence during upcasts that the surface is at about -0.5db. Pressure should be recalibrated.

The cross-reference check was compared with the log book and the only problem found was 1 cast with the wrong station name (#28); that was corrected in the full profile files. There was no bottle file for that cast.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. For 1 CLN file a spike was misinterpreted, but the altimetry traces shows a clear trend, so the reading was replaced with an estimate.

The Water Depth header was also examined. There were deviations from the log book entries or entries that were less than the maximum depth sampled; most look like the ship drifted between the CTD reaching the bottom and the bottom depth reading. Where the difference was >1db and supported by the altimetry the readings were changed to match the log book. In 1 case the log and header look clearly wrong, so an estimate based on the maximum pressure and altimetry reading were used to replace it.

No depth headers needed changing in the bottle files.

##### Shift

Fluorescence

Tests were run on two casts to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence traces look like that of the temperature trace. The results varied somewhat from cast to cast and the upcast temperature is very noisy making the comparison difficult. An advance of +12 records looks best overall for the ECO and +18 to +24 for the SeaPoint. During other recent uses of this equipment the settings used were +24 for the SeaPoint and either +6 or +12 records for the ECO.

SHIFT was run twice on all casts to advance the SeaPoint fluorescence by +24 records and the ECO fluorescence channel by +12 records.

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 3 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.5s was applied to the primary and a shift of -1.0 to the secondary conductivity.

SHIFT was run twice on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

pH

The pH sensor was not pumped. In the past a wide variety of alignment variables have been applied to this sensor, most recently +76 records. Tests run on these data suggest a value of +65 records provides good results overall. SHIFT was run applying that advance to the pH channel.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings are for cast #84 from the end of the upcast, so are of no concern to the downcast profile.

Header Check was repeated on the DEL files and the minimum fluorescence value is -0.016mL/L; this looks like noise in the signal and should disappear in bin-averaging.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late March 2011 and were used for 2011-44, 2011-16, 2011-26, 2011-09, 2011-27 and 2011-17. For the first 4 of these cruises there were few calibration samples and/or the bottle calibration was not trusted. No corrections were applied. For 2011-27 both sensors were low by about 0.0015 but the samples had been stored for 2.5 months which can lead to evaporation, making the CTD look low compared to bottles. For 2011-17 which was analyzed within a month the primary was found to be high by just 0.0001 and the secondary low by 0.0004. This supports the notion that evaporation affected the earlier analysis. There was also a cruise after this one, 2011-57, which has been partly processed, and it shows both sensors being low, by 0.008 and 0.006, but again the analysis was done more than 2 months after collection.

Dissolved Oxygen

The method for calibrating the DO sensors changed between the previous uses and this cruise.

The DO sensor was repaired and recalibrated in March 2011. It was used for 2011-44, 2011-16, 2011-26, 2011-09, 2011-27 and 2011-17. There appeared to be some time dependence in the 2011-26 data, but not in any of the others. 2011-27 required a 2nd calibration to further correct hysteresis.

Pressure

The sensor was recalibrated in April 2011 and was used for 2011-44, 2011-16, 2011-26, 2011-09 and 2011-27 and 2011-17. No further offset was applied to any of those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The only excursions from the ranges for the salinity data were some low surface values near the Fraser River and high values between 100 and 200db north of Orcas Island early in the cruise but low values between 100 and 120db for the final cast of the cruise which in the same general area. Temperature values were generally lower than the range minima at mid-depths north of 49.3°N.  All temperatures were within the ranges south of that. The ranges are considered too severe for these near-shore casts, so this is not indicative of instrumental problems. This likely reflects real conditions.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration - There were no post-cruise calibrations available.

##### DETAILED EDITING

The bottle comparison shows both sensors are close to bottles with the secondary being especially close. The secondary sensors were chosen for archiving for all 2011 cruises using this equipment. The secondary salinity is also a little less noisy than the primary, so the secondary T and S channels were chosen for editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. Editing was mostly light except for many of the Juan de Fuca casts where the CTD descent rate was very noisy with many complete reversals of direction.

All EDU files were copied to EDT.

##### . Dissolved Oxygen and fluorescence corrections

It was decided late in processing (after some editing had been done) that a new method suggested by SeaBird should be used for dissolved oxygen calibration for these data. This involved using a configuration file that contained different values of Voffset and SOC different. The change has no effect on any variables other than dissolved oxygen concentration. An error was also found in the edited data; the two fluorescence channels had not been shifted as planned. So those channels were re-processed, putting files with Pressure, temperature, conductivity, ECO fluorescence, SeaPoint fluorescence and DO only through all the relevant earlier steps including DELETE., but with the corrected configuration file.

Those files were then merged with the EDT files replacing the original DO channels (voltage and concentration) and fluorescence channels with the new versions. The merged files are named EDTMRG. (The bottle files were redone from the beginning using the new method, so no patching was needed for those.) Using this approach, the effects of editing are retained.

##### Initial Recalibration

The pressure requires recalibration by adding 0.5db.

Salinity will not be recalibrated. While there is scatter, the differences are not large, and the scatter is likely due to Autosal stability problems.

The dissolved oxygen concentration needs to be recalibrated using equation:

DO corrected = 1.0704 * CTD DO

CALIBRATE was run using file 2012-05-recal1.ccf to apply the corrections to the SAM files. COMPARE was run to ensure the correction worked well and it did. (See 2012-05-DO-comp2.xls.) CALIBRATE was then applied to the EDTMRG and MRGCLN2 files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure they were higher at the surface, small at depth. When plotted against DO concentration and outliers were excluded, the trendline was:

CTD DO– Bottle DO = 0.0117 * CTD DO + 0.037

and the average differences was +0.087mL/L.

This shows the data should be corrected as follows:

CTD DO Corrected = 0.9883 *CTD DO Uncorrected – 0.037

That correction was applied to the thinned files and COMPARE was rerun.

The results looked good with a very flat trendline for the plot of differences against DO concentration and an average difference of 0.0002mL/L. The slope versus pressure is slightly higher with the deepest samples reading slightly low, but by no more than 0.08mL/L. This looks as good as possible since the error associated with response time is related to the descent rate, so what suits the data from the bottom of casts will not do as well above.

(See 2012-05-dox-comp3.xls and 2012-05-dox-comp4.xls for details.)

CALIBRATE was run on the COR1 files to apply the above correction. (Output: COR2)

##### Special Fluorometer Processing

Special files were prepared for Dr. Peña by clipping the COR2 files to 150db. The clipped files were then processed in 2 ways: one set were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 while the other was first filtered (size 11, both fluorescence channels) and then put through REMOVE and HEADEDIT and named*.FCTD2. The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, size 11, was applied to the two fluorescence channels in the COR2 files.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. T-S plots generally looked fine though there were a few unstable features. Those were examined again to see if further editing was needed, but there was no evidence that the unstable features were caused by shed wakes or instrument problems, so no further editing was applied. Profile plots show many cases of near-surface low transmissivity (0-20%), mostly between the southern part of Texada Island and Pt. Roberts. There is again no evidence of instrumental problems; the Fraser River run-off may account for these features. The fluorescence sometimes peaks at the same level but is not particularly high for the cases with lowest transmissivity.

The Surface PAR data from this cruise look very odd. Near the beginning and end of the cruise, the traces may be ok, but for much of the cruise there are a lot of traces with no signal, or just a few spikes. We expect that at night, but many of these are in daylight hours. The other odd thing is that when there are sections of "no signal" the value differs. Early on the value is ~7.5, which seems reasonable. But from cast #13 to the end the value is between 245 and 265, varying from one cast to another. Even among the group between cast #78 and 87 where the signals look reasonable, the values may off if the "null" value is off. Since some of the traces may be useful, the channel will be left in the files, but a warning will be added to the header.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and Surface PAR data are nominal and unedited

except that some records were removed in editing temperature and salinity.

The Surface PAR data looks suspicious for some casts, though the sea-going

crew checked it occasionally and it appeared to be functioning properly.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision. It is not expected that this will make a

significant difference in the results.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.40mL/L from 0 to 50db

±0.15mL/L from 50 to 150db

±0.10mL/L below 150db (below 150db SBE DO values tended to be low)

For details on the processing see processing report: 2012-05-proc.doc.

A second set of files was produced in the same way except that the pH:SBE channel was not removed and the following comment was added to the header:

THIS FILE WAS PREPARED FOR THE USE OF THE CHIEF SCIENTIST. IT CONTAINS SBE pH

DATA WHICH WERE REMOVED FROM THE FILE TO BE PLACED IN THE IOS DATA LIBRARY,

BECAUSE THERE WAS NO FIELD CALIBRATION INFORMATION AVAILABLE AT THE TIME OF

PROCESSING, AND THERE ARE SOME CONCERNS ABOUT HOW THE SENSOR PERFORMS.

For cast #73, the pH:SBE channel was removed from this second set because the pH probe cap was left on. A note was added to the headers to explain this.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. There was a lot of variability which may reflect local mixing. The Juan de Fuca Strait and Haro Strait had values between 70% and 100%. To the north values were mostly between 90% and 110%. In the southern Strait of Georgia values were mostly between 100% and 120%. The wide range of values argues against a systematic error in dissolved oxygen, but does illustrate unusual variability. This may be due to wind mixing in Juan de Fuca where sampling was done between storms, while other areas were likely affected by the Fraser River outflow.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second set of casts (*.REMpH) to include pH was prepared by running REMOVE to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

For cast #73 which had the probe cap left on, pH was removed from this second set as well.

From this point on the two sets of files were treated in the same way.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and a few errors were found and fixed.

A header check was run on the final files and no problems were found.

Track plots look ok.

Plots of each file were examined to ensure no problems had crept in, and none were found.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. The only discrepancies were that there were some DO and Salinity samples not reported on the rosette sheets.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars (Notes from log)

# 13. Bottle 17 did not fire.

# 20. Continued increase in variation in fluorescence values.

# 41. Bottles 7&8 snapped @125m. Some doubt if bottle #11 snapped at 75m or at surface. There was no water in the bottle, so must have closed out of water. Vents opened before sampling for 30-45s affecting DO sampling.

# 46. Bottle #11 did not fire.

# 54. Air temp 12.7 deg C.

# 63. Depth deeper than in station file – stopped 10m off bottom. Weights on.

# 63-68 Weights on.

64-67. Raining during sampling noted by DO analyst.

# 69. Weights off

# 73. pH probe cap left on.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 1Apr2011 | Factory |  |  |
| Conductivity | 3396 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2668 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2754 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | Jan2012 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 1Apr11 | Factory |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| SBE Fluorometer | 2356 |  |  |  |  |
| Eco-AFL Fluorometer | 2215 | Aug2012 |  |  |  |
| Pressure Sensor | 75636 | 13Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |