## Notes
Analysis: | Glenn Cooper, 21-JUNE-12 to 25-JUNE-12 using Water Properties' Guildline 8400B Salinometer Serial Number 68572. Data processing by Kenny Scozzafava
Chief Scientist: | Peter Chandler
2012-07-05 |  | KS
Salinometer Calibration Information
Sample ID | Event | Station | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | New Calculated Salinity | Flag | Comments
Calibration#1 |  |  | K15 | 2012-06-21 13:24 | 24C | 0.99988 | 4.4783e-06 | -4.3035e-05 | 0.99984 |  |  | #VALUE!
Calibration#2 |  |  | P150std K0.9978 | 2012-06-22 14:21 | 24C | 0.99978 | 2.7426e-06 | -3.0398e-06 | 0.99978 |  |  | #VALUE!
2012-05#61 |  |  | P150std K0.9978 | 2012-06-22 14:23 | 24C | 0.99977 | 7.2639e-06 | -3.0398e-06 | 0.99977 | 34.9908 | 0.00028424 | #VALUE! |  | Ran standard to see what salinity value was calculated. Salinity value on std bottle 34.991
Continued to use P24 deep water reference to check the Salinometer in-between and within analysis runs
On June 22, between samples 147 and 146, analyst cleaned cell with CLR for 20 minutes followed by a rinse with 400 mL of dH2O and 400 mL of salt water. System was recalibrated after cleaning (above).
Salinometer readings were erratic and produced slightly higher than normal standard deviations of calculated salinity
Analyst noted that the salinometer was more stable after cleaning and following a weekend rest.
No replicate samples were collected therefore a precision analysis was not completed.
Peter Chandler's hydrometer calibration sample is included at the bottom of the QF worksheet.
2012-11-15 |  | KS
Outliers identified by the comparison to the CTD are noted in the QF worksheet and flagged appropriately.
In general, it seems as though the salinometer stability problems may have in fact contributed to increased scatter in the data in addition to some variablity in the T/S profiles themselves.
2015-03-30 |  | KS
Amended header in the QF worksheet to show that silicone inserts were used in place of the typical nylon inserts supplied by Guildline Instruments.

## QF 2012-05 SAL FINAL
Header: | Salinity samples were collected in 200 mL type ll glass bottles with reusable silicone inserts supplied by Xiamen Cuten Accessories Industry, Ltd. And
screw caps supplied by Ocean Scientific International Limited. They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer
standardized with IAPSO standard seawater 04 days after collection. For details including a duplicate analysis, see document QF2012-05SAL.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Salinity:Bottle | Flag:Salinity Bottle | Comments
2012-05 | 1 | SI | 1 | 30.6094 | 3 | SAL: Significant outlier in COMPARE not explained by noise in CTD signal
2012-05 | 2 | 59 | 10 | 32.3461
2012-05 | 2 | 59 | 11 | 32.3006
2012-05 | 2 | 59 | 12 | 32.2448
2012-05 | 2 | 59 | 13 | 31.972
2012-05 | 2 | 59 | 14 | 31.7172
2012-05 | 2 | 59 | 15 | 30.9768
2012-05 | 2 | 59 | 16 | 30.6031
2012-05 | 2 | 59 | 17 | 30.358
2012-05 | 2 | 59 | 18 | 30.1106
2012-05 | 2 | 59 | 19 | 30.0992
2012-05 | 2 | 59 | 20 | 30.0633
2012-05 | 2 | 59 | 21 | 29.5745
2012-05 | 2 | 59 | 22 | 29.1022
2012-05 | 2 | 59 | 23 | 29.0803
2012-05 | 7 | 56 | 25 | 31.7925
2012-05 | 13 | 14 | 38 | 30.5116
2012-05 | 13 | 14 | 39 | 30.4714
2012-05 | 13 | 14 | 40 | 30.4376
2012-05 | 13 | 14 | 41 | 30.3838
2012-05 | 13 | 14 | 42 | 30.3275
2012-05 | 13 | 14 | 43 | 30.2659
2012-05 | 13 | 14 | 44 | 30.0901 | 3 | SAL: Significant outlier in COMPARE not explained by noise in CTD signal
2012-05 | 13 | 14 | 45 | 29.8646
2012-05 | 13 | 14 | 46 | 29.3338
2012-05 | 13 | 14 | 47 | 29.2013
2012-05 | 13 | 14 | 48 | 29.1356
2012-05 | 13 | 14 | 49 | 28.9193
2012-05 | 13 | 14 | 50 | 27.9054
2012-05 | 13 | 14 | 51 | 26.5369
2012-05 | 13 | 14 | 52 | 25.8072
2012-05 | 15 | 16 | 54 | 30.355
2012-05 | 23 | 12 | 64 | 30.559
2012-05 | 23 | 12 | 65 | 30.5273
2012-05 | 23 | 12 | 66 | 30.4637
2012-05 | 23 | 12 | 67 | 30.4326
2012-05 | 23 | 12 | 68 | 30.3941
2012-05 | 23 | 12 | 69 | 30.3002
2012-05 | 23 | 12 | 70 | 30.1753
2012-05 | 23 | 12 | 71 | 30.0257 | 3 | SAL: Significant outlier in COMPARE not explained by noise in CTD signal
2012-05 | 23 | 12 | 72 | 29.8572 | 3 | SAL: Significant outlier in COMPARE not explained by noise in CTD signal
2012-05 | 23 | 12 | 73 | 29.6438
2012-05 | 23 | 12 | 74 | 29.4937
2012-05 | 23 | 12 | 75 | 29.2531
2012-05 | 23 | 12 | 76 | 28.6593
2012-05 | 23 | 12 | 77 | 27.3199
2012-05 | 23 | 12 | 78 | 25.2466
2012-05 | 23 | 12 | 79 | 24.1267
2012-05 | 25 | 9 | 81 | 30.4787 | 3 | SAL: Significant outlier in COMPARE not explained by noise in CTD signal
2012-05 | 27 | 6 | 93 | 30.4675
2012-05 | 33 | 3 | 106 | 30.7751
2012-05 | 37 | 2 | 122 | 30.6623
2012-05 | 41 | 27 | 136 | 30.8232
2012-05 | 41 | 27 | 137 | 30.7936
2012-05 | 41 | 27 | 138 | 30.7448
2012-05 | 41 | 27 | 139 | 30.6392
2012-05 | 41 | 27 | 140 | 30.5384
2012-05 | 41 | 27 | 141 | 30.4386
2012-05 | 41 | 27 | 142 | 30.1803
2012-05 | 41 | 27 | 143 | 30.0424
2012-05 | 41 | 27 | 144 | 29.8627
2012-05 | 41 | 27 | 146 | 29.6109
2012-05 | 41 | 27 | 147 | 29.5157 | 2 | SAL: Cleaned cell with CLR in the hopes of reducing the erratic readings
2012-05 | 41 | 27 | 148 | 29.1782
2012-05 | 41 | 27 | 149 | 26.6374
2012-05 | 41 | 27 | 150 | 22.4941
2012-05 | 46 | 39 | 153 | 30.8043
2012-05 | 51 | 42 | 169 | 30.7434
2012-05 | 51 | 42 | 170 | 30.7421
2012-05 | 51 | 42 | 171 | 30.7142
2012-05 | 51 | 42 | 172 | 30.6208
2012-05 | 51 | 42 | 173 | 30.5137
2012-05 | 51 | 42 | 174 | 30.5944 | 3 | SAL: Significant outlier in COMPARE not explained by noise in CTD signal
2012-05 | 51 | 42 | 175 | 30.1892
2012-05 | 51 | 42 | 176 | 30.0219
2012-05 | 51 | 42 | 177 | 29.7784
2012-05 | 51 | 42 | 178 | 29.4976 | 3 | SAL: Significant outlier in COMPARE not explained by noise in CTD signal
2012-05 | 51 | 42 | 179 | 29.4317
2012-05 | 51 | 42 | 180 | 29.4302
2012-05 | 51 | 42 | 181 | 29.4132
2012-05 | 51 | 42 | 182 | 29.2076
2012-05 | 51 | 42 | 183 | 28.5756
2012-05 | 51 | 42 | 184 | 25.9355
2012-05 | 51 | 42 | 185 | 6.1316
2012-05 | 57 | 46 | 187 | 30.7362
2012-05 | 64 | 102 | 198 | 33.957
2012-05 | 64 | 102 | 199 | 33.9254
2012-05 | 64 | 102 | 200 | 33.8811
2012-05 | 64 | 102 | 201 | 33.8309
2012-05 | 64 | 102 | 202 | 33.7687
2012-05 | 64 | 102 | 203 | 33.5747
2012-05 | 64 | 102 | 204 | 33.3624
2012-05 | 64 | 102 | 205 | 32.8387
2012-05 | 64 | 102 | 206 | 32.3523 | 2 | SAL: Flushed cell with dH2O prior to next sample. Autosal erratic.
2012-05 | 64 | 102 | 207 | 32.2225
2012-05 | 64 | 102 | 208 | 31.8639
2012-05 | 64 | 102 | 209 | 31.1318
2012-05 | 64 | 102 | 210 | 31.0527
2012-05 | 64 | 102 | 211 | 31.0519
2012-05 | 67 | 75 | 213 | 33.9168
2012-05 | 70 | 72 | 226 | 33.8934
2012-05 | 70 | 72 | 227 | 33.8929
2012-05 | 70 | 72 | 228 | 33.8908
2012-05 | 70 | 72 | 229 | 33.8853
2012-05 | 70 | 72 | 230 | 33.8906
2012-05 | 70 | 72 | 231 | 32.9491
2012-05 | 70 | 72 | 232 | 32.1689
2012-05 | 70 | 72 | 233 | 31.6699
2012-05 | 70 | 72 | 234 | 31.476
2012-05 | 70 | 72 | 235 | 31.0102
2012-05 | 70 | 72 | 236 | 30.7772
2012-05 | 70 | 72 | 237 | 30.7311
2012-05 | 70 | 72 | 238 | 30.6907
2012-05 | 73 | 69 | 240 | 33.8643
2012-05 | 76 | ADCP | 252 | 33.0186
2012-05 | 80 | 65 | 260 | 33.4245
2012-05 | 80 | 65 | 261 | 32.7934
2012-05 | 80 | 65 | 262 | 32.0068
2012-05 | 80 | 65 | 263 | 31.3856
2012-05 | 80 | 65 | 264 | 31.2573
2012-05 | 80 | 65 | 265 | 30.7617
2012-05 | 80 | 65 | 266 | 30.1367 | 2 | SAL: Lost communication between Autosal and computer. Re-established communication and continued analysis
2012-05 | 80 | 65 | 267 | 29.8423
2012-05 | 80 | 65 | 268 | 29.8072
2012-05 | 80 | 65 | 269 | 29.7953
2012-05 | 83 | 62 | 271 | 32.7366
2012-05 |  |  | Test | 31.1437 |  | SAL: Peter Chandler hydrometer standardization

## Flags
Institute of Ocean Sciences - Quality Flags for Water Sampling | Last Updated: October 2011
Double click document to use.

## Salinity Data Processing
Sample ID | Event | Station | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | New Calculated Salinity | Flag | Comments
2012-05#124 | 1 | SI | 1 | 2012-06-25 12:01 | 24C | 0.88731 | 4.6106e-06 | -3.0994e-06 | 0.88731 | 30.6094 | 0.00017741 | #VALUE!
2012-05#123 | 2 | 59 | 10 | 2012-06-25 11:59 | 24C | 0.93216 | 2.4754e-06 | -2.9802e-06 | 0.93216 | 32.3461 | 9.6423e-05 | #VALUE!
2012-05#122 | 2 | 59 | 11 | 2012-06-25 11:57 | 24C | 0.93099 | 4.5983e-06 | -2.9802e-06 | 0.93099 | 32.3006 | 0.00017701 | #VALUE!
2012-05#121 | 2 | 59 | 12 | 2012-06-25 11:55 | 24C | 0.92956 | 3.4657e-06 | -2.9802e-06 | 0.92955 | 32.2448 | 0.00013386 | #VALUE!
2012-05#120 | 2 | 59 | 13 | 2012-06-25 11:53 | 24C | 0.92253 | 5.007e-06 | -3.0994e-06 | 0.92253 | 31.972 | 0.00019031 | #VALUE!
2012-05#119 | 2 | 59 | 14 | 2012-06-25 11:51 | 24C | 0.91596 | 3.862e-06 | -2.9802e-06 | 0.91596 | 31.7172 | 0.00014806 | #VALUE!
2012-05#118 | 2 | 59 | 15 | 2012-06-25 11:49 | 24C | 0.89683 | 4.837e-06 | -3.0994e-06 | 0.89683 | 30.9768 | 0.00018563 | #VALUE!
2012-05#117 | 2 | 59 | 16 | 2012-06-25 11:47 | 24C | 0.88714 | 3.6105e-06 | -2.9802e-06 | 0.88714 | 30.6031 | 0.00014202 | #VALUE!
2012-05#116 | 2 | 59 | 17 | 2012-06-25 11:44 | 24C | 0.88078 | 5.8621e-06 | -2.9802e-06 | 0.88078 | 30.358 | 0.00022744 | #VALUE!
2012-05#115 | 2 | 59 | 18 | 2012-06-25 11:42 | 24C | 0.87436 | 4.3699e-06 | -3.0994e-06 | 0.87435 | 30.1106 | 0.00017031 | #VALUE!
2012-05#114 | 2 | 59 | 19 | 2012-06-25 11:38 | 24C | 0.87406 | 3.4494e-06 | -3.0398e-06 | 0.87406 | 30.0992 | 0.0001343 | #VALUE!
2012-05#113 | 2 | 59 | 20 | 2012-06-25 11:35 | 24C | 0.87313 | 3.8543e-06 | -2.9802e-06 | 0.87312 | 30.0633 | 0.00015089 | #VALUE!
2012-05#112 | 2 | 59 | 21 | 2012-06-25 11:33 | 24C | 0.8604 | 4.2092e-06 | -2.9802e-06 | 0.8604 | 29.5745 | 0.00016147 | #VALUE!
2012-05#111 | 2 | 59 | 22 | 2012-06-25 11:31 | 24C | 0.84808 | 4.952e-06 | -2.9802e-06 | 0.84807 | 29.1022 | 0.00018882 | #VALUE!
2012-05#110 | 2 | 59 | 23 | 2012-06-25 11:29 | 24C | 0.8475 | 3.85e-06 | -2.9802e-06 | 0.8475 | 29.0803 | 0.00014604 | #VALUE!
2012-05#109 | 7 | 56 | 25 | 2012-06-25 11:27 | 24C | 0.9179 | 5.0205e-06 | -3.0398e-06 | 0.9179 | 31.7925 | 0.00019469 | #VALUE!
2012-05#108 | 13 | 14 | 38 | 2012-06-25 11:25 | 24C | 0.88477 | 5.2797e-06 | -2.9802e-06 | 0.88477 | 30.5116 | 0.00020323 | #VALUE!
2012-05#107 | 13 | 14 | 39 | 2012-06-25 11:23 | 24C | 0.88373 | 3.865e-06 | -3.0398e-06 | 0.88373 | 30.4714 | 0.00014655 | #VALUE!
2012-05#106 | 13 | 14 | 40 | 2012-06-25 11:20 | 24C | 0.88285 | 4.4783e-06 | -3.0994e-06 | 0.88285 | 30.4376 | 0.00017423 | #VALUE!
2012-05#105 | 13 | 14 | 41 | 2012-06-25 11:17 | 24C | 0.88145 | 5.1607e-06 | -3.0398e-06 | 0.88145 | 30.3838 | 0.00019536 | #VALUE!
2012-05#104 | 13 | 14 | 42 | 2012-06-25 11:15 | 24C | 0.87999 | 4.949e-06 | -3.0398e-06 | 0.87999 | 30.3275 | 0.00019108 | #VALUE!
2012-05#103 | 13 | 14 | 43 | 2012-06-25 11:13 | 24C | 0.87839 | 5.3479e-06 | -3.0398e-06 | 0.87839 | 30.2659 | 0.00020399 | #VALUE!
2012-05#101 | 13 | 14 | 44 | 2012-06-25 11:09 | 24C | 0.87382 | 5.5671e-06 | -2.9802e-06 | 0.87382 | 30.0901 | 0.00021551 | #VALUE!
2012-05#102 | 13 | 14 | 45 | 2012-06-25 11:11 | 24C | 0.86796 | 4.322e-06 | -3.0398e-06 | 0.86795 | 29.8646 | 0.00016668 | #VALUE!
2012-05#93 | 13 | 14 | 46 | 2012-06-22 15:58 | 24C | 0.85413 | 8.6316e-06 | -3.0398e-06 | 0.85412 | 29.3338 | 0.00033154 | #VALUE!
2012-05#92 | 13 | 14 | 47 | 2012-06-22 15:56 | 24C | 0.85066 | 6.5372e-06 | -3.0398e-06 | 0.85066 | 29.2013 | 0.00025104 | #VALUE!
2012-05#95 | 13 | 14 | 48 | 2012-06-22 16:04 | 24C | 0.84895 | 1.1514e-05 | -3.0398e-06 | 0.84895 | 29.1356 | 0.000443 | #VALUE!
2012-05#94 | 13 | 14 | 49 | 2012-06-22 16:01 | 24C | 0.8433 | 6.1735e-06 | -2.9802e-06 | 0.84329 | 28.9193 | 0.00023894 | #VALUE!
2012-05#96 | 13 | 14 | 50 | 2012-06-22 16:06 | 24C | 0.81672 | 6.311e-06 | -2.9802e-06 | 0.81671 | 27.9054 | 0.00023915 | #VALUE!
2012-05#132 | 13 | 14 | 51 | 2012-06-25 12:23 | 24C | 0.78062 | 4.7302e-06 | -2.861e-06 | 0.78062 | 26.5369 | 0.0001764 | #VALUE!
2012-05#98 | 13 | 14 | 52 | 2012-06-22 16:10 | 24C | 0.76128 | 1.0334e-05 | -3.0994e-06 | 0.76128 | 25.8072 | 0.0003895 | #VALUE!
2012-05#97 | 15 | 16 | 54 | 2012-06-22 16:08 | 24C | 0.88071 | 1.0789e-05 | -2.9802e-06 | 0.8807 | 30.355 | 0.00041391 | #VALUE! |  | Corrected Station ID
2012-05#85 | 23 | 12 | 64 | 2012-06-22 15:23 | 24C | 0.886 | 1.1266e-05 | -3.0398e-06 | 0.886 | 30.559 | 0.00043418 | #VALUE!
2012-05#84 | 23 | 12 | 65 | 2012-06-22 15:20 | 24C | 0.88518 | 1.4623e-05 | -2.9206e-06 | 0.88518 | 30.5273 | 0.00056388 | #VALUE!
2012-05#86 | 23 | 12 | 66 | 2012-06-22 15:25 | 24C | 0.88353 | 1.0312e-05 | -2.9802e-06 | 0.88353 | 30.4637 | 0.00039633 | #VALUE!
2012-05#88 | 23 | 12 | 67 | 2012-06-22 15:47 | 24C | 0.88272 | 1.0967e-05 | -3.0994e-06 | 0.88272 | 30.4326 | 0.00042292 | #VALUE!
2012-05#87 | 23 | 12 | 68 | 2012-06-22 15:44 | 24C | 0.88172 | 1.3754e-05 | -3.0398e-06 | 0.88172 | 30.3941 | 0.00052881 | #VALUE!
2012-05#89 | 23 | 12 | 69 | 2012-06-22 15:49 | 24C | 0.87928 | 1.321e-05 | -2.9802e-06 | 0.87928 | 30.3002 | 0.00050908 | #VALUE!
2012-05#90 | 23 | 12 | 70 | 2012-06-22 15:51 | 24C | 0.87604 | 5.4475e-06 | -2.9802e-06 | 0.87604 | 30.1753 | 0.0002091 | #VALUE!
2012-05#81 | 23 | 12 | 71 | 2012-06-22 15:14 | 24C | 0.87215 | 1.0861e-05 | -3.0994e-06 | 0.87214 | 30.0257 | 0.00041704 | #VALUE!
2012-05#80 | 23 | 12 | 72 | 2012-06-22 15:12 | 24C | 0.86777 | 1.362e-05 | -3.0398e-06 | 0.86776 | 29.8572 | 0.00052627 | #VALUE!
2012-05#79 | 23 | 12 | 73 | 2012-06-22 15:10 | 24C | 0.86221 | 2.2258e-06 | -3.0398e-06 | 0.8622 | 29.6438 | 8.5943e-05 | #VALUE!
2012-05#78 | 23 | 12 | 74 | 2012-06-22 15:08 | 24C | 0.8583 | 9.2117e-06 | -3.0398e-06 | 0.85829 | 29.4841 | 0.00035556 | #VALUE!
2012-05#76 | 23 | 12 | 75 | 2012-06-22 15:03 | 24C | 0.85202 | 8.252e-06 | -2.9802e-06 | 0.85201 | 29.2531 | 0.00031617 | #VALUE!
2012-05#77 | 23 | 12 | 76 | 2012-06-22 15:06 | 24C | 0.83649 | 8.4361e-06 | -2.9206e-06 | 0.83649 | 28.6592 | 0.00032096 | #VALUE!
2012-05#91 | 23 | 12 | 77 | 2012-06-22 15:54 | 24C | 0.80131 | 7.0799e-06 | -3.0398e-06 | 0.8013 | 27.3199 | 0.00026844 | #VALUE!
2012-05#83 | 23 | 12 | 78 | 2012-06-22 15:18 | 24C | 0.74636 | 1.2921e-05 | -3.0398e-06 | 0.74636 | 25.2466 | 0.00048579 | #VALUE!
2012-05#82 | 23 | 12 | 79 | 2012-06-22 15:16 | 24C | 0.71644 | 2.4908e-06 | -2.9802e-06 | 0.71643 | 24.1267 | 9.2752e-05 | #VALUE!
2012-05#59 | 25 | 9 | 81 | 2012-06-22 13:32 | 24C | 0.88396 | 1.0341e-05 | -4.2915e-05 | 0.88391 | 30.4787 | 0.00039641 | #VALUE!
2012-05#58 | 27 | 6 | 93 | 2012-06-22 13:30 | 24C | 0.88367 | 1.4043e-05 | -4.2915e-05 | 0.88362 | 30.4675 | 0.00054103 | #VALUE!
2012-05#57 | 33 | 3 | 106 | 2012-06-22 13:27 | 24C | 0.89164 | 6.8255e-06 | -4.2796e-05 | 0.8916 | 30.7751 | 0.00026198 | #VALUE!
2012-05#56 | 37 | 2 | 122 | 2012-06-22 13:25 | 24C | 0.88872 | 7.7727e-06 | -4.2915e-05 | 0.88868 | 30.6623 | 0.00030321 | #VALUE!
2012-05#52 | 41 | 27 | 136 | 2012-06-22 13:15 | 24C | 0.89289 | 1.089e-05 | -4.2915e-05 | 0.89285 | 30.8232 | 0.00041927 | #VALUE!
2012-05#53 | 41 | 27 | 137 | 2012-06-22 13:18 | 24C | 0.89212 | 2.2401e-06 | -4.2915e-05 | 0.89208 | 30.7936 | 8.5313e-05 | #VALUE!
2012-05#54 | 41 | 27 | 138 | 2012-06-22 13:21 | 24C | 0.89086 | 1.2971e-05 | -4.2915e-05 | 0.89082 | 30.7448 | 0.0005022 | #VALUE!
2012-05#55 | 41 | 27 | 139 | 2012-06-22 13:22 | 24C | 0.88812 | 1.0349e-05 | -4.2915e-05 | 0.88808 | 30.6392 | 0.00039779 | #VALUE!
2012-05#65 | 41 | 27 | 140 | 2012-06-22 14:33 | 24C | 0.88547 | 1.122e-05 | -2.9802e-06 | 0.88547 | 30.5384 | 0.00043195 | #VALUE!
2012-05#67 | 41 | 27 | 141 | 2012-06-22 14:42 | 24C | 0.88288 | 7.3032e-06 | -2.9206e-06 | 0.88288 | 30.4386 | 0.00027802 | #VALUE!
2012-05#66 | 41 | 27 | 142 | 2012-06-22 14:40 | 24C | 0.87617 | 1.1033e-05 | -3.0398e-06 | 0.87617 | 30.1803 | 0.0004249 | #VALUE!
2012-05#64 | 41 | 27 | 143 | 2012-06-22 14:31 | 24C | 0.87258 | 8.6767e-06 | -3.0398e-06 | 0.87258 | 30.0424 | 0.00032921 | #VALUE!
2012-05#63 | 41 | 27 | 144 | 2012-06-22 14:29 | 24C | 0.86791 | 8.4303e-06 | -2.9802e-06 | 0.86791 | 29.8627 | 0.00032325 | #VALUE!
2012-05#62 | 41 | 27 | 146 | 2012-06-22 14:27 | 24C | 0.86135 | 1.2557e-05 | -3.0994e-06 | 0.86135 | 29.6109 | 0.00048169 | #VALUE!
2012-05#60 | 41 | 27 | 147 | 2012-06-22 13:34 | 24C | 0.85891 | 9.4813e-06 | -4.2915e-05 | 0.85886 | 29.5157 | 0.00036562 | #VALUE! | 2 | Cleaned cell with CLR in the hopes of reducing the erratic readings
2012-05#73 | 41 | 27 | 148 | 2012-06-22 14:56 | 24C | 0.85006 | 2.7425e-06 | -2.9206e-06 | 0.85006 | 29.1782 | 0.0001097 | #VALUE!
2012-05#74 | 41 | 27 | 149 | 2012-06-22 14:58 | 24C | 0.78329 | 8.4963e-06 | -2.9802e-06 | 0.78328 | 26.6374 | 0.00032103 | #VALUE!
2012-05#75 | 41 | 27 | 150 | 2012-06-22 15:01 | 24C | 0.67248 | 6.7855e-06 | -3.0398e-06 | 0.67248 | 22.4941 | 0.0002516 | #VALUE!
2012-05#72 | 46 | 39 | 153 | 2012-06-22 14:53 | 24C | 0.89236 | 7.5885e-06 | -2.9802e-06 | 0.89236 | 30.8043 | 0.00029127 | #VALUE!
2012-05#69 | 51 | 42 | 169 | 2012-06-22 14:46 | 24C | 0.89078 | 9.394e-06 | -2.9802e-06 | 0.89078 | 30.7434 | 0.00036425 | #VALUE!
2012-05#70 | 51 | 42 | 170 | 2012-06-22 14:48 | 24C | 0.89075 | 9.2917e-06 | -2.9802e-06 | 0.89075 | 30.7421 | 0.00035775 | #VALUE!
2012-05#71 | 51 | 42 | 171 | 2012-06-22 14:50 | 24C | 0.89003 | 1.372e-06 | -2.9802e-06 | 0.89002 | 30.7142 | 5.2246e-05 | #VALUE!
2012-05#68 | 51 | 42 | 172 | 2012-06-22 14:44 | 24C | 0.8876 | 8.054e-06 | -3.0994e-06 | 0.8876 | 30.6208 | 0.00031247 | #VALUE!
2012-05#33 | 51 | 42 | 173 | 2012-06-22 09:31 | 24C | 0.88487 | 4.9693e-06 | -4.2915e-05 | 0.88482 | 30.5137 | 0.00019375 | #VALUE!
2012-05#32 | 51 | 42 | 174 | 2012-06-22 09:29 | 24C | 0.88696 | 2.4744e-06 | -4.3035e-05 | 0.88692 | 30.5944 | 9.4885e-05 | #VALUE!
2012-05#31 | 51 | 42 | 175 | 2012-06-22 09:26 | 24C | 0.87644 | 2.2391e-06 | -4.2915e-05 | 0.8764 | 30.1892 | 8.5299e-05 | #VALUE!
2012-05#29 | 51 | 42 | 176 | 2012-06-22 09:03 | 24C | 0.87209 | 3.3595e-06 | -4.3035e-05 | 0.87205 | 30.0219 | 0.00012705 | #VALUE!
2012-05#30 | 51 | 42 | 177 | 2012-06-22 09:06 | 24C | 0.86575 | 3.8747e-06 | -4.2915e-05 | 0.86571 | 29.7784 | 0.00014713 | #VALUE!
2012-05#28 | 51 | 42 | 178 | 2012-06-22 09:00 | 24C | 0.85844 | 5.3443e-06 | -4.2915e-05 | 0.85839 | 29.4976 | 0.00020321 | #VALUE!
2012-05#27 | 51 | 42 | 179 | 2012-06-22 08:56 | 24C | 0.85672 | 2.4747e-06 | -4.2975e-05 | 0.85668 | 29.4317 | 9.4885e-05 | #VALUE!
2012-05#26 | 51 | 42 | 180 | 2012-06-22 08:53 | 24C | 0.85668 | 3.9464e-06 | -4.2975e-05 | 0.85664 | 29.4302 | 0.00015308 | #VALUE!
2012-05#42 | 51 | 42 | 181 | 2012-06-22 09:58 | 24C | 0.85624 | 1.3341e-05 | -4.2975e-05 | 0.85619 | 29.4132 | 0.00051305 | #VALUE! |  | Corrected station ID
2012-05#40 | 51 | 42 | 182 | 2012-06-22 09:50 | 24C | 0.85087 | 3.4657e-06 | -4.3035e-05 | 0.85083 | 29.2076 | 0.00013637 | #VALUE!
2012-05#39 | 51 | 42 | 183 | 2012-06-22 09:47 | 24C | 0.83434 | 4.8104e-06 | -4.2915e-05 | 0.8343 | 28.5756 | 0.0001834 | #VALUE!
2012-05#38 | 51 | 42 | 184 | 2012-06-22 09:44 | 24C | 0.76472 | 5.3433e-06 | -4.3035e-05 | 0.76468 | 25.9355 | 0.00020329 | #VALUE!
2012-05#37 | 51 | 42 | 185 | 2012-06-22 09:40 | 24C | 0.20402 | 6.4355e-06 | -4.3035e-05 | 0.20398 | 6.1316 | 0.0002085 | #VALUE!
2012-05#36 | 57 | 46 | 187 | 2012-06-22 09:38 | 24C | 0.89064 | 4.8375e-06 | -4.3154e-05 | 0.89059 | 30.7362 | 0.00018329 | #VALUE! |  | Corrected station ID
2012-05#35 | 64 | 102 | 198 | 2012-06-22 09:36 | 24C | 0.97348 | 3.2147e-06 | -4.3035e-05 | 0.97344 | 33.957 | 0.00012429 | #VALUE!
2012-05#34 | 64 | 102 | 199 | 2012-06-22 09:33 | 24C | 0.97267 | 3.4551e-06 | -4.2915e-05 | 0.97263 | 33.9254 | 0.00013134 | #VALUE!
2012-05#51 | 64 | 102 | 200 | 2012-06-22 12:04 | 24C | 0.97154 | 1.059e-05 | -4.2915e-05 | 0.9715 | 33.8811 | 0.00041717 | #VALUE!
2012-05#49 | 64 | 102 | 201 | 2012-06-22 12:00 | 24C | 0.97026 | 1.2732e-05 | -4.2915e-05 | 0.97022 | 33.8309 | 0.00049484 | #VALUE!
2012-05#50 | 64 | 102 | 202 | 2012-06-22 12:02 | 24C | 0.96867 | 1.1286e-05 | -4.2856e-05 | 0.96863 | 33.7687 | 0.000443 | #VALUE!
2012-05#48 | 64 | 102 | 203 | 2012-06-22 11:58 | 24C | 0.96371 | 1.1895e-05 | -4.2915e-05 | 0.96367 | 33.5747 | 0.00046415 | #VALUE!
2012-05#47 | 64 | 102 | 204 | 2012-06-22 11:56 | 24C | 0.95828 | 1.1603e-05 | -4.2915e-05 | 0.95824 | 33.3624 | 0.00045334 | #VALUE!
2012-05#46 | 64 | 102 | 205 | 2012-06-22 11:53 | 24C | 0.94486 | 1.0154e-05 | -4.2796e-05 | 0.94481 | 32.8387 | 0.00039621 | #VALUE!
2012-05#44 | 64 | 102 | 206 | 2012-06-22 10:04 | 24C | 0.93236 | 1.5236e-05 | -4.2915e-05 | 0.93232 | 32.3523 | 0.00059097 | #VALUE! | 2 | Flushed cell with dH2O prior to next sample. Autosal erratic.
2012-05#43 | 64 | 102 | 207 | 2012-06-22 10:00 | 24C | 0.92902 | 1.6626e-05 | -4.2915e-05 | 0.92898 | 32.2225 | 0.00064505 | #VALUE!
2012-05#9 | 64 | 102 | 208 | 2012-06-21 13:47 | 24C | 0.91979 | 3.6104e-06 | -4.3035e-05 | 0.91974 | 31.8639 | 0.00013874 | #VALUE!
2012-05#8 | 64 | 102 | 209 | 2012-06-21 13:45 | 24C | 0.90088 | 3.8543e-06 | -4.3094e-05 | 0.90084 | 31.1318 | 0.00014848 | #VALUE!
2012-05#7 | 64 | 102 | 210 | 2012-06-21 13:43 | 24C | 0.89883 | 4.2256e-06 | -4.2975e-05 | 0.89879 | 31.0527 | 0.00016065 | #VALUE!
2012-05#6 | 64 | 102 | 211 | 2012-06-21 13:41 | 24C | 0.89881 | 4.3044e-06 | -4.2975e-05 | 0.89877 | 31.0519 | 0.00016796 | #VALUE!
2012-05#5 | 67 | 75 | 213 | 2012-06-21 13:38 | 24C | 0.97245 | 5.3443e-06 | -4.3035e-05 | 0.97241 | 33.9168 | 0.00020985 | #VALUE!
2012-05#4 | 70 | 72 | 226 | 2012-06-21 13:33 | 24C | 0.97186 | 2.8904e-06 | -4.3035e-05 | 0.97181 | 33.8934 | 0.00011317 | #VALUE!
2012-05#3 | 70 | 72 | 227 | 2012-06-21 13:30 | 24C | 0.97184 | 5.9279e-06 | -4.2915e-05 | 0.9718 | 33.8929 | 0.00023572 | #VALUE!
2012-05#2 | 70 | 72 | 228 | 2012-06-21 13:28 | 24C | 0.97179 | 4.73e-06 | -4.3035e-05 | 0.97175 | 33.8908 | 0.00018758 | #VALUE!
2012-05#17 | 70 | 72 | 229 | 2012-06-21 14:04 | 24C | 0.97165 | 2.7425e-06 | -4.3035e-05 | 0.97161 | 33.8853 | 0.00010656 | #VALUE!
2012-05#16 | 70 | 72 | 230 | 2012-06-21 14:02 | 24C | 0.97178 | 6.5729e-06 | -4.2975e-05 | 0.97174 | 33.8906 | 0.0002579 | #VALUE!
2012-05#15 | 70 | 72 | 231 | 2012-06-21 14:00 | 24C | 0.94769 | 1.1199e-06 | -4.2975e-05 | 0.94765 | 32.9491 | 4.4364e-05 | #VALUE!
2012-05#14 | 70 | 72 | 232 | 2012-06-21 13:57 | 24C | 0.92764 | 3.5301e-06 | -4.3154e-05 | 0.9276 | 32.1689 | 0.00013763 | #VALUE!
2012-05#13 | 70 | 72 | 233 | 2012-06-21 13:55 | 24C | 0.91478 | 4.3221e-06 | -4.2975e-05 | 0.91474 | 31.6699 | 0.00016682 | #VALUE!
2012-05#12 | 70 | 72 | 234 | 2012-06-21 13:53 | 24C | 0.90978 | 3.1417e-06 | -4.3035e-05 | 0.90973 | 31.476 | 0.00012104 | #VALUE!
2012-05#11 | 70 | 72 | 235 | 2012-06-21 13:51 | 24C | 0.89773 | 6.5631e-06 | -4.2975e-05 | 0.89769 | 31.0102 | 0.00025268 | #VALUE!
2012-05#10 | 70 | 72 | 236 | 2012-06-21 13:49 | 24C | 0.8917 | 3.4516e-06 | -4.3035e-05 | 0.89166 | 30.7772 | 0.00013157 | #VALUE!
2012-05#25 | 70 | 72 | 237 | 2012-06-21 14:21 | 24C | 0.89051 | 4.8368e-06 | -4.2975e-05 | 0.89046 | 30.7311 | 0.00018769 | #VALUE!
2012-05#24 | 70 | 72 | 238 | 2012-06-21 14:19 | 24C | 0.88946 | 4.9398e-06 | -4.3035e-05 | 0.88941 | 30.6907 | 0.00019284 | #VALUE!
2012-05#23 | 73 | 69 | 240 | 2012-06-21 14:17 | 24C | 0.97111 | 7.2858e-06 | -4.3035e-05 | 0.97107 | 33.8643 | 0.00028361 | #VALUE!
2012-05#22 | 76 | ADCP | 252 | 2012-06-21 14:15 | 24C | 0.94947 | 3.8619e-06 | -4.3035e-05 | 0.94943 | 33.0186 | 0.00015041 | #VALUE!
2012-05#21 | 80 | 65 | 260 | 2012-06-21 14:12 | 24C | 0.95987 | 3.4495e-06 | -4.3035e-05 | 0.95983 | 33.4245 | 0.00013655 | #VALUE!
2012-05#19 | 80 | 65 | 261 | 2012-06-21 14:08 | 24C | 0.9437 | 4.3336e-06 | -4.3035e-05 | 0.94365 | 32.7934 | 0.00017344 | #VALUE!
2012-05#20 | 80 | 65 | 262 | 2012-06-21 14:10 | 24C | 0.92347 | 4.5985e-06 | -4.2975e-05 | 0.92343 | 32.0068 | 0.00017986 | #VALUE!
2012-05#18 | 80 | 65 | 263 | 2012-06-21 14:06 | 24C | 0.90744 | 3.3591e-06 | -4.3035e-05 | 0.9074 | 31.3856 | 0.00013214 | #VALUE!
2012-05#125 | 80 | 65 | 264 | 2012-06-25 12:04 | 24C | 0.90409 | 4.1735e-06 | -3.0398e-06 | 0.90408 | 31.2573 | 0.00016082 | #VALUE!
2012-05#126 | 80 | 65 | 265 | 2012-06-25 12:07 | 24C | 0.89126 | 1.1199e-06 | -2.9802e-06 | 0.89126 | 30.7617 | 4.265e-05 | #VALUE!
2012-05#127 | 80 | 65 | 266 | 2012-06-25 12:10 | 24C | 0.87503 | 5.8674e-06 | -2.9206e-06 | 0.87503 | 30.1367 | 0.00022328 | #VALUE! | 2 | Lost communication between Autosal and computer. Re-established communication and continued analysis
2012-05#128 | 80 | 65 | 267 | 2012-06-25 12:13 | 24C | 0.86738 | 6.425e-06 | -2.861e-06 | 0.86737 | 29.8423 | 0.00024802 | #VALUE!
2012-05#129 | 80 | 65 | 268 | 2012-06-25 12:15 | 24C | 0.86646 | 3.8607e-06 | -2.861e-06 | 0.86646 | 29.8072 | 0.00014954 | #VALUE!
2012-05#130 | 80 | 65 | 269 | 2012-06-25 12:18 | 24C | 0.86615 | 5.0134e-06 | -2.861e-06 | 0.86615 | 29.7953 | 0.00019111 | #VALUE!
2012-05#131 | 83 | 62 | 271 | 2012-06-25 12:20 | 24C | 0.9422 | 2.2394e-06 | -2.861e-06 | 0.94219 | 32.7366 | 8.7102e-05 | #VALUE!
2012-05#41 |  |  | Test | 2012-06-22 09:54 | 24C | 0.90119 | 1.8257e-05 | -4.2915e-05 | 0.90114 | 31.1437 | 0.00070721 | #VALUE! |  | Peter Chandler hydrometer standardization

## RAW Salinometer Output
Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 | K15 | 2012-06-21 13:24 | 24C | 0.99988 | 4.4783e-06 | -4.3035e-05 | 0.99984
2012-05#1 | P24-DWR | 2012-06-21 13:26 | 24C | 0.98934 | 5.0844e-06 | -4.3035e-05 | 0.9893 | 34.5793 | 0.00020192
2012-05#2 | sta 72-228 | 2012-06-21 13:28 | 24C | 0.97179 | 4.73e-06 | -4.3035e-05 | 0.97175 | 33.8908 | 0.00018758
2012-05#3 | sta72-227 | 2012-06-21 13:30 | 24C | 0.97184 | 5.9279e-06 | -4.2915e-05 | 0.9718 | 33.8929 | 0.00023572
2012-05#4 | sta 72-226 | 2012-06-21 13:33 | 24C | 0.97186 | 2.8904e-06 | -4.3035e-05 | 0.97181 | 33.8934 | 0.00011317
2012-05#5 | Sta 75-213 | 2012-06-21 13:38 | 24C | 0.97245 | 5.3443e-06 | -4.3035e-05 | 0.97241 | 33.9168 | 0.00020985
2012-05#6 | Sta 102-211 | 2012-06-21 13:41 | 24C | 0.89881 | 4.3044e-06 | -4.2975e-05 | 0.89877 | 31.0519 | 0.00016796
2012-05#7 | Sta 102-210 | 2012-06-21 13:43 | 24C | 0.89883 | 4.2256e-06 | -4.2975e-05 | 0.89879 | 31.0527 | 0.00016065
2012-05#8 | Sta 102-209 | 2012-06-21 13:45 | 24C | 0.90088 | 3.8543e-06 | -4.3094e-05 | 0.90084 | 31.1318 | 0.00014848
2012-05#9 | Sta 102-208 | 2012-06-21 13:47 | 24C | 0.91979 | 3.6104e-06 | -4.3035e-05 | 0.91974 | 31.8639 | 0.00013874
2012-05#10 | Sta72-236 | 2012-06-21 13:49 | 24C | 0.8917 | 3.4516e-06 | -4.3035e-05 | 0.89166 | 30.7772 | 0.00013157
2012-05#11 | Sta 72-235 | 2012-06-21 13:51 | 24C | 0.89773 | 6.5631e-06 | -4.2975e-05 | 0.89769 | 31.0102 | 0.00025268
2012-05#12 | Sta 72-234 | 2012-06-21 13:53 | 24C | 0.90978 | 3.1417e-06 | -4.3035e-05 | 0.90973 | 31.476 | 0.00012104
2012-05#13 | Sta 72-233 | 2012-06-21 13:55 | 24C | 0.91478 | 4.3221e-06 | -4.2975e-05 | 0.91474 | 31.6699 | 0.00016682
2012-05#14 | Sta 72-232 | 2012-06-21 13:57 | 24C | 0.92764 | 3.5301e-06 | -4.3154e-05 | 0.9276 | 32.1689 | 0.00013763
2012-05#15 | Sta 72-231 | 2012-06-21 14:00 | 24C | 0.94769 | 1.1199e-06 | -4.2975e-05 | 0.94765 | 32.9491 | 4.4364e-05
2012-05#16 | Sta 72-230 | 2012-06-21 14:02 | 24C | 0.97178 | 6.5729e-06 | -4.2975e-05 | 0.97174 | 33.8906 | 0.0002579
2012-05#17 | Sta 72-229 | 2012-06-21 14:04 | 24C | 0.97165 | 2.7425e-06 | -4.3035e-05 | 0.97161 | 33.8853 | 0.00010656
2012-05#18 | Sta 65-263 | 2012-06-21 14:06 | 24C | 0.90744 | 3.3591e-06 | -4.3035e-05 | 0.9074 | 31.3856 | 0.00013214
2012-05#19 | Sta 65-261 | 2012-06-21 14:08 | 24C | 0.9437 | 4.3336e-06 | -4.3035e-05 | 0.94365 | 32.7934 | 0.00017344
2012-05#20 | Sta 65-262 | 2012-06-21 14:10 | 24C | 0.92347 | 4.5985e-06 | -4.2975e-05 | 0.92343 | 32.0068 | 0.00017986
2012-05#21 | Sta 65-260 | 2012-06-21 14:12 | 24C | 0.95987 | 3.4495e-06 | -4.3035e-05 | 0.95983 | 33.4245 | 0.00013655
2012-05#22 | ADCP-252 | 2012-06-21 14:15 | 24C | 0.94947 | 3.8619e-06 | -4.3035e-05 | 0.94943 | 33.0186 | 0.00015041
2012-05#23 | Sta 69-240 | 2012-06-21 14:17 | 24C | 0.97111 | 7.2858e-06 | -4.3035e-05 | 0.97107 | 33.8643 | 0.00028361
2012-05#24 | Sta 72-238 | 2012-06-21 14:19 | 24C | 0.88946 | 4.9398e-06 | -4.3035e-05 | 0.88941 | 30.6907 | 0.00019284
2012-05#25 | Sta 72-237 | 2012-06-21 14:21 | 24C | 0.89051 | 4.8368e-06 | -4.2975e-05 | 0.89046 | 30.7311 | 0.00018769
2012-05#26 | Sta 42-180 | 2012-06-22 08:53 | 24C | 0.85668 | 3.9464e-06 | -4.2975e-05 | 0.85664 | 29.4302 | 0.00015308
2012-05#27 | Sta 42-179 | 2012-06-22 08:56 | 24C | 0.85672 | 2.4747e-06 | -4.2975e-05 | 0.85668 | 29.4317 | 9.4885e-05
2012-05#28 | Sta 42-178 | 2012-06-22 09:00 | 24C | 0.85844 | 5.3443e-06 | -4.2915e-05 | 0.85839 | 29.4976 | 0.00020321
2012-05#29 | Sta 42-176 | 2012-06-22 09:03 | 24C | 0.87209 | 3.3595e-06 | -4.3035e-05 | 0.87205 | 30.0219 | 0.00012705
2012-05#30 | Sta42-177 | 2012-06-22 09:06 | 24C | 0.86575 | 3.8747e-06 | -4.2915e-05 | 0.86571 | 29.7784 | 0.00014713
2012-05#31 | Sta42-175 | 2012-06-22 09:26 | 24C | 0.87644 | 2.2391e-06 | -4.2915e-05 | 0.8764 | 30.1892 | 8.5299e-05
2012-05#32 | Sta42-174 | 2012-06-22 09:29 | 24C | 0.88696 | 2.4744e-06 | -4.3035e-05 | 0.88692 | 30.5944 | 9.4885e-05
2012-05#33 | Sta42-173 | 2012-06-22 09:31 | 24C | 0.88487 | 4.9693e-06 | -4.2915e-05 | 0.88482 | 30.5137 | 0.00019375
2012-05#34 | Sta102-199 | 2012-06-22 09:33 | 24C | 0.97267 | 3.4551e-06 | -4.2915e-05 | 0.97263 | 33.9254 | 0.00013134
2012-05#35 | Sta102-198 | 2012-06-22 09:36 | 24C | 0.97348 | 3.2147e-06 | -4.3035e-05 | 0.97344 | 33.957 | 0.00012429
2012-05#36 | Sta102-187 | 2012-06-22 09:38 | 24C | 0.89064 | 4.8375e-06 | -4.3154e-05 | 0.89059 | 30.7362 | 0.00018329
2012-05#37 | Sta42-185 | 2012-06-22 09:40 | 24C | 0.20402 | 6.4355e-06 | -4.3035e-05 | 0.20398 | 6.1316 | 0.0002085
2012-05#38 | Sta42-184 | 2012-06-22 09:44 | 24C | 0.76472 | 5.3433e-06 | -4.3035e-05 | 0.76468 | 25.9355 | 0.00020329
2012-05#39 | Sta42-183 | 2012-06-22 09:47 | 24C | 0.83434 | 4.8104e-06 | -4.2915e-05 | 0.8343 | 28.5756 | 0.0001834
2012-05#40 | Sta42-182 | 2012-06-22 09:50 | 24C | 0.85087 | 3.4657e-06 | -4.3035e-05 | 0.85083 | 29.2076 | 0.00013637
2012-05#41 | Test | 2012-06-22 09:54 | 24C | 0.90119 | 1.8257e-05 | -4.2915e-05 | 0.90114 | 31.1437 | 0.00070721 | Peter Chandler hydrometer standardization
2012-05#42 | Stn41-181 | 2012-06-22 09:58 | 24C | 0.85624 | 1.3341e-05 | -4.2975e-05 | 0.85619 | 29.4132 | 0.00051305
2012-05#43 | Stn102-207 | 2012-06-22 10:00 | 24C | 0.92902 | 1.6626e-05 | -4.2915e-05 | 0.92898 | 32.2225 | 0.00064505
2012-05#44 | Sta102-206 | 2012-06-22 10:04 | 24C | 0.93236 | 1.5236e-05 | -4.2915e-05 | 0.93232 | 32.3523 | 0.00059097
2012-05#45 | P24-DWR | 2012-06-22 11:50 | 24C | 0.98946 | 8.2179e-06 | -4.2915e-05 | 0.98942 | 34.5839 | 0.0003216
2012-05#46 | Stn102-205 | 2012-06-22 11:53 | 24C | 0.94486 | 1.0154e-05 | -4.2796e-05 | 0.94481 | 32.8387 | 0.00039621
2012-05#47 | Sta102-204 | 2012-06-22 11:56 | 24C | 0.95828 | 1.1603e-05 | -4.2915e-05 | 0.95824 | 33.3624 | 0.00045334
2012-05#48 | Sta102-203 | 2012-06-22 11:58 | 24C | 0.96371 | 1.1895e-05 | -4.2915e-05 | 0.96367 | 33.5747 | 0.00046415
2012-05#49 | Sta102-201 | 2012-06-22 12:00 | 24C | 0.97026 | 1.2732e-05 | -4.2915e-05 | 0.97022 | 33.8309 | 0.00049484
2012-05#50 | Sta102-202 | 2012-06-22 12:02 | 24C | 0.96867 | 1.1286e-05 | -4.2856e-05 | 0.96863 | 33.7687 | 0.000443
2012-05#51 | Sta102-200 | 2012-06-22 12:04 | 24C | 0.97154 | 1.059e-05 | -4.2915e-05 | 0.9715 | 33.8811 | 0.00041717
2012-05#52 | Sta27-136 | 2012-06-22 13:15 | 24C | 0.89289 | 1.089e-05 | -4.2915e-05 | 0.89285 | 30.8232 | 0.00041927
2012-05#53 | Sta27-137 | 2012-06-22 13:18 | 24C | 0.89212 | 2.2401e-06 | -4.2915e-05 | 0.89208 | 30.7936 | 8.5313e-05
2012-05#54 | Sta27-138 | 2012-06-22 13:21 | 24C | 0.89086 | 1.2971e-05 | -4.2915e-05 | 0.89082 | 30.7448 | 0.0005022
2012-05#55 | Sta27-139 | 2012-06-22 13:22 | 24C | 0.88812 | 1.0349e-05 | -4.2915e-05 | 0.88808 | 30.6392 | 0.00039779
2012-05#56 | Sta2-122 | 2012-06-22 13:25 | 24C | 0.88872 | 7.7727e-06 | -4.2915e-05 | 0.88868 | 30.6623 | 0.00030321
2012-05#57 | Sta3-106 | 2012-06-22 13:27 | 24C | 0.89164 | 6.8255e-06 | -4.2796e-05 | 0.8916 | 30.7751 | 0.00026198
2012-05#58 | Sta6-93 | 2012-06-22 13:30 | 24C | 0.88367 | 1.4043e-05 | -4.2915e-05 | 0.88362 | 30.4675 | 0.00054103
2012-05#59 | Sta9-81 | 2012-06-22 13:32 | 24C | 0.88396 | 1.0341e-05 | -4.2915e-05 | 0.88391 | 30.4787 | 0.00039641
2012-05#60 | Sta27-147 | 2012-06-22 13:34 | 24C | 0.85891 | 9.4813e-06 | -4.2915e-05 | 0.85886 | 29.5157 | 0.00036562 | Cleaned cell with CLR in the hopes of reducing the erratic readings
Calibration#2 | P150std K0.9978 | 2012-06-22 14:21 | 24C | 0.99978 | 2.7426e-06 | -3.0398e-06 | 0.99978
2012-05#61 | P150std K0.9978 | 2012-06-22 14:23 | 24C | 0.99977 | 7.2639e-06 | -3.0398e-06 | 0.99977 | 34.9908 | 0.00028424 | Ran standard to see what salinity value was calculated. Salinity value on std bottle 34.991
2012-05#62 | Sta27-146 | 2012-06-22 14:27 | 24C | 0.86135 | 1.2557e-05 | -3.0994e-06 | 0.86135 | 29.6109 | 0.00048169
2012-05#63 | Sta27-144 | 2012-06-22 14:29 | 24C | 0.86791 | 8.4303e-06 | -2.9802e-06 | 0.86791 | 29.8627 | 0.00032325
2012-05#64 | Sta27-143 | 2012-06-22 14:31 | 24C | 0.87258 | 8.6767e-06 | -3.0398e-06 | 0.87258 | 30.0424 | 0.00032921
2012-05#65 | Sta27-140 | 2012-06-22 14:33 | 24C | 0.88547 | 1.122e-05 | -2.9802e-06 | 0.88547 | 30.5384 | 0.00043195
2012-05#66 | Sta27-142 | 2012-06-22 14:40 | 24C | 0.87617 | 1.1033e-05 | -3.0398e-06 | 0.87617 | 30.1803 | 0.0004249
2012-05#67 | Sta27-141 | 2012-06-22 14:42 | 24C | 0.88288 | 7.3032e-06 | -2.9206e-06 | 0.88288 | 30.4386 | 0.00027802
2012-05#68 | Sta42-172 | 2012-06-22 14:44 | 24C | 0.8876 | 8.054e-06 | -3.0994e-06 | 0.8876 | 30.6208 | 0.00031247
2012-05#69 | Sta42-169 | 2012-06-22 14:46 | 24C | 0.89078 | 9.394e-06 | -2.9802e-06 | 0.89078 | 30.7434 | 0.00036425
2012-05#70 | Sta42-170 | 2012-06-22 14:48 | 24C | 0.89075 | 9.2917e-06 | -2.9802e-06 | 0.89075 | 30.7421 | 0.00035775
2012-05#71 | Sta42-171 | 2012-06-22 14:50 | 24C | 0.89003 | 1.372e-06 | -2.9802e-06 | 0.89002 | 30.7142 | 5.2246e-05 | Sta42-171
2012-05#72 | Sta39-153 | 2012-06-22 14:53 | 24C | 0.89236 | 7.5885e-06 | -2.9802e-06 | 0.89236 | 30.8043 | 0.00029127
2012-05#73 | Sta27-148 | 2012-06-22 14:56 | 24C | 0.85006 | 2.7425e-06 | -2.9206e-06 | 0.85006 | 29.1782 | 0.0001097
2012-05#74 | Sta27-149 | 2012-06-22 14:58 | 24C | 0.78329 | 8.4963e-06 | -2.9802e-06 | 0.78328 | 26.6374 | 0.00032103
2012-05#75 | Sta27-150 | 2012-06-22 15:01 | 24C | 0.67248 | 6.7855e-06 | -3.0398e-06 | 0.67248 | 22.4941 | 0.0002516
2012-05#76 | Sta12-75 | 2012-06-22 15:03 | 24C | 0.85202 | 8.252e-06 | -2.9802e-06 | 0.85201 | 29.2531 | 0.00031617
2012-05#77 | Sta12-76 | 2012-06-22 15:06 | 24C | 0.83649 | 8.4361e-06 | -2.9206e-06 | 0.83649 | 28.6592 | 0.00032096
2012-05#78 | Sta12-74 | 2012-06-22 15:08 | 24C | 0.8583 | 9.2117e-06 | -3.0398e-06 | 0.85829 | 29.4841 | 0.00035556
2012-05#79 | Sta12-73 | 2012-06-22 15:10 | 24C | 0.86221 | 2.2258e-06 | -3.0398e-06 | 0.8622 | 29.6438 | 8.5943e-05
2012-05#80 | Sta12-72 | 2012-06-22 15:12 | 24C | 0.86777 | 1.362e-05 | -3.0398e-06 | 0.86776 | 29.8572 | 0.00052627
2012-05#81 | Sta12-71 | 2012-06-22 15:14 | 24C | 0.87215 | 1.0861e-05 | -3.0994e-06 | 0.87214 | 30.0257 | 0.00041704
2012-05#82 | Sta12-79 | 2012-06-22 15:16 | 24C | 0.71644 | 2.4908e-06 | -2.9802e-06 | 0.71643 | 24.1267 | 9.2752e-05
2012-05#83 | Sta12-78 | 2012-06-22 15:18 | 24C | 0.74636 | 1.2921e-05 | -3.0398e-06 | 0.74636 | 25.2466 | 0.00048579
2012-05#84 | Sta12-65 | 2012-06-22 15:20 | 24C | 0.88518 | 1.4623e-05 | -2.9206e-06 | 0.88518 | 30.5273 | 0.00056388
2012-05#85 | Sta12-64 | 2012-06-22 15:23 | 24C | 0.886 | 1.1266e-05 | -3.0398e-06 | 0.886 | 30.559 | 0.00043418
2012-05#86 | Sta12-66 | 2012-06-22 15:25 | 24C | 0.88353 | 1.0312e-05 | -2.9802e-06 | 0.88353 | 30.4637 | 0.00039633
2012-05#87 | Sta12-68 | 2012-06-22 15:44 | 24C | 0.88172 | 1.3754e-05 | -3.0398e-06 | 0.88172 | 30.3941 | 0.00052881
2012-05#88 | Sta12-67 | 2012-06-22 15:47 | 24C | 0.88272 | 1.0967e-05 | -3.0994e-06 | 0.88272 | 30.4326 | 0.00042292
2012-05#89 | Sta12-69 | 2012-06-22 15:49 | 24C | 0.87928 | 1.321e-05 | -2.9802e-06 | 0.87928 | 30.3002 | 0.00050908
2012-05#90 | Sta12-70 | 2012-06-22 15:51 | 24C | 0.87604 | 5.4475e-06 | -2.9802e-06 | 0.87604 | 30.1753 | 0.0002091
2012-05#91 | Sta12-77 | 2012-06-22 15:54 | 24C | 0.80131 | 7.0799e-06 | -3.0398e-06 | 0.8013 | 27.3199 | 0.00026844
2012-05#92 | Sta14-47 | 2012-06-22 15:56 | 24C | 0.85066 | 6.5372e-06 | -3.0398e-06 | 0.85066 | 29.2013 | 0.00025104
2012-05#93 | Sta14-46 | 2012-06-22 15:58 | 24C | 0.85413 | 8.6316e-06 | -3.0398e-06 | 0.85412 | 29.3338 | 0.00033154
2012-05#94 | Sta14-49 | 2012-06-22 16:01 | 24C | 0.8433 | 6.1735e-06 | -2.9802e-06 | 0.84329 | 28.9193 | 0.00023894
2012-05#95 | Sta14-48 | 2012-06-22 16:04 | 24C | 0.84895 | 1.1514e-05 | -3.0398e-06 | 0.84895 | 29.1356 | 0.000443
2012-05#96 | Sta14-50 | 2012-06-22 16:06 | 24C | 0.81672 | 6.311e-06 | -2.9802e-06 | 0.81671 | 27.9054 | 0.00023915
2012-05#97 | Sta14-54 | 2012-06-22 16:08 | 24C | 0.88071 | 1.0789e-05 | -2.9802e-06 | 0.8807 | 30.355 | 0.00041391
2012-05#98 | Sta14-52 | 2012-06-22 16:10 | 24C | 0.76128 | 1.0334e-05 | -3.0994e-06 | 0.76128 | 25.8072 | 0.0003895
2012-05#99 | P24 DWR Bot9/10 320/321 | 2012-06-22 16:19 | 24C | 0.98947 | 1.3264e-05 | -2.9802e-06 | 0.98946 | 34.5857 | 0.00052196
2012-05#100 | P24 DWR Bot9/10 320/321 | 2012-06-25 11:06 | 24C | 0.98946 | 4.9696e-06 | -2.9206e-06 | 0.98946 | 34.5853 | 0.00019581
2012-05#101 | Sta14-44 | 2012-06-25 11:09 | 24C | 0.87382 | 5.5671e-06 | -2.9802e-06 | 0.87382 | 30.0901 | 0.00021551
2012-05#102 | Sta14-45 | 2012-06-25 11:11 | 24C | 0.86796 | 4.322e-06 | -3.0398e-06 | 0.86795 | 29.8646 | 0.00016668
2012-05#103 | Sta14-43 | 2012-06-25 11:13 | 24C | 0.87839 | 5.3479e-06 | -3.0398e-06 | 0.87839 | 30.2659 | 0.00020399
2012-05#104 | Sta14-42 | 2012-06-25 11:15 | 24C | 0.87999 | 4.949e-06 | -3.0398e-06 | 0.87999 | 30.3275 | 0.00019108
2012-05#105 | Sta14-41 | 2012-06-25 11:17 | 24C | 0.88145 | 5.1607e-06 | -3.0398e-06 | 0.88145 | 30.3838 | 0.00019536
2012-05#106 | Sta14-40 | 2012-06-25 11:20 | 24C | 0.88285 | 4.4783e-06 | -3.0994e-06 | 0.88285 | 30.4376 | 0.00017423
2012-05#107 | Sta14-39 | 2012-06-25 11:23 | 24C | 0.88373 | 3.865e-06 | -3.0398e-06 | 0.88373 | 30.4714 | 0.00014655
2012-05#108 | Sta14-38 | 2012-06-25 11:25 | 24C | 0.88477 | 5.2797e-06 | -2.9802e-06 | 0.88477 | 30.5116 | 0.00020323
2012-05#109 | Sta56-25 | 2012-06-25 11:27 | 24C | 0.9179 | 5.0205e-06 | -3.0398e-06 | 0.9179 | 31.7925 | 0.00019469
2012-05#110 | Sta59-23 | 2012-06-25 11:29 | 24C | 0.8475 | 3.85e-06 | -2.9802e-06 | 0.8475 | 29.0803 | 0.00014604
2012-05#111 | Sta59-22 | 2012-06-25 11:31 | 24C | 0.84808 | 4.952e-06 | -2.9802e-06 | 0.84807 | 29.1022 | 0.00018882
2012-05#112 | Sta59-21 | 2012-06-25 11:33 | 24C | 0.8604 | 4.2092e-06 | -2.9802e-06 | 0.8604 | 29.5745 | 0.00016147
2012-05#113 | Sta59-20 | 2012-06-25 11:35 | 24C | 0.87313 | 3.8543e-06 | -2.9802e-06 | 0.87312 | 30.0633 | 0.00015089
2012-05#114 | Sta59-19 | 2012-06-25 11:38 | 24C | 0.87406 | 3.4494e-06 | -3.0398e-06 | 0.87406 | 30.0992 | 0.0001343
2012-05#115 | Sta59-18 | 2012-06-25 11:42 | 24C | 0.87436 | 4.3699e-06 | -3.0994e-06 | 0.87435 | 30.1106 | 0.00017031
2012-05#116 | Sta59-17 | 2012-06-25 11:44 | 24C | 0.88078 | 5.8621e-06 | -2.9802e-06 | 0.88078 | 30.358 | 0.00022744
2012-05#117 | Sta59-16 | 2012-06-25 11:47 | 24C | 0.88714 | 3.6105e-06 | -2.9802e-06 | 0.88714 | 30.6031 | 0.00014202
2012-05#118 | Sta59-15 | 2012-06-25 11:49 | 24C | 0.89683 | 4.837e-06 | -3.0994e-06 | 0.89683 | 30.9768 | 0.00018563
2012-05#119 | Sta59-14 | 2012-06-25 11:51 | 24C | 0.91596 | 3.862e-06 | -2.9802e-06 | 0.91596 | 31.7172 | 0.00014806
2012-05#120 | Sta59-13 | 2012-06-25 11:53 | 24C | 0.92253 | 5.007e-06 | -3.0994e-06 | 0.92253 | 31.972 | 0.00019031
2012-05#121 | Sta59-12 | 2012-06-25 11:55 | 24C | 0.92956 | 3.4657e-06 | -2.9802e-06 | 0.92955 | 32.2448 | 0.00013386
2012-05#122 | Sta59-11 | 2012-06-25 11:57 | 24C | 0.93099 | 4.5983e-06 | -2.9802e-06 | 0.93099 | 32.3006 | 0.00017701
2012-05#123 | Sta59-10 | 2012-06-25 11:59 | 24C | 0.93216 | 2.4754e-06 | -2.9802e-06 | 0.93216 | 32.3461 | 9.6423e-05
2012-05#124 | StaSInlet-01 | 2012-06-25 12:01 | 24C | 0.88731 | 4.6106e-06 | -3.0994e-06 | 0.88731 | 30.6094 | 0.00017741
2012-05#125 | Sta65-264 | 2012-06-25 12:04 | 24C | 0.90409 | 4.1735e-06 | -3.0398e-06 | 0.90408 | 31.2573 | 0.00016082
2012-05#126 | Sta65-265 | 2012-06-25 12:07 | 24C | 0.89126 | 1.1199e-06 | -2.9802e-06 | 0.89126 | 30.7617 | 4.265e-05
2012-05#127 | Sta65-266 | 2012-06-25 12:10 | 24C | 0.87503 | 5.8674e-06 | -2.9206e-06 | 0.87503 | 30.1367 | 0.00022328 | Lost communication between Autosal and computer. Re-established communication and continued analysis
2012-05#128 | Sta65-267 | 2012-06-25 12:13 | 24C | 0.86738 | 6.425e-06 | -2.861e-06 | 0.86737 | 29.8423 | 0.00024802
2012-05#129 | Sta65-268 | 2012-06-25 12:15 | 24C | 0.86646 | 3.8607e-06 | -2.861e-06 | 0.86646 | 29.8072 | 0.00014954
2012-05#130 | Sta65-269 | 2012-06-25 12:18 | 24C | 0.86615 | 5.0134e-06 | -2.861e-06 | 0.86615 | 29.7953 | 0.00019111
2012-05#131 | Sta62-271 | 2012-06-25 12:20 | 24C | 0.9422 | 2.2394e-06 | -2.861e-06 | 0.94219 | 32.7366 | 8.7102e-05
2012-05#132 | Sta14-51 | 2012-06-25 12:23 | 24C | 0.78062 | 4.7302e-06 | -2.861e-06 | 0.78062 | 26.5369 | 0.0001764