## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2012-09

Agency: OSD

Location: North Pacific	/ Bering Sea / Chukchi Sea

Project: C3O/IPY

Party Chief: Vagle S.

Platform: Sir Wilfrid Laurier

Date: July 6, 2012 - July 22, 2012

Processed by: Germaine Gatien

Date of Processing: 17 June 2014 – 22 March 2018

Number of original HEX files: 26

Number of CTD files:  26	Number of bottle casts:	26

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0941) was used for this cruise. It was mounted in a custom-built compact 24-bottle rosette sampler and attached were a Wetlabs CSTAR transmissometer (#993DR), an SBE 43 DO sensor (#1117), a SeaPoint Fluorometer (#2745), a Biospherical QSP-200L4S PAR sensor (#70123) and an altimeter (#40853).

24 OceanTest Equipment 10L bottles were mounted on the rosette.

A thermosalinograph (SeaBird 21 S/N 3274) was mounted with a fluorometer (SCF2859) and a remote temperature sensor #0271. The fluorometer cable was changed several times. These data are not discussed in this report as they were processed by other personnel.

The data logging computer was W9B6IOS101V55.

The data acquisition program was Seasave 7.20c.

The deck unit was a Seabird model 11, serial # 11P53201-0800; it included a NMEA board to automatically add GPS positions into the header of the data files.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572. Salinity samples were analyzed on 14 and 16 November 2012.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette log sheets as well as spreadsheets detailing sampling were provided.

The CTD was dropped to about 7m for a soak, returned to the surface; then, after a stop of 5s to 15s, a full cast was run. The practice of acquiring data during the initial drop and return to the surface makes processing awkward. If acquisition began at the start of the full cast it would make processing simpler. It is also recommended that the stop at the surface before the full cast be longer, at least 1 minute, to allow the near-surface water to recover from the effects of the raising of the rosette package.

Salinity samples were analyzed 4 months after the end of the cruise. The comparison with the CTD suggests poor flushing of Niskin bottles and/or evaporation of samples; both are likely. A post-cruise calibration indicates that the two conductivity sensors drifted minimally while secondary temperature drifted enough to lead to salinity being slightly high, ~0.0005.

There was no dissolved oxygen calibration sampling. A post-cruise calibration suggests that the sensor was reading low by about 2% in November 2012 but it is unknown how much of that change had occurred at the time of this cruise.

An error in the PAR calibration parameters was discovered after most processing had been done; it was corrected.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Files had non-standard names, but fixing them is simple. This will be done after conversion.

##### Preliminary Steps

The Daily Science Log Book and rosette log sheets were obtained. A cruise report was available.

Spreadsheet 2012-09_SWL_Chem.xls contains a sampling log with details on bottles fired and contains results of salinity analysis. Nutrients analyzed at IOS were available in another spreadsheet; there were flag channels but there were no entries despite comments that might justify flagging. There were duplicates taken but the first value was selected by the analyst, though in some cases it looks like the 2nd was more likely correct. A duplicate analysis is included. DIC and Alkalinity were available in a separate spreadsheet; there were no flag channels.

Other bottle data (nutrients, ammonium, O18 and chlorophyll) were analyzed at the University of Maryland and were delivered in a separate spreadsheet.

The cruise summary sheet was completed.

The pressure sensor, conductivity and DO sensors were only used for this cruise since the last factory calibration.

The calibration constants were checked for all instruments. Most sensors had been recalibrated in late 2012 and there were post-cruise calibrations from late 2013 for the temperature, conductivity and dissolved oxygen sensors. There was an error in either the serial number or the parameters for the PAR sensor. It was confirmed by the CTD technician that the serial number was correct, so the parameters are wrong. The correct ones were entered in file 2012-09-ctd.xmlcon.

The pressure offset differed from that in the last factory recalibration, but was deliberately changed at sea, so the adjusted setting will be used.

There was no deep sampling, so no hysteresis tests were done.

Voltage channels were included for Transmissivity, SeaPoint Fluorescence, Dissolved Oxygen, altimeter and PAR channels.

There was a drop to about 7m before a return to the surface and then casts were run. The acquisition system was on during the soak period so data from those sections will have to be removed before running DELETE.

##### BOTTLE FILE PREPARATION

The ROS files were converted in the same way as the full files except that bottle number and bottle position channels were included and oxygen concentration and salinity were derived.

The files were then converted to IOS Header files with BOT extensions.

The BOT files were averaged on bottle number and those files were used to prepare an ADDSAMP. That file was edited to add sample numbers. On two occasions a single sample number (#77 and 82) was assigned to 3 Niskin bottles fired at the same depth. Niskin bottles #2 and 3 were sampled just for water, so this should not cause any confusion.

Temperature and salinity were plotted for all BOT files. There were a few spikes in the secondary salinity. CTDEDIT was used to remove such spikes from casts #19 and 22. The edited files were copied to BOT.

A preliminary header check turned up no problems. There was no off-scale fluorescence.

The BOT files were averaged on bottle number to enable preparation of file ADDSAMP.csv. That file was edited to add sample numbers and was then used to add sample numbers to the BOT files, creating SAM files. Those were bin-averaged on bottle number to create SAMAVG files.

Next, text file 2012-09-bot-hdr.txt was prepared to add an explanation of quality flags and some general comments from analysts.

Spreadsheet “2012-09_SWL_Chemistry.xls” contained final nutrient data and salinity bottle data, but the latter appeared to be corrupted. The file was simplified and saved as “2012-09-bottle data.csv” in preparation for combining with the CTD data from the rosette files.

The salinity bottle data was found in file “Salts 2012-09.xls”. There were no flags though there were duplicates and some comments. The duplicates were averaged and flag 6 added and an initial flag of 2 was added to the cases where leaky liners were reported. This should be reviewed after COMPARE is run – a 3 flag may be appropriate. DIC and Alkalinity data were delivered in a separate spreadsheet, “2012-09 DIC and ALK data.xls”; this contained duplicates and no final value or flags, so the duplicates were averaged and flag 6 added. That information was then copied to file “2012-09-bottle data .csv”.

File 2012-09-bottle data.csv was then converted to individual MRG1 files.

Next a spreadsheet containing the University of Maryland data (Nutrients, Ammonium, CHL and O18) were simplified and saved as file 2012-09_UMCES_samples.csv. One bad silicate2 value was replaced with a pad value and comment. That file was converted to MRG2 files, which were then merged with the MRG1 files, with output *.MRG.

These files were put through CLEAN to remove SeaBird headers and comments from the secondary file. They were then put through SORT to arrange the files in bottle # order and then merged with the SAMAVG files with output MRG.

CLEAN was run to add 0s to empty flag channels and to remove channels with all pad values.

The altimeter readings and bottom depths from the headers of the CLN files were exported to a spreadsheet.

Since the bottom depths were missing for many casts and others differed from the entries in the log and the spreadsheet provided with depths, file 2012-09_header_info.csv was prepared with file names and bottom depths. That file was used with routine “MERGE: csv file to Headers” to produce MRGMRH files.

The altimeter readings were then exported again along with bottom depths and maximum pressure for the cast. To check that the altimetry header readings are reasonable the following calculation was done:

Check Value = Maximum Pressure + Altimeter – Bottom Depth

We do not expect 0 values due to pressure not being equal to depth and due to altimetry header being averaged over the bottom 2m and due to variations in water depth. Where values are >4 further investigation is made.

For cast #20 the value was 4.6. A plot of the cast shows the altimetry trace was clear and the header value a good estimate, the ship may have drifted into slightly shallower water during the cast. All other check values were <2.

The bottle data were exported to a spreadsheet. Checks were made that all the samples indicated on the rosette logs were present; all samples that were analyzed at IOS were present with the exception of the samples from cast #26 which the analyst noted could not be found. It is presumed that they were not taken.

##### Compare

Salinity

Compare was run with pressure as reference channel. There was a lot of scatter in the plot of differences between salinity from bottles and CTD versus CTD salinity. This is not a surprise in such shallow sampling. The CTD salinity reads lower than the bottles near the surface which is likely due to incomplete flushing in a high vertical gradient zone. When bottles are excluded from above 10m and any bottle for which the CTD salinity standard deviation is >0.001, the primary salinity is found to be low by an average of 0.0069 and the secondary low by 0.0060; the standard deviation in the fit is 0.01 for both.

There were notes about 3 bottles having leaky liners, so those were investigated to see if they were notable outliers. They were not out of line, so the flags were left as 2, and the comment was adjusted to note the COMPARE results.

The only outliers are associated with high standard deviations or were very close to the surface, so no flags are recommended for the bottle data. The COMPARE results do not imply problems with sampling protocols or analysis, but rather, likely reflect that the water drawn is likely to have come from deeper in the water column so that salinity doesn’t match ambient values at the bottle closing level.

For full details for the COMPARE run see file 2012-12-sal-comp1.xls.

##### Conversion of Full files from Raw Data

The file names were changed to standard format before conversion.

All files were converted using 2012-09-ctd.xmlcon. The hysteresis correction was not selected since all casts were shallow; the Tau correction will be used.

All channels were plotted for a few casts to check for problems in the conversion.  All casts were shallow so variability is high throughout. The temperature and conductivity channels have sections where they are close, but generally there are significant differences and the downcasts seem as bad as the upcasts from that point of view. The altimetry looks useful. PAR and SPAR signals have similar ranges and look normal. The dissolved oxygen has the usual offset, but the high temperature gradients in a shallow cast will challenge the instrument, and we can expect a problem aligning the data, especially since there are frequent stops for all casts so the upcast cannot be expected to match the downcast. The transmissivity and fluorescence look reasonable. Descent rates are as high and steady as can be expected from shallow casts.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

There are so many stops for bottles and a lot of variability in temperature and dissolved oxygen, so the usual tests are not very helpful in deciding how to align the data, except to indicate that some shift is needed. The value used for 2010-05 was applied since this was the same sensor as used then. The same sensor was also used for 2011-18 when a lower shift was used but the sensor was found to need repairs after that cruise.

ALIGNCTD was used to advance the DO Voltage by 5s relative to the pressure.

##### CELLTM

As for ALIGNCTD tests are not helpful for these shallow casts with so many stops and high variability, so settings were used that are always found reasonable, and often the best choice.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration (using the Tau correction). The usual step of deriving differences between channels for a few test casts was skipped as they are not useful in such shallow water.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.

- There are some very high SBE dissolved oxygen values during casts #4 and 5 and many spikes in the DO. The largest values are seen in the upcast at the level of a bottle stop, but the spiky data were not included in the 10s window used to create the bottle files. The downcast values are ~12mL/L at the levels where the upcast spikes go to 16mL/L.

- Fluorescence went off scale during cast #1 reaching a value of 50ug/L; this occurred as a spike during the stop at the bottom so should be removed with normal processing steps. There were no other casts with off-scale fluorescence.

- Speeds look reasonable.

- No other problems were noted.

Surface check was run and shows an average surface pressure for the cruise was 0.09db but the associated salinity values were mostly close to 0.

There is no clear evidence of whether the pressure is correct or not because the pumps were off when the lowest pressures were recorded, but there are some low salinity values associated with pressures of ~0.2db. While the pumps were off, the two conductivity channels were close to each other, so probably reflect ambient conditions. This shows that the pressures are likely reliable.

The cross-reference check was compared with the log book and no problems were found. The times and positions were slightly different but this is likely because the log entries were made at the beginning of the full cast, rather than the beginning of the file.  Cast #1 was checked and the pre-cast soak period lasted about 11min and the difference between the header and log times was 8min.

The altimeter readings and bottom depths from the headers of the CLN files were exported to a spreadsheet.

As was explained in section 4, file 2012-09_header_info.csv  was used with routine “MERGE: csv file to Headers” to add bottom depths to the MRG files. (Output: MRH files)

The altimeter readings were then exported again along with bottom depths and maximum pressure for the cast. To check that the altimetry header readings are reasonable the following calculation was done:

Check Value = Maximum Pressure + Altimeter – Bottom Depth

As was found for the bottle files only cast #20 has a check value >2. The altimetry signal was good and the header value a good estimate, so it is likely that the ship drifted into slightly shallower water during the cast.

The cruise tracks (with event #s and station names) were plotted and added to the end of this report. No problems were found.

At this point the early records before the full cast were removed because DELETE will patch the initial drop to 7m to the full cast from 7m to the bottom.

A graphical examination of the files was used to make an estimate of the first record # to be included and CLIP was used to remove all records before that. Plots were made after CLIP and some settings were fine-tuned until enough data are removed to ensure that DELETE will select the appropriate data.

##### Shift

The casts are too shallow, have too many stops for bottles and data too variable to allow meaningful testing of alignment for conductivity and fluorescence. The same sensors were used for 2010-05 and 2011-18, so the settings used in alignment for those cruises were applied for this cruise. Plots were examined before and after and the alignment was improved by these steps.

Fluorescence

SHIFT was run on all casts to advance the SeaPoint fluorescence channel by +24 records.

Conductivity

SHIFT was run applying an advance of -0.4 records to the primary conductivity and -0.6 records to the secondary conductivity.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. There is no obvious further adjustment that would improve the comparison between downcast and upcast traces.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was to note that cast #26 is <20m deep which is correct.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The sensors were both recalibrated in December 2011 and there were post-cruise calibrations in November 2012. This was the only cruise between those calibrations.

2. Dissolved Oxygen

The DO sensor was recalibrated in Nov. 2011 and a post-cruise calibration in November 2012. This cruise was the only one using this sensor between the two calibrations.

3. Pressure

The sensor was recalibrated in December 2009 and there is no post-cruise calibration available. The offset used at sea has been steadily changed from -0.9039db in 2009 and 2010 to -2.2db in 2011 and -3.4db in 2012 and -4.4db in 2013.

Historic ranges – There was no local climatology available.

Post-cruise calibrations

1. Salinity: The drifts reported were ~0 for both conductivity sensors and -0.00001C° per year for the primary temperature and -0.00054C° per year for the secondary.

2. Dissolved Oxygen: Three casts were recalibrated using the later settings and the “before” and “after” cruise results were compared. The data are very noisy due to temperature variability, and the range of values is fairly low, so a simple answer did not emerge. The post-cruise calibration did lead to higher values by about 2.5% though this varied from 2 to 5% and in one case post-cruise value was lower. Some of the drift may have occurred after the cruise.

Repeat Casts –

There were no repeat casts.

##### DETAILED EDITING

The post-cruise calibration showed insignificant drift in the primary salinity and primary T-S plots look better than the secondary, with fewer unstable features.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. The CTD was lowered to about 7db and returned to the surface. After a short stop it was then run through the full cast. The stop was only about 10 to 15s which may explain the very noisy data in the top 5m.  A longer stop is recommended. Editing could not remove all unstable features and some may be real.

All files were put through Bin-Average using 0.5db bins. Plots were made to see if further editing was required. Many unstable features remain but most are fine-scale. A few were re-examined in CTDEDIT to see if further editing would be useful but the features do not have an obvious instrumental source, so were left in case they are real.

##### Initial Recalibration

The pressure looks ok.

We have insufficient information on which to base a recalibration of dissolved oxygen.

From COMPARE, the primary salinity was found to be low by 0.0069 and the secondary by 0.0060. The post-cruise calibrations showed that the primary salinity had insignificant drift while the secondary salinity would have been high by ~0.0005C°.  The post-cruise calibration report is more reliable than the bottle comparison given the clear evidence of poor flushing. Since the primary channels were selected, no recalibration is needed.

No recalibration was applied.

##### Special Fluorometer Processing

There were no off-scale fluorescence data after editing.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 0.5	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. Many unstable features remain as described in section 15, but most are only slightly unstable and may be real. No further editing was applied.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on the AVG files to remove the following channels (Output *.REM):

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate,Voltage:0, Voltage:1, Voltage:2, Voltage:3, Voltage:4 and Flag

A second SBE DO channel (with umol/kg units) was added to both the REM and REMSTD.

REORDER was run to get the two DO channels together.

CALIBRATE was run to fix an error in the PAR calibration.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence:SeaPoint, PAR and PAR:Reference are nominal and unedited

except that some records were removed in editing temperature and salinity.

There was no calibration sampling for dissolved oxygen, so values are nominal and

are reported with 1 less significant figure than usual.

The post-cruise factory calibration indicates salinity drift was insignificant.

For further processing details see the processing report 2012-09_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and the only report was about the dissolved oxygen formats being non-standard, but that was done deliberately.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 93% to 130% with the highest values in the Chukchi Sea.

##### Final Bottle Files

The MRGCLN2 files were put through SORT to order on increasing pressure.

CALIBRATE was run to correct the PAR data as described in section 19.

REMOVE was run on the MRGSORT files to remove the following channels (Output *.MRGREM):

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate,Voltage:0, Voltage:1, Voltage:2, Voltage:3, Voltage:4 and Flag.

A second SBE DO channel was added with mass units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, correct the name of DIC, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

An acknowledgement was added that this was part of a joint cruise with the University of Maryland Center for Environmental Sciences.

A header check was run on the final files and the only problems found were formats that were purposefully changed from the standard ones.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no errors were found.

Standards check was run on all files and the only warnings concerned formats that were purposefully set to non-standard ones.

The track plot was produced on screen and no further errors were found.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0941 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5048 | 2Dec2011 | Factory | 2Nov2012 |  |
| Conductivity | 3579 | 1Dec2011 | Factory | 2Nov2012 | Factory |
| Secondary Temp. | 5073 | 2Dec2011 | Factory | 2Nov2012 |  |
| Secondary Cond. | 3581 | 1Dec2011 | Factory | 2Nov2012 | Factory |
| Transmissometer | 1050DR | 2June2012 |  |  |  |
| SBE 43 DO sensor | 1117 | 16Dec2011 | Factory | 24Dec2013 | Factory |
| SeaPoint Fluorometer | 2745 |  |  |  |  |
| PAR | 70123 | 31Aug2009 |  |  |  |
| Surface PAR | 20281 | 1Apr2007 |  |  |  |
| Pressure Sensor | 941 | 07Dec2009 | Factory |  |  |
| Altimeter | 40853 |  |  |  |  |