## comparison-1119
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
22 | 601.83 | 4.4663 | 4.4675 | 34.1861 | 34.1903 | 7.2032 | 0.6064 | 6 | 6 | 191 | 7.7645 | 0.045721 | 0.355 |  | 1.7971e-06 | 0.4499 | -0.2271
22 | 201.944 | 6.7424 | 6.7443 | 33.878 | 33.8822 | 6.8415 | 1.4724 | 10 | 10 | 195 | 6.9735 | 0.43536 | 3.036 |  | 2.6451e-05 | [5 empty] | Summary of Soc Voffset including the original values in the factory calibration when E=0.03725
15 | 601.784 | 4.6983 | 4.6992 | 34.1968 | 34.2014 | 7.1627 | 0.60203 | 6 | 6 | 122 | 7.7185 | 0.043791 | 0.338 |  | 3.7714e-05 | [7 empty] | m | b | Soc | Voffset | R2
15 | 802.22 | 4.1438 | 4.1448 | 34.3073 | 34.3119 | 7.2534 | 0.56269 | 5 | 5 | 121 | 8.0354 | 0.02601 | 0.209 |  | 4.4277e-05 | [5 empty] | Bottles used | Original | 0.4278 | -0.2186 | 0.4278 | -0.5109 |  | 2.776
46 | 601.861 | 3.7584 | 3.7594 | 34.1758 | 34.1809 | 7.328 | 0.6542 | 11 | 11 | 380 | 7.9066 | 0.067285 | 0.532 |  | 6.0795e-05 | [5 empty] | 168 | Using 2012-15 fit | 0.449 | -0.2244 | 0.449 | -0.49978 | 0.988
15 | 6.6 | 11.0321 | 11.034 | 31.7741 | 31.7754 | 6.3043 | 2.9025 | 19 | 19 | 135 | 6.244 | 1.0788 | 6.736 |  | 7.1945e-05 | [5 empty] | 158 | all except cast #1 | 0.4493 | -0.2246 | 0.4493 | -0.49989 | 0.998
15 | 11.449 | 11.0087 | 11.0105 | 31.7788 | 31.7831 | 6.3072 | 2.9001 | 18 | 18 | 134 | 6.2509 | 1.0778 | 6.737 |  | 9.2485e-05 | [5 empty] | 157 | excluding cast #1 and 1 outlier | 0.4497 | -0.2265 | 0.4497 | -0.50367 | 0.9997 | 2.9205
15 | 125.765 | 7.8334 | 7.8351 | 33.5067 | 33.5107 | 6.6894 | 1.7025 | 13 | 13 | 129 | 6.7437 | 0.53872 | 3.633 |  | 0.00013811 | [5 empty] | 149 | excluding diff >0.01 | 0.4499 | -0.2271 | 0.4499 | -0.50478 | 0.9999
3 | 5.795 | 10.5404 | 10.5453 | 31.7641 | 31.7674 | 6.3718 | 2.9244 | 8 | 8 | 155 | 6.3118 | 1.0888 | 6.872 |  | 0.00015725 | [5 empty] | 136 | excluding diff >0.005 | 0.4496 | -0.2272 | 0.4496 | -0.50534 | 1
61 | 3500.216 | 1.5245 | 1.525 | 34.6648 | 34.6712 | 7.7203 | 1.0229 | 2 | 2 | 485 | 12.3827 | 0.23291 | 2.884 | 2 | 0.00020603 | [5 empty] | 126 | excluding diff >0.004 | 0.4495 | -0.2274 | 0.4495 | -0.5059 | 1
46 | 6.198 | 8.0572 | 8.0583 | 32.7099 | 32.7145 | 6.6904 | 2.8386 | 24 | 24 | 393 | 6.6376 | 1.0498 | 6.968 | 6 | 0.00021688 | [5 empty] | 121 | excluding diff >0.0035 | 0.4496 | -0.2276 | 0.4496 | -0.50623 | 1
78 | 25.719 | 6.3269 | 6.33 | 32.7028 | 32.7071 | 6.9615 | 2.8111 | 22 | 22 | 625 | 6.9337 | 1.0374 | 7.193 | 2 | 0.00021982 | [6 empty] | used for recal of 2012-26 | 0.4483 | -0.2246 | 0.4483 | -0.501
15 | 1251.992 | 3.0982 | 3.1005 | 34.4582 | 34.4627 | 7.4332 | 0.61197 | 2 | 2 | 118 | 8.7631 | 0.048499 | 0.425 |  | 0.00027342
46 | 3502.228 | 1.5725 | 1.5729 | 34.6597 | 34.666 | 7.7112 | 0.9908 | 3 | 3 | 372 | 12.3696 | 0.21836 | 2.701 | 6 | 0.00030344
78 | 176.22 | 4.3256 | 4.3278 | 33.664 | 33.6683 | 7.253 | 1.6492 | 16 | 16 | 619 | 7.3854 | 0.51453 | 3.8 |  | 0.00035661
15 | 1251.83 | 3.1006 | 3.1012 | 34.4579 | 34.4627 | 7.4328 | 0.61179 | 3 | 3 | 119 | 8.7624 | 0.048503 | 0.425 |  | 0.00035841
29 | 1251.696 | 2.8336 | 2.8347 | 34.4376 | 34.4422 | 7.4829 | 0.59097 | 7 | 7 | 238 | 8.8256 | 0.038411 | 0.339 |  | 0.00036646
22 | 801.132 | 3.8091 | 3.8101 | 34.2922 | 34.2964 | 7.3132 | 0.56978 | 5 | 5 | 190 | 8.1044 | 0.028873 | 0.234 |  | 0.00037094
46 | 202.859 | 5.2348 | 5.2364 | 33.8345 | 33.8395 | 7.089 | 1.4209 | 15 | 15 | 384 | 7.2371 | 0.41177 | 2.98 |  | 0.00039432
3 | 27.092 | 8.4504 | 8.4525 | 32.522 | 32.5251 | 6.6395 | 2.6077 | 6 | 6 | 153 | 6.6037 | 0.94569 | 6.245 | 6 | 0.00041113
29 | 251.355 | 5.81 | 5.8109 | 33.8664 | 33.8708 | 6.992 | 1.4025 | 13 | 13 | 244 | 7.1803 | 0.40347 | 2.897 |  | 0.0004151
61 | 50.949 | 6.8953 | 6.9024 | 32.6696 | 32.6739 | 6.8719 | 2.7955 | 21 | 21 | 504 | 6.8643 | 1.0311 | 7.078 |  | 0.00053059
22 | 127.302 | 7.4203 | 7.4216 | 33.6463 | 33.6507 | 6.7463 | 1.7539 | 13 | 13 | 198 | 6.8047 | 0.56256 | 3.828 |  | 0.00056256
46 | 3686.661 | 1.5689 | 1.5692 | 34.6626 | 34.6691 | 7.7117 | 0.99288 | 1 | 1 | 370 | 12.6839 | 0.22028 | 2.794 |  | 0.00068211
61 | 3501.057 | 1.5249 | 1.5253 | 34.6648 | 34.6712 | 7.7202 | 1.0227 | 3 | 3 | 486 | 12.384 | 0.23232 | 2.877 | 2 | 0.00071476
15 | 1001.432 | 3.5549 | 3.5555 | 34.3998 | 34.4045 | 7.3533 | 0.57614 | 4 | 4 | 120 | 8.3747 | 0.032837 | 0.275 | 6 | 0.00073153
29 | 3276.45 | 1.6715 | 1.6718 | 34.6431 | 34.6486 | 7.6928 | 0.93449 | 1 | 1 | 232 | 11.9645 | 0.19407 | 2.322 |  | 0.00074638
29 | 2501.274 | 1.7508 | 1.7511 | 34.6226 | 34.6278 | 7.6785 | 0.89666 | 4 | 4 | 235 | 10.749 | 0.17555 | 1.887 |  | 0.00075605
61 | 151.564 | 5.4873 | 5.4884 | 33.7871 | 33.792 | 7.049 | 1.6581 | 17 | 17 | 500 | 7.1452 | 0.51965 | 3.713 |  | 0.0007632
29 | 3001.327 | 1.6482 | 1.6485 | 34.6426 | 34.648 | 7.6973 | 0.94797 | 2 | 2 | 233 | 11.5342 | 0.19863 | 2.291 |  | 0.00076473
78 | 76.518 | 5.1761 | 5.1779 | 32.713 | 32.7168 | 7.1523 | 2.7336 | 20 | 20 | 623 | 7.1797 | 1.0035 | 7.205 |  | 0.00077743
78 | 251.173 | 3.9262 | 3.9272 | 33.8215 | 33.8261 | 7.3157 | 1.1549 | 14 | 14 | 617 | 7.5281 | 0.29171 | 2.196 |  | 0.00079086
29 | 3001.159 | 1.6484 | 1.6487 | 34.6425 | 34.648 | 7.6973 | 0.94785 | 3 | 3 | 234 | 11.5339 | 0.19855 | 2.29 | 6 | 0.00079178 | [6 empty] | all data except 3 and 4 flags |  |  | 0 | #DIV/0!
107 | 1251.456 | 3.0697 | 3.0716 | 34.4625 | 34.4674 | 7.4382 | 0.61454 | 2 | 2 | 813 | 8.7688 | 0.048581 | 0.426 |  | 0.00080034 | [6 empty] | excl. outliers diff>0.02 |  |  | 0 | #DIV/0! |  | 0
46 | 3502.546 | 1.5723 | 1.5727 | 34.6597 | 34.6661 | 7.7112 | 0.99106 | 2 | 2 | 371 | 12.3702 | 0.21794 | 2.696 |  | 0.0008353 | [12 empty] | 0
78 | 1250.903 | 2.5836 | 2.5845 | 34.4403 | 34.4457 | 7.5293 | 0.63108 | 8 | 8 | 611 | 8.8835 | 0.055947 | 0.497 |  | 0.00087634
29 | 1001.354 | 3.2095 | 3.2106 | 34.3624 | 34.3668 | 7.4177 | 0.57121 | 8 | 8 | 239 | 8.4527 | 0.028985 | 0.245 |  | 0.00090254 | [6 empty] | excl. outliers diff>0.001 |  |  | 0 | #DIV/0!
61 | 176.828 | 5.4128 | 5.4139 | 33.8352 | 33.8399 | 7.0591 | 1.5457 | 16 | 16 | 499 | 7.1803 | 0.46739 | 3.356 |  | 0.00090495 | [12 empty] | 0
15 | 151.95 | 7.537 | 7.5381 | 33.7524 | 33.7573 | 6.7237 | 1.5211 | 12 | 12 | 128 | 6.8035 | 0.45815 | 3.117 | 6 | 0.00090912
78 | 3001.236 | 1.5711 | 1.5717 | 34.6506 | 34.6567 | 7.7119 | 0.99611 | 4 | 4 | 607 | 11.5585 | 0.2201 | 2.544 |  | 0.00095133
61 | 3001.222 | 1.5782 | 1.5787 | 34.6498 | 34.6559 | 7.7106 | 0.99087 | 4 | 4 | 487 | 11.5562 | 0.21772 | 2.516 |  | 0.00097369
46 | 2001.08 | 1.9486 | 1.9491 | 34.5822 | 34.5879 | 7.6424 | 0.79323 | 6 | 6 | 375 | 9.9929 | 0.12879 | 1.287 |  | 0.00098254 | [7 empty] | Soc | Voffset
46 | 1001.163 | 3.0572 | 3.0582 | 34.3769 | 34.382 | 7.4448 | 0.57326 | 9 | 9 | 378 | 8.4855 | 0.029816 | 0.253 |  | 0.00099416 | [6 empty] | factory | 0.4278 | -0.5109
22 | 1001.2 | 3.4014 | 3.4026 | 34.3806 | 34.3849 | 7.3819 | 0.57048 | 4 | 4 | 189 | 8.4091 | 0.030562 | 0.257 | 6 | 0.0010031 | [6 empty] | 2012-04 | 0.4442 | -0.4822
46 | 302.467 | 4.5866 | 4.5879 | 33.89 | 33.8948 | 7.1967 | 1.1016 | 13 | 13 | 382 | 7.451 | 0.26748 | 1.993 |  | 0.0010211 | [6 empty] | 2012-25 | 0.4483 | -0.501
15 | 401.671 | 5.4554 | 5.4565 | 34.0634 | 34.068 | 7.0413 | 0.76632 | 7 | 7 | 123 | 7.3804 | 0.11869 | 0.876 |  | 0.0010253 | [6 empty] | 2012-12 | 0.4494 | -0.5031
22 | 11.139 | 10.5313 | 10.5336 | 31.9837 | 31.9874 | 6.3641 | 2.8829 | 18 | 20 | 203 | 6.3086 | 1.0689 | 6.743 |  | 0.0010581
22 | 7.112 | 10.5335 | 10.5354 | 31.9829 | 31.9867 | 6.3639 | 2.8811 | 19 | 21 | 204 | 6.305 | 1.068 | 6.734 |  | 0.0010637
46 | 3002.397 | 1.5941 | 1.5945 | 34.6494 | 34.6554 | 7.7075 | 0.9795 | 4 | 4 | 373 | 11.5529 | 0.2125 | 2.455 |  | 0.001076
29 | 2000.287 | 1.9531 | 1.9537 | 34.5824 | 34.5873 | 7.6415 | 0.80823 | 5 | 5 | 236 | 9.9905 | 0.13543 | 1.353 |  | 0.0010946
78 | 3501.917 | 1.5109 | 1.5115 | 34.6659 | 34.6722 | 7.7229 | 1.033 | 3 | 3 | 606 | 12.3903 | 0.23656 | 2.931 |  | 0.0011079
78 | 101.342 | 4.7729 | 4.7794 | 32.7372 | 32.7414 | 7.2203 | 2.6722 | 19 | 19 | 622 | 7.2751 | 0.97401 | 7.086 |  | 0.0011139
61 | 401.146 | 4.1999 | 4.2012 | 34.0341 | 34.0391 | 7.2569 | 0.88082 | 12 | 12 | 495 | 7.6171 | 0.16804 | 1.28 |  | 0.0011379
78 | 2500.819 | 1.7084 | 1.7091 | 34.6234 | 34.6293 | 7.6867 | 0.91238 | 5 | 5 | 608 | 10.761 | 0.18223 | 1.961 | 6 | 0.0011471
46 | 2500.769 | 1.7311 | 1.7315 | 34.6243 | 34.6301 | 7.6822 | 0.90277 | 5 | 5 | 374 | 10.754 | 0.17789 | 1.913 |  | 0.0011696
29 | 175.892 | 6.8758 | 6.8776 | 33.8522 | 33.8567 | 6.8217 | 1.8054 | 15 | 15 | 246 | 6.9284 | 0.58397 | 4.046 |  | 0.0011802
107 | 151.726 | 7.2478 | 7.2499 | 33.7813 | 33.786 | 6.7669 | 1.6817 | 12 | 12 | 823 | 6.8486 | 0.52829 | 3.618 | 6 | 0.0012069
78 | 1000.85 | 2.8909 | 2.8921 | 34.3751 | 34.3803 | 7.4755 | 0.63917 | 9 | 9 | 612 | 8.5225 | 0.059255 | 0.505 | 6 | 0.001208
78 | 126.302 | 4.0963 | 4.0965 | 32.8423 | 32.8459 | 7.3339 | 2.5811 | 18 | 18 | 621 | 7.4198 | 0.93291 | 6.922 | 6 | 0.0012086
29 | 26.559 | 9.0383 | 9.0394 | 32.6517 | 32.656 | 6.5476 | 2.8249 | 21 | 21 | 252 | 6.5092 | 1.0425 | 6.786 |  | 0.0012928
29 | 51.266 | 7.9989 | 8.0006 | 32.668 | 32.6717 | 6.701 | 2.8141 | 20 | 20 | 251 | 6.6882 | 1.0376 | 6.94 |  | 0.0013204
107 | 401.461 | 5.3215 | 5.3231 | 34.0016 | 34.0063 | 7.0665 | 0.86084 | 7 | 7 | 818 | 7.4077 | 0.16159 | 1.197 |  | 0.0013959
29 | 600.811 | 3.9952 | 3.9952 | 34.1353 | 34.1402 | 7.288 | 0.6705 | 10 | 10 | 241 | 7.8598 | 0.073157 | 0.575 |  | 0.0014006
15 | 302.342 | 5.8263 | 5.8276 | 33.996 | 34.0052 | 6.9833 | 0.90861 | 8 | 8 | 124 | 7.2202 | 0.1831 | 1.322 |  | 0.0014129
78 | 601.704 | 3.6285 | 3.6296 | 34.1991 | 34.2042 | 7.3501 | 0.69071 | 11 | 11 | 614 | 7.9317 | 0.082202 | 0.652 |  | 0.0014488
61 | 1502.197 | 2.3476 | 2.3485 | 34.4989 | 34.5044 | 7.5707 | 0.6357 | 7 | 7 | 490 | 9.245 | 0.057436 | 0.531 |  | 0.0014652
29 | 802.144 | 3.5641 | 3.5651 | 34.2591 | 34.2635 | 7.3587 | 0.59358 | 9 | 9 | 240 | 8.1589 | 0.038486 | 0.314 |  | 0.001466
61 | 2001.892 | 1.9376 | 1.9382 | 34.5821 | 34.5878 | 7.6445 | 0.79466 | 6 | 6 | 489 | 9.997 | 0.12894 | 1.289 |  | 0.001479
107 | 1001.104 | 3.6378 | 3.6384 | 34.388 | 34.393 | 7.339 | 0.57213 | 4 | 4 | 815 | 8.357 | 0.03183 | 0.266 | 6 | 0.0015283
78 | 4002.275 | 1.5035 | 1.5039 | 34.6742 | 34.6809 | 7.7239 | 1.0365 | 2 | 2 | 605 | 13.2623 | 0.23767 | 3.152 |  | 0.0015735
107 | 601.432 | 4.7601 | 4.7614 | 34.1797 | 34.1845 | 7.1529 | 0.61086 | 6 | 6 | 817 | 7.707 | 0.049306 | 0.38 |  | 0.0015799
15 | 1317.447 | 2.9489 | 2.9489 | 34.4758 | 34.4805 | 7.4597 | 0.62596 | 1 | 1 | 117 | 8.8748 | 0.056114 | 0.498 |  | 0.0015948
78 | 4317.031 | 1.5199 | 1.5202 | 34.6775 | 34.6843 | 7.7205 | 1.0249 | 1 | 1 | 604 | 13.8339 | 0.2324 | 3.215 |  | 0.0016063
22 | 176.355 | 6.9346 | 6.936 | 33.8603 | 33.8645 | 6.8121 | 1.5035 | 11 | 11 | 196 | 6.9187 | 0.45095 | 3.12 |  | 0.0016104
3 | 53.403 | 7.8702 | 7.8709 | 32.8403 | 32.8446 | 6.713 | 2.315 | 5 | 5 | 152 | 6.7027 | 0.81609 | 5.47 |  | 0.0016861
78 | 201.208 | 4.0197 | 4.0215 | 33.7407 | 33.745 | 7.3031 | 1.3627 | 15 | 15 | 618 | 7.464 | 0.38773 | 2.894 |  | 0.0017584
29 | 1501.409 | 2.4198 | 2.4206 | 34.5051 | 34.5097 | 7.5568 | 0.66303 | 6 | 6 | 237 | 9.2257 | 0.069371 | 0.64 |  | 0.0018258
78 | 2002.265 | 1.9266 | 1.9274 | 34.5807 | 34.5863 | 7.6467 | 0.79291 | 6 | 6 | 609 | 10.0006 | 0.12779 | 1.278 |  | 0.0018381
29 | 401.365 | 4.552 | 4.553 | 33.9511 | 33.9552 | 7.1997 | 0.96797 | 11 | 11 | 242 | 7.554 | 0.20651 | 1.56 |  | 0.0018767
29 | 152.017 | 7.0264 | 7.0272 | 33.7893 | 33.7952 | 6.8009 | 1.9076 | 16 | 16 | 247 | 6.8845 | 0.62924 | 4.332 |  | 0.0018849
61 | 2501.023 | 1.7224 | 1.723 | 34.6232 | 34.629 | 7.684 | 0.90406 | 5 | 5 | 488 | 10.7571 | 0.17774 | 1.912 | 6 | 0.0018935
61 | 126.859 | 5.6253 | 5.6267 | 33.5798 | 33.5847 | 7.0358 | 1.9602 | 18 | 18 | 501 | 7.1074 | 0.65678 | 4.668 | 6 | 0.0019715
78 | 51.677 | 6.024 | 6.0349 | 32.7063 | 32.711 | 7.0108 | 2.7853 | 21 | 21 | 624 | 7.0088 | 1.028 | 7.205 |  | 0.0019897
107 | 800.793 | 4.1386 | 4.1398 | 34.3057 | 34.3106 | 7.2544 | 0.56391 | 5 | 5 | 816 | 8.035 | 0.028625 | 0.23 |  | 0.0020217
22 | 76.156 | 6.9161 | 6.9155 | 32.706 | 32.7082 | 6.867 | 2.6751 | 15 | 15 | 200 | 6.8823 | 0.97439 | 6.706 | 6 | 0.0020571
29 | 301.367 | 5.146 | 5.1476 | 33.8696 | 33.8739 | 7.1022 | 1.276 | 12 | 12 | 243 | 7.3475 | 0.34488 | 2.534 | 6 | 0.0020818
46 | 802.154 | 3.3278 | 3.3289 | 34.2887 | 34.2938 | 7.3999 | 0.60322 | 10 | 10 | 379 | 8.2075 | 0.042156 | 0.346 | 6 | 0.0021322
46 | 1251.135 | 2.6675 | 2.6683 | 34.4443 | 34.4496 | 7.5134 | 0.58627 | 8 | 8 | 377 | 8.8636 | 0.034523 | 0.306 |  | 0.0021398
61 | 298.794 | 4.5171 | 4.5175 | 33.9343 | 33.9397 | 7.2066 | 1.0818 | 13 | 13 | 496 | 7.4582 | 0.25744 | 1.92 |  | 0.0021574
29 | 11.802 | 9.0636 | 9.0653 | 32.6503 | 32.6543 | 6.5439 | 2.8313 | 22 | 22 | 253 | 6.4929 | 1.0445 | 6.782 |  | 0.0021719
29 | 101.54 | 6.9486 | 6.9499 | 32.6877 | 32.691 | 6.8627 | 2.6999 | 18 | 18 | 249 | 6.901 | 0.98536 | 6.8 |  | 0.0022372
61 | 201.869 | 5.1826 | 5.1838 | 33.8358 | 33.8408 | 7.0977 | 1.4596 | 15 | 15 | 498 | 7.2454 | 0.42731 | 3.096 |  | 0.0022735
46 | 11.675 | 8.0583 | 8.0606 | 32.7095 | 32.7141 | 6.6902 | 2.8372 | 23 | 23 | 392 | 6.6422 | 1.0471 | 6.955 |  | 0.0022796
78 | 1501.176 | 2.3235 | 2.3245 | 34.4994 | 34.5049 | 7.5752 | 0.63264 | 7 | 7 | 610 | 9.2497 | 0.055245 | 0.511 |  | 0.0022799
46 | 75.93 | 6.5345 | 6.5349 | 32.7301 | 32.7342 | 6.9268 | 2.7359 | 20 | 20 | 389 | 6.9442 | 1.0014 | 6.954 | 3 | 0.0023473
46 | 50.988 | 6.9508 | 6.9553 | 32.7155 | 32.7197 | 6.8611 | 2.8167 | 21 | 21 | 390 | 6.8532 | 1.0378 | 7.112 |  | 0.0023553
22 | 402.232 | 5.1345 | 5.1359 | 34.0131 | 34.0173 | 7.0974 | 0.84325 | 7 | 7 | 192 | 7.4424 | 0.15465 | 1.151 |  | 0.0023761
29 | 5.986 | 9.0603 | 9.0623 | 32.6504 | 32.6542 | 6.5444 | 2.8341 | 23 | 23 | 254 | 6.4884 | 1.0456 | 6.784 |  | 0.0023939
61 | 600.366 | 3.716 | 3.7171 | 34.1813 | 34.1864 | 7.3353 | 0.68297 | 11 | 11 | 494 | 7.9134 | 0.077717 | 0.615 |  | 0.0024516
61 | 251.5 | 4.813 | 4.8144 | 33.8868 | 33.8916 | 7.158 | 1.2312 | 14 | 14 | 497 | 7.3586 | 0.32438 | 2.387 |  | 0.0024517
61 | 6.409 | 7.0535 | 7.0546 | 32.6624 | 32.6668 | 6.8472 | 2.8087 | 24 | 24 | 507 | 6.7984 | 1.0341 | 7.03 |  | 0.0024538
46 | 151.552 | 6.1755 | 6.1775 | 33.8362 | 33.8414 | 6.9339 | 1.6871 | 17 | 17 | 386 | 7.024 | 0.52947 | 3.719 |  | 0.0024577
78 | 402.235 | 3.8855 | 3.8865 | 34.0297 | 34.0344 | 7.3126 | 0.84945 | 12 | 12 | 615 | 7.6797 | 0.15261 | 1.172 |  | 0.0024578
15 | 200.199 | 6.836 | 6.8381 | 33.8545 | 33.8591 | 6.8278 | 1.5843 | 10 | 10 | 126 | 6.9573 | 0.48323 | 3.362 |  | 0.0024641
61 | 11.641 | 7.041 | 7.0423 | 32.6629 | 32.6673 | 6.8492 | 2.8071 | 23 | 23 | 506 | 6.8051 | 1.0333 | 7.032 |  | 0.0024797
61 | 1001.251 | 3.0219 | 3.0231 | 34.3578 | 34.363 | 7.4523 | 0.59783 | 9 | 9 | 492 | 8.4946 | 0.039319 | 0.334 | 6 | 0.0025447
61 | 76.009 | 5.978 | 5.9777 | 32.6936 | 32.6976 | 7.019 | 2.7477 | 20 | 20 | 503 | 7.0401 | 1.0065 | 7.086 |  | 0.0025628
46 | 1501.355 | 2.3809 | 2.3817 | 34.4981 | 34.5036 | 7.5644 | 0.63275 | 7 | 7 | 376 | 9.2357 | 0.055004 | 0.508 |  | 0.0025705
61 | 26.547 | 7.0247 | 7.0262 | 32.6631 | 32.6678 | 6.8518 | 2.8014 | 22 | 22 | 505 | 6.8213 | 1.0306 | 7.03 | 2 | 0.0026482
78 | 302.137 | 3.8614 | 3.8624 | 33.8907 | 33.8954 | 7.3238 | 1.0002 | 13 | 13 | 616 | 7.5888 | 0.22019 | 1.671 |  | 0.0026894
107 | 1322.816 | 2.8818 | 2.881 | 34.4829 | 34.4881 | 7.4717 | 0.63446 | 1 | 1 | 812 | 8.8966 | 0.061035 | 0.543 |  | 0.002691
107 | 76.997 | 7.4897 | 7.4922 | 32.6758 | 32.6791 | 6.7786 | 2.6578 | 15 | 15 | 826 | 6.7914 | 0.97138 | 6.597 |  | 0.0027289
78 | 801.257 | 3.2401 | 3.2412 | 34.298 | 34.3032 | 7.4154 | 0.65638 | 10 | 10 | 613 | 8.2248 | 0.065412 | 0.538 |  | 0.0027935
61 | 800.214 | 3.3183 | 3.3195 | 34.2824 | 34.2875 | 7.4019 | 0.63456 | 10 | 10 | 493 | 8.2077 | 0.055557 | 0.456 | 2 | 0.0028313
61 | 1251.063 | 2.6636 | 2.6646 | 34.4388 | 34.4442 | 7.5145 | 0.58743 | 8 | 8 | 491 | 8.8648 | 0.03418 | 0.303 |  | 0.0030046
107 | 51.647 | 8.266 | 8.2675 | 32.6015 | 32.6056 | 6.6637 | 2.6965 | 16 | 16 | 827 | 6.65 | 0.98917 | 6.578 |  | 0.0031208
22 | 251.956 | 6.0844 | 6.0859 | 33.9039 | 33.9078 | 6.9455 | 1.2889 | 9 | 9 | 194 | 7.1312 | 0.3559 | 2.538 |  | 0.0031285
3 | 106.599 | 7.0364 | 7.0366 | 33.7913 | 33.7948 | 6.7993 | 1.1856 | 1 | 1 | 148 | 6.8414 | 0.30944 | 2.117 | 6 | 0.0031299
46 | 251.243 | 4.8512 | 4.8527 | 33.8558 | 33.8607 | 7.1529 | 1.2501 | 14 | 14 | 383 | 7.3528 | 0.33212 | 2.442 |  | 0.0031944
22 | 2002.255 | 1.9168 | 1.9172 | 34.5893 | 34.594 | 7.6482 | 0.82239 | 1 | 1 | 186 | 10.0027 | 0.14616 | 1.462 | 6 | 0.0032673
29 | 75.454 | 7.1961 | 7.1977 | 32.681 | 32.6844 | 6.824 | 2.7402 | 19 | 19 | 250 | 6.837 | 1.0023 | 6.853 | 6 | 0.0033687
22 | 1500.73 | 2.4672 | 2.4678 | 34.5065 | 34.511 | 7.5478 | 0.66429 | 2 | 2 | 187 | 9.213 | 0.07522 | 0.693 |  | 0.0034557
22 | 151.66 | 7.212 | 7.2136 | 33.7914 | 33.7956 | 6.772 | 1.6284 | 12 | 12 | 197 | 6.8539 | 0.50906 | 3.489 |  | 0.0035162
107 | 1251.613 | 3.0717 | 3.0724 | 34.4623 | 34.4673 | 7.4378 | 0.61438 | 3 | 3 | 814 | 8.7686 | 0.05303 | 0.465 |  | 0.0037209
61 | 101.395 | 5.3688 | 5.3691 | 32.875 | 32.8789 | 7.1119 | 2.5299 | 19 | 19 | 502 | 7.1616 | 0.90733 | 6.498 |  | 0.0037689
107 | 125.921 | 7.461 | 7.468 | 33.6701 | 33.6725 | 6.739 | 1.7091 | 13 | 13 | 824 | 6.7958 | 0.54563 | 3.708 |  | 0.0037829
46 | 402.365 | 4.2352 | 4.2359 | 34.0049 | 34.0101 | 7.2522 | 0.8376 | 12 | 12 | 381 | 7.613 | 0.14593 | 1.111 |  | 0.0038019
22 | 50.624 | 7.316 | 7.3162 | 32.6607 | 32.6639 | 6.8062 | 2.7712 | 16 | 16 | 201 | 6.7961 | 1.0156 | 6.902 |  | 0.0040845
22 | 1251.761 | 2.924 | 2.9252 | 34.4636 | 34.468 | 7.4649 | 0.61182 | 3 | 3 | 188 | 8.803 | 0.052255 | 0.46 |  | 0.0040971
15 | 52.234 | 8.5593 | 8.5615 | 32.5556 | 32.5593 | 6.6219 | 2.5795 | 16 | 16 | 132 | 6.6076 | 0.92924 | 6.14 |  | 0.0041861
46 | 176.96 | 5.5908 | 5.5975 | 33.8316 | 33.8362 | 7.0297 | 1.5351 | 16 | 16 | 385 | 7.1492 | 0.46788 | 3.345 |  | 0.0043428
46 | 27.614 | 7.7406 | 7.7422 | 32.7069 | 32.7114 | 6.7386 | 2.834 | 22 | 22 | 391 | 6.706 | 1.0435 | 6.998 |  | 0.0043651
22 | 300.937 | 5.6799 | 5.6807 | 33.9472 | 33.9514 | 7.0096 | 1.0519 | 8 | 8 | 193 | 7.2472 | 0.25058 | 1.816 |  | 0.0044311
78 | 151.922 | 4.5024 | 4.5036 | 33.4443 | 33.4473 | 7.2329 | 1.9787 | 17 | 17 | 620 | 7.3395 | 0.66762 | 4.9 |  | 0.0044984
15 | 25.525 | 10.0994 | 10.1116 | 31.9138 | 31.915 | 6.427 | 2.8535 | 17 | 17 | 133 | 6.3845 | 1.0615 | 6.777 |  | 0.0047975
107 | 176.362 | 7.0276 | 7.0302 | 33.852 | 33.8568 | 6.7979 | 1.5304 | 11 | 11 | 822 | 6.9038 | 0.46626 | 3.219 |  | 0.0048455
15 | 176.451 | 7.2558 | 7.2561 | 33.818 | 33.8227 | 6.764 | 1.5502 | 11 | 11 | 127 | 6.8682 | 0.46534 | 3.196 |  | 0.0049898
107 | 11.461 | 10.8242 | 10.8265 | 31.8205 | 31.8247 | 6.3306 | 2.8821 | 18 | 20 | 829 | 6.2747 | 1.0746 | 6.743 | 6 | 0.0050561
22 | 27.979 | 9.2112 | 9.2111 | 32.6262 | 32.6306 | 6.5236 | 2.8882 | 17 | 18 | 202 | 6.4859 | 1.0672 | 6.922 |  | 0.0050819
107 | 26.754 | 10.7009 | 10.7052 | 31.8401 | 31.8439 | 6.3467 | 2.8655 | 17 | 17 | 828 | 6.3037 | 1.0675 | 6.729 |  | 0.0054081
3 | 78.032 | 7.7408 | 7.7414 | 33.3401 | 33.3444 | 6.7108 | 1.756 | 4 | 4 | 151 | 6.7231 | 0.56983 | 3.831 |  | 0.0069054
78 | 5.981 | 6.3942 | 6.3956 | 32.7025 | 32.7067 | 6.9506 | 2.8189 | 24 | 24 | 627 | 6.9043 | 1.048 | 7.236 | 2 | 0.0069212
78 | 11.895 | 6.3781 | 6.379 | 32.7025 | 32.7068 | 6.9532 | 2.8175 | 23 | 23 | 626 | 6.9124 | 1.0475 | 7.241 |  | 0.0070432
15 | 251.391 | 6.3226 | 6.3217 | 33.9485 | 33.9531 | 6.9051 | 1.076 | 9 | 9 | 125 | 7.0875 | 0.26413 | 1.872 |  | 0.0071337
107 | 300.999 | 5.7827 | 5.7856 | 33.9263 | 33.9308 | 6.9937 | 1.1346 | 8 | 8 | 819 | 7.23 | 0.29073 | 2.102 |  | 0.0073952
46 | 101.254 | 6.2406 | 6.2417 | 32.889 | 32.8949 | 6.967 | 2.5428 | 19 | 19 | 388 | 7.0099 | 0.9247 | 6.482 |  | 0.0078075
3 | 103.39 | 7.0426 | 7.0435 | 33.7891 | 33.7928 | 6.7984 | 1.1888 | 2 | 2 | 149 | 6.8376 | 0.31619 | 2.162 |  | 0.0084357
3 | 78.14 | 7.7393 | 7.7407 | 33.3428 | 33.3458 | 6.7109 | 1.7484 | 3 | 3 | 150 | 6.7233 | 0.56862 | 3.823 |  | 0.0091166
29 | 200.723 | 6.5726 | 6.5736 | 33.8646 | 33.8691 | 6.869 | 1.7042 | 14 | 14 | 245 | 7.0014 | 0.52989 | 3.71 |  | 0.0097225
61 | 4024.135 | 1.5383 | 1.5386 | 34.6713 | 34.678 | 7.7172 | 1.0118 | 1 | 1 | 484 | 13.2887 | 0.2381 | 3.164 | 2 | 0.0099754
15 | 75.445 | 8.3103 | 8.3114 | 32.6606 | 32.6635 | 6.6545 | 2.4832 | 15 | 15 | 131 | 6.6616 | 0.90038 | 5.998 | 2 | 0.010314
3 | 11.419 | 9.9088 | 9.9118 | 31.8665 | 31.869 | 6.4558 | 2.971 | 7 | 7 | 154 | 6.4019 | 1.0902 | 6.979 |  | 0.019391
107 | 200.454 | 6.7389 | 6.7484 | 33.863 | 33.8672 | 6.8427 | 1.5582 | 10 | 10 | 821 | 6.9734 | 0.4946 | 3.449 | 3 | 0.020653
15 | 100.932 | 7.4669 | 7.4673 | 32.8401 | 32.8439 | 6.7748 | 2.4253 | 14 | 14 | 130 | 6.8093 | 0.84311 | 5.741 |  | 0.020918
29 | 126.208 | 6.6939 | 6.6932 | 32.8545 | 32.8548 | 6.8955 | 2.561 | 17 | 17 | 248 | 6.9584 | 0.90036 | 6.265 |  | 0.024743
22 | 101.397 | 7.3386 | 7.3425 | 33.2683 | 33.2735 | 6.7757 | 2.1283 | 14 | 14 | 199 | 6.8112 | 0.75904 | 5.17 |  | 0.028603
107 | 101.735 | 7.7031 | 7.7032 | 33.2876 | 33.2892 | 6.7188 | 1.9066 | 14 | 14 | 825 | 6.7525 | 0.67442 | 4.554 |  | 0.043715
107 | 251.741 | 6.221 | 6.2261 | 33.8918 | 33.8966 | 6.924 | 1.3508 | 9 | 9 | 820 | 7.108 | 0.58892 | 4.186 |  | 0.20827
1 | 6.363 | 11.4514 | 11.4659 | 27.6984 | 27.6946 | 6.4113 | 3.7276 | 11 | 11 | 11 | 6.3485 |  | 9.119 |  | 1.4499
1 | 11.597 | 10.0028 | 10.0102 | 28.151 | 28.1461 | 6.5977 | 3.3148 | 10 | 10 | 10 | 6.5424 | 1.2638 | 8.268 |  | 0.00045353
1 | 26.556 | 9.0692 | 9.0713 | 29.1328 | 29.1329 | 6.6935 | 2.9641 | 9 | 9 | 9 | 6.6542 | 1.0708 | 7.125 | 2 | 0.035711
1 | 51.666 | 7.7865 | 7.7903 | 29.5735 | 29.5736 | 6.8708 | 1.9882 | 8 | 8 | 8 | 6.8591 | 0.66219 | 4.542 | 6 | 0.0052049
1 | 76.528 | 7.8869 | 7.8867 | 30.2593 | 30.2609 | 6.8244 | 1.4978 | 7 | 7 | 7 | 6.8348 | 0.44317 | 3.029 |  | 0.0035886
1 | 101.293 | 8.2436 | 8.24 | 30.6548 | 30.6537 | 6.752 | 0.9495 | 6 | 6 | 6 | 6.7827 | 0.21201 | 1.438 |  | 0.011929
1 | 125.723 | 8.847 | 8.8472 | 31.0148 | 31.0169 | 6.6455 | 0.54591 | 5 | 5 | 5 | 6.6945 | 0.030025 | 0.201 |  | 0.01152
1 | 150.761 | 9.1346 | 9.1351 | 31.2092 | 31.2113 | 6.5947 | 0.52486 | 4 | 4 | 4 | 6.6641 | 0.0072028 | 0.048 |  | 0.0018317
1 | 177.047 | 9.1366 | 9.1371 | 31.249 | 31.2511 | 6.5927 | 0.52536 | 3 | 3 | 3 | 6.6852 | 0 | 0 |  | 0.0092595
1 | 202.433 | 9.1358 | 9.1362 | 31.2647 | 31.2667 | 6.5922 | 0.52622 | 2 | 2 | 2 | 6.7071 | 0 | 0 |  | 0.0096464
2.971
0.52486

## bottle data
Event_Number | Station | Sample_Number | Oxygen:Dissolved | Temperature:Draw | Comment | Event_Number | Sample_Number
1 | SI03 | 1 | -99 | 11.1 | OXY: Spigot open, Bubble in burrette. Overall a bad sample. Should be anoxic as well | 1 | 9 | bit off but not bad for this cast
1 | SI03 | 2 | 0 | 9.3 | OXY: Anoxic sample | 15 | 131 | mild outlier
1 | SI03 | 3 | 0 | 9.3 | OXY: Anoxic sample | 61 | 484 | little off
1 | SI03 | 4 | 0.048 | 9.3 |  | 61 | 485 | ok
1 | SI03 | 5 | 0.201 | 9.3 |  | 61 | 486 | ok
1 | SI03 | 6 | 1.438 | 9.1 |  | 61 | 493
1 | SI03 | 7 | 3.029 | 8.5 |  | 61 | 505 | ok
1 | SI03 | 8 | 4.542 | 8.4 |  | 78 | 625 | ok
1 | SI03 | 9 | 7.125 | 9.2 | OXY: Bubble on lens, however; Titration clean and ended at 100%. | 78 | 627 | ok
1 | SI03 | 10 | 8.268 | 10.2 |  | 46 | 389 | ok in compare
1 | SI03 | 11 | 9.119 | 11.4 |  | 107 | 821 | ok
3 | P2 | 148 | 2.117 | 7.5
3 | P2 | 149 | 2.162 | 7.5
3 | P2 | 150 | 3.823 | 8.1
3 | P2 | 151 | 3.831 | 8.1
3 | P2 | 152 | 5.47 | 8.2
3 | P2 | 153 | 6.245 | 8.8
3 | P2 | 154 | 6.979 | 9.9
3 | P2 | 155 | 6.872 | 10.6
15 | P4 | 117 | 0.498 | 4.2
15 | P4 | 118 | 0.425 | 4.2
15 | P4 | 119 | 0.425 | 4.2
15 | P4 | 120 | 0.275 | 5
15 | P4 | 121 | 0.209 | 5.7
15 | P4 | 122 | 0.338 | 5.2
15 | P4 | 123 | 0.876 | 6.4
15 | P4 | 124 | 1.322 | 6.7
15 | P4 | 125 | 1.872 | 7
15 | P4 | 126 | 3.362 | 7.4
15 | P4 | 127 | 3.196 | 7.7
15 | P4 | 128 | 3.117 | 8.1
15 | P4 | 129 | 3.633 | 8.3
15 | P4 | 130 | 5.741 | 8
15 | P4 | 131 | 5.998 | 8.6 | OXY: Bad end point. Started bad
15 | P4 | 132 | 6.14 | 8.9
15 | P4 | 133 | 6.777 | 10.3
15 | P4 | 134 | 6.737 | 10.8
15 | P4 | 135 | 6.736 | 10.9
22 | P8 | 186 | 1.462 | 3.5
22 | P8 | 187 | 0.693 | 3.6
22 | P8 | 188 | 0.46 | 4
22 | P8 | 189 | 0.257 | 4.6
22 | P8 | 190 | 0.234 | 5
22 | P8 | 191 | 0.355 | 5.6
22 | P8 | 192 | 1.151 | 6.1
22 | P8 | 193 | 1.816 | 6.4
22 | P8 | 194 | 2.538 | 6.8
22 | P8 | 195 | 3.036 | 7.3
22 | P8 | 196 | 3.12 | 7.4
22 | P8 | 197 | 3.489 | 7.6
22 | P8 | 198 | 3.828 | 7.8
22 | P8 | 199 | 5.17 | 7.7
22 | P8 | 200 | 6.706 | 7.5
22 | P8 | 201 | 6.902 | 7.8
22 | P8 | 202 | 6.922 | 9.4
22 | P8 | 203 | 6.743 | 10.3
22 | P8 | 204 | 6.734 | 10.3
29 | P12 | 232 | 2.322 | 2.8
29 | P12 | 233 | 2.291 | 2.6
29 | P12 | 234 | 2.29 | 2.7
29 | P12 | 235 | 1.887 | 2.7
29 | P12 | 236 | 1.353 | 3
29 | P12 | 237 | 0.64 | 3.4
29 | P12 | 238 | 0.339 | 3.9
29 | P12 | 239 | 0.245 | 4.2
29 | P12 | 240 | 0.314 | 4.6
29 | P12 | 241 | 0.575 | 5.2
29 | P12 | 242 | 1.56 | 5.6
29 | P12 | 243 | 2.534 | 6.1
29 | P12 | 244 | 2.897 | 6.5
29 | P12 | 245 | 3.71 | 7.1
29 | P12 | 246 | 4.046 | 7.3
29 | P12 | 247 | 4.332 | 7.4
29 | P12 | 248 | 6.265 | 7.3
29 | P12 | 249 | 6.8 | 7.4
29 | P12 | 250 | 6.853 | 7.6
29 | P12 | 251 | 6.94 | 8.2
29 | P12 | 252 | 6.786 | 9.1
29 | P12 | 253 | 6.782 | 9
29 | P12 | 254 | 6.784 | 9.1
46 | P16 | 370 | 2.794 | 2.6
46 | P16 | 371 | 2.696 | 2.4
46 | P16 | 372 | 2.701 | 2.4
46 | P16 | 373 | 2.455 | 2.3
46 | P16 | 374 | 1.913 | 2.4
46 | P16 | 375 | 1.287 | 2.6
46 | P16 | 376 | 0.508 | 3.1
46 | P16 | 377 | 0.306 | 3.3
46 | P16 | 378 | 0.253 | 3.7
46 | P16 | 379 | 0.346 | 4.1
46 | P16 | 380 | 0.532 | 4.5
46 | P16 | 381 | 1.111 | 5.1
46 | P16 | 382 | 1.993 | 5.5
46 | P16 | 383 | 2.442 | 5.7
46 | P16 | 384 | 2.98 | 5.9
46 | P16 | 385 | 3.345 | 6.2
46 | P16 | 386 | 3.719 | 6.7
46 | P16 | 387 | -99 | 6.9 | OXY: Titration incomplete. Likely a bubble in light path
46 | P16 | 388 | 6.482 | 6.7
46 | P16 | 389 | 6.954 | 6.9 | OXY: Bubble on lens. Messy curve. End point at only 96%. Vlaue not trusted
46 | P16 | 390 | 7.112 | 7.2
46 | P16 | 391 | 6.998 | 7.8
46 | P16 | 392 | 6.955 | 8.1
46 | P16 | 393 | 6.968 | 8.1
61 | P26 | 484 | 3.164 | 2.4 | OXY: Bubbles from MnCl2 Dispenser
61 | P26 | 485 | 2.884 | 2.3 | OXY: Bubbles from MnCl2 Dispenser
61 | P26 | 486 | 2.877 | 2.3 | OXY: Bubbles from MnCl2 Dispenser
61 | P26 | 487 | 2.516 | 2.4
61 | P26 | 488 | 1.912 | 2.6
61 | P26 | 489 | 1.289 | 2.7
61 | P26 | 490 | 0.531 | 3.1
61 | P26 | 491 | 0.303 | 3.4
61 | P26 | 492 | 0.334 | 3.8
61 | P26 | 493 | 0.456 | 4.1 | OXY: Draw temperature may be incorrect
61 | P26 | 494 | 0.615 | 4.7
61 | P26 | 495 | 1.28 | 5.2
61 | P26 | 496 | 1.92 | 5.5
61 | P26 | 497 | 2.387 | 5.7
61 | P26 | 498 | 3.096 | 6
61 | P26 | 499 | 3.356 | 6.2
61 | P26 | 500 | 3.713 | 6.3
61 | P26 | 501 | 4.668 | 6.4
61 | P26 | 502 | 6.498 | 6.2
61 | P26 | 503 | 7.086 | 6.6
61 | P26 | 504 | 7.078 | 7.3
61 | P26 | 505 | 7.03 | 7.4 | OXY: Top valve open slightly
61 | P26 | 506 | 7.032 | 7.5
61 | P26 | 507 | 7.03 | 7.4
78 | P26 | 604 | 3.215 | 2.4
78 | P26 | 605 | 3.152 | 2.1
78 | P26 | 606 | 2.931 | 2.1
78 | P26 | 607 | 2.544 | 2.2
78 | P26 | 608 | 1.961 | 2.4
78 | P26 | 609 | 1.278 | 2.4
78 | P26 | 610 | 0.511 | 2.8
78 | P26 | 611 | 0.497 | 3.1
78 | P26 | 612 | 0.505 | 3.4
78 | P26 | 613 | 0.538 | 3.8
78 | P26 | 614 | 0.652 | 4.2
78 | P26 | 615 | 1.172 | 4.5
78 | P26 | 616 | 1.671 | 4.5
78 | P26 | 617 | 2.196 | 4.5
78 | P26 | 618 | 2.894 | 4.7
78 | P26 | 619 | 3.8 | 5
78 | P26 | 620 | 4.9 | 5.2
78 | P26 | 621 | 6.922 | 4.9
78 | P26 | 622 | 7.086 | 5.6
78 | P26 | 623 | 7.205 | 5.9
78 | P26 | 624 | 7.205 | 6.4
78 | P26 | 625 | 7.193 | 6.7 | OXY: Top valve leaking
78 | P26 | 626 | 7.241 | 6.7
78 | P26 | 627 | 7.236 | 6.7 | OXY: Top valve leaking
107 | P4 | 812 | 0.543 | 4
107 | P4 | 813 | 0.426 | 4
107 | P4 | 814 | 0.465 | 4
107 | P4 | 815 | 0.266 | 4.7
107 | P4 | 816 | 0.23 | 5.2
107 | P4 | 817 | 0.38 | 5.7
107 | P4 | 818 | 1.197 | 6.2
107 | P4 | 819 | 2.102 | 6.5
107 | P4 | 820 | 4.186 | 6.9
107 | P4 | 821 | 3.449 | 7.3 | OXY: Messy curve. End point ambiguous
107 | P4 | 822 | 3.219 | 7.5
107 | P4 | 823 | 3.618 | 7.7
107 | P4 | 824 | 3.708 | 7.8
107 | P4 | 825 | 4.554 | 8
107 | P4 | 826 | 6.597 | 7.9
107 | P4 | 827 | 6.578 | 8.5
107 | P4 | 828 | 6.729 | 10.6
107 | P4 | 829 | 6.743 | 10.6
107 | P4 | 830 | -99 | 10.7 | OXY: Analyst did not add acid to sample. 5 mL of Thio added. Sample lost.

## ctd data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Remove where no DO sampling
1 | 215.396 | 9.1382 | 9.1385 | 31.2732 | 31.2753 | 6.5915 | 0.52679 | 1 | 1 | 1
1 | 202.433 | 9.1358 | 9.1362 | 31.2647 | 31.2667 | 6.5922 | 0.52622 | 2 | 2 | 2
1 | 177.047 | 9.1366 | 9.1371 | 31.249 | 31.2511 | 6.5927 | 0.52536 | 3 | 3 | 3
1 | 150.761 | 9.1346 | 9.1351 | 31.2092 | 31.2113 | 6.5947 | 0.52486 | 4 | 4 | 4
1 | 125.723 | 8.847 | 8.8472 | 31.0148 | 31.0169 | 6.6455 | 0.54591 | 5 | 5 | 5
1 | 101.293 | 8.2436 | 8.24 | 30.6548 | 30.6537 | 6.752 | 0.9495 | 6 | 6 | 6
1 | 76.528 | 7.8869 | 7.8867 | 30.2593 | 30.2609 | 6.8244 | 1.4978 | 7 | 7 | 7
1 | 51.666 | 7.7865 | 7.7903 | 29.5735 | 29.5736 | 6.8708 | 1.9882 | 8 | 8 | 8
1 | 26.556 | 9.0692 | 9.0713 | 29.1328 | 29.1329 | 6.6935 | 2.9641 | 9 | 9 | 9
1 | 11.597 | 10.0028 | 10.0102 | 28.151 | 28.1461 | 6.5977 | 3.3148 | 10 | 10 | 10
1 | 6.363 | 11.4514 | 11.4659 | 27.6984 | 27.6946 | 6.4113 | 3.7276 | 11 | 11 | 11
3 | 106.599 | 7.0364 | 7.0366 | 33.7913 | 33.7948 | 6.7993 | 1.1856 | 1 | 1 | 148
3 | 103.39 | 7.0426 | 7.0435 | 33.7891 | 33.7928 | 6.7984 | 1.1888 | 2 | 2 | 149
3 | 78.14 | 7.7393 | 7.7407 | 33.3428 | 33.3458 | 6.7109 | 1.7484 | 3 | 3 | 150
3 | 78.032 | 7.7408 | 7.7414 | 33.3401 | 33.3444 | 6.7108 | 1.756 | 4 | 4 | 151
3 | 53.403 | 7.8702 | 7.8709 | 32.8403 | 32.8446 | 6.713 | 2.315 | 5 | 5 | 152
3 | 27.092 | 8.4504 | 8.4525 | 32.522 | 32.5251 | 6.6395 | 2.6077 | 6 | 6 | 153
3 | 11.419 | 9.9088 | 9.9118 | 31.8665 | 31.869 | 6.4558 | 2.971 | 7 | 7 | 154
3 | 5.795 | 10.5404 | 10.5453 | 31.7641 | 31.7674 | 6.3718 | 2.9244 | 8 | 8 | 155
15 | 1317.447 | 2.9489 | 2.9489 | 34.4758 | 34.4805 | 7.4597 | 0.62596 | 1 | 1 | 117
15 | 1251.992 | 3.0982 | 3.1005 | 34.4582 | 34.4627 | 7.4332 | 0.61197 | 2 | 2 | 118
15 | 1251.83 | 3.1006 | 3.1012 | 34.4579 | 34.4627 | 7.4328 | 0.61179 | 3 | 3 | 119
15 | 1001.432 | 3.5549 | 3.5555 | 34.3998 | 34.4045 | 7.3533 | 0.57614 | 4 | 4 | 120
15 | 802.22 | 4.1438 | 4.1448 | 34.3073 | 34.3119 | 7.2534 | 0.56269 | 5 | 5 | 121
15 | 601.784 | 4.6983 | 4.6992 | 34.1968 | 34.2014 | 7.1627 | 0.60203 | 6 | 6 | 122
15 | 401.671 | 5.4554 | 5.4565 | 34.0634 | 34.068 | 7.0413 | 0.76632 | 7 | 7 | 123
15 | 302.342 | 5.8263 | 5.8276 | 33.996 | 34.0052 | 6.9833 | 0.90861 | 8 | 8 | 124
15 | 251.391 | 6.3226 | 6.3217 | 33.9485 | 33.9531 | 6.9051 | 1.076 | 9 | 9 | 125
15 | 200.199 | 6.836 | 6.8381 | 33.8545 | 33.8591 | 6.8278 | 1.5843 | 10 | 10 | 126
15 | 176.451 | 7.2558 | 7.2561 | 33.818 | 33.8227 | 6.764 | 1.5502 | 11 | 11 | 127
15 | 151.95 | 7.537 | 7.5381 | 33.7524 | 33.7573 | 6.7237 | 1.5211 | 12 | 12 | 128
15 | 125.765 | 7.8334 | 7.8351 | 33.5067 | 33.5107 | 6.6894 | 1.7025 | 13 | 13 | 129
15 | 100.932 | 7.4669 | 7.4673 | 32.8401 | 32.8439 | 6.7748 | 2.4253 | 14 | 14 | 130
15 | 75.445 | 8.3103 | 8.3114 | 32.6606 | 32.6635 | 6.6545 | 2.4832 | 15 | 15 | 131
15 | 52.234 | 8.5593 | 8.5615 | 32.5556 | 32.5593 | 6.6219 | 2.5795 | 16 | 16 | 132
15 | 25.525 | 10.0994 | 10.1116 | 31.9138 | 31.915 | 6.427 | 2.8535 | 17 | 17 | 133
15 | 11.449 | 11.0087 | 11.0105 | 31.7788 | 31.7831 | 6.3072 | 2.9001 | 18 | 18 | 134
15 | 6.6 | 11.0321 | 11.034 | 31.7741 | 31.7754 | 6.3043 | 2.9025 | 19 | 19 | 135
22 | 2002.255 | 1.9168 | 1.9172 | 34.5893 | 34.594 | 7.6482 | 0.82239 | 1 | 1 | 186
22 | 1500.73 | 2.4672 | 2.4678 | 34.5065 | 34.511 | 7.5478 | 0.66429 | 2 | 2 | 187
22 | 1251.761 | 2.924 | 2.9252 | 34.4636 | 34.468 | 7.4649 | 0.61182 | 3 | 3 | 188
22 | 1001.2 | 3.4014 | 3.4026 | 34.3806 | 34.3849 | 7.3819 | 0.57048 | 4 | 4 | 189
22 | 801.132 | 3.8091 | 3.8101 | 34.2922 | 34.2964 | 7.3132 | 0.56978 | 5 | 5 | 190
22 | 601.83 | 4.4663 | 4.4675 | 34.1861 | 34.1903 | 7.2032 | 0.6064 | 6 | 6 | 191
22 | 402.232 | 5.1345 | 5.1359 | 34.0131 | 34.0173 | 7.0974 | 0.84325 | 7 | 7 | 192
22 | 300.937 | 5.6799 | 5.6807 | 33.9472 | 33.9514 | 7.0096 | 1.0519 | 8 | 8 | 193
22 | 251.956 | 6.0844 | 6.0859 | 33.9039 | 33.9078 | 6.9455 | 1.2889 | 9 | 9 | 194
22 | 201.944 | 6.7424 | 6.7443 | 33.878 | 33.8822 | 6.8415 | 1.4724 | 10 | 10 | 195
22 | 176.355 | 6.9346 | 6.936 | 33.8603 | 33.8645 | 6.8121 | 1.5035 | 11 | 11 | 196
22 | 151.66 | 7.212 | 7.2136 | 33.7914 | 33.7956 | 6.772 | 1.6284 | 12 | 12 | 197
22 | 127.302 | 7.4203 | 7.4216 | 33.6463 | 33.6507 | 6.7463 | 1.7539 | 13 | 13 | 198
22 | 101.397 | 7.3386 | 7.3425 | 33.2683 | 33.2735 | 6.7757 | 2.1283 | 14 | 14 | 199
22 | 76.156 | 6.9161 | 6.9155 | 32.706 | 32.7082 | 6.867 | 2.6751 | 15 | 15 | 200
22 | 50.624 | 7.316 | 7.3162 | 32.6607 | 32.6639 | 6.8062 | 2.7712 | 16 | 16 | 201
22 | 27.979 | 9.2112 | 9.2111 | 32.6262 | 32.6306 | 6.5236 | 2.8882 | 17 | 18 | 202
22 | 11.139 | 10.5313 | 10.5336 | 31.9837 | 31.9874 | 6.3641 | 2.8829 | 18 | 20 | 203
22 | 7.112 | 10.5335 | 10.5354 | 31.9829 | 31.9867 | 6.3639 | 2.8811 | 19 | 21 | 204
29 | 3276.45 | 1.6715 | 1.6718 | 34.6431 | 34.6486 | 7.6928 | 0.93449 | 1 | 1 | 232
29 | 3001.327 | 1.6482 | 1.6485 | 34.6426 | 34.648 | 7.6973 | 0.94797 | 2 | 2 | 233
29 | 3001.159 | 1.6484 | 1.6487 | 34.6425 | 34.648 | 7.6973 | 0.94785 | 3 | 3 | 234
29 | 2501.274 | 1.7508 | 1.7511 | 34.6226 | 34.6278 | 7.6785 | 0.89666 | 4 | 4 | 235
29 | 2000.287 | 1.9531 | 1.9537 | 34.5824 | 34.5873 | 7.6415 | 0.80823 | 5 | 5 | 236
29 | 1501.409 | 2.4198 | 2.4206 | 34.5051 | 34.5097 | 7.5568 | 0.66303 | 6 | 6 | 237
29 | 1251.696 | 2.8336 | 2.8347 | 34.4376 | 34.4422 | 7.4829 | 0.59097 | 7 | 7 | 238
29 | 1001.354 | 3.2095 | 3.2106 | 34.3624 | 34.3668 | 7.4177 | 0.57121 | 8 | 8 | 239
29 | 802.144 | 3.5641 | 3.5651 | 34.2591 | 34.2635 | 7.3587 | 0.59358 | 9 | 9 | 240
29 | 600.811 | 3.9952 | 3.9952 | 34.1353 | 34.1402 | 7.288 | 0.6705 | 10 | 10 | 241
29 | 401.365 | 4.552 | 4.553 | 33.9511 | 33.9552 | 7.1997 | 0.96797 | 11 | 11 | 242
29 | 301.367 | 5.146 | 5.1476 | 33.8696 | 33.8739 | 7.1022 | 1.276 | 12 | 12 | 243
29 | 251.355 | 5.81 | 5.8109 | 33.8664 | 33.8708 | 6.992 | 1.4025 | 13 | 13 | 244
29 | 200.723 | 6.5726 | 6.5736 | 33.8646 | 33.8691 | 6.869 | 1.7042 | 14 | 14 | 245
29 | 175.892 | 6.8758 | 6.8776 | 33.8522 | 33.8567 | 6.8217 | 1.8054 | 15 | 15 | 246
29 | 152.017 | 7.0264 | 7.0272 | 33.7893 | 33.7952 | 6.8009 | 1.9076 | 16 | 16 | 247
29 | 126.208 | 6.6939 | 6.6932 | 32.8545 | 32.8548 | 6.8955 | 2.561 | 17 | 17 | 248
29 | 101.54 | 6.9486 | 6.9499 | 32.6877 | 32.691 | 6.8627 | 2.6999 | 18 | 18 | 249
29 | 75.454 | 7.1961 | 7.1977 | 32.681 | 32.6844 | 6.824 | 2.7402 | 19 | 19 | 250
29 | 51.266 | 7.9989 | 8.0006 | 32.668 | 32.6717 | 6.701 | 2.8141 | 20 | 20 | 251
29 | 26.559 | 9.0383 | 9.0394 | 32.6517 | 32.656 | 6.5476 | 2.8249 | 21 | 21 | 252
29 | 11.802 | 9.0636 | 9.0653 | 32.6503 | 32.6543 | 6.5439 | 2.8313 | 22 | 22 | 253
29 | 5.986 | 9.0603 | 9.0623 | 32.6504 | 32.6542 | 6.5444 | 2.8341 | 23 | 23 | 254
46 | 3686.661 | 1.5689 | 1.5692 | 34.6626 | 34.6691 | 7.7117 | 0.99288 | 1 | 1 | 370
46 | 3502.546 | 1.5723 | 1.5727 | 34.6597 | 34.6661 | 7.7112 | 0.99106 | 2 | 2 | 371
46 | 3502.228 | 1.5725 | 1.5729 | 34.6597 | 34.666 | 7.7112 | 0.9908 | 3 | 3 | 372
46 | 3002.397 | 1.5941 | 1.5945 | 34.6494 | 34.6554 | 7.7075 | 0.9795 | 4 | 4 | 373
46 | 2500.769 | 1.7311 | 1.7315 | 34.6243 | 34.6301 | 7.6822 | 0.90277 | 5 | 5 | 374
46 | 2001.08 | 1.9486 | 1.9491 | 34.5822 | 34.5879 | 7.6424 | 0.79323 | 6 | 6 | 375
46 | 1501.355 | 2.3809 | 2.3817 | 34.4981 | 34.5036 | 7.5644 | 0.63275 | 7 | 7 | 376
46 | 1251.135 | 2.6675 | 2.6683 | 34.4443 | 34.4496 | 7.5134 | 0.58627 | 8 | 8 | 377
46 | 1001.163 | 3.0572 | 3.0582 | 34.3769 | 34.382 | 7.4448 | 0.57326 | 9 | 9 | 378
46 | 802.154 | 3.3278 | 3.3289 | 34.2887 | 34.2938 | 7.3999 | 0.60322 | 10 | 10 | 379
46 | 601.861 | 3.7584 | 3.7594 | 34.1758 | 34.1809 | 7.328 | 0.6542 | 11 | 11 | 380
46 | 402.365 | 4.2352 | 4.2359 | 34.0049 | 34.0101 | 7.2522 | 0.8376 | 12 | 12 | 381
46 | 302.467 | 4.5866 | 4.5879 | 33.89 | 33.8948 | 7.1967 | 1.1016 | 13 | 13 | 382
46 | 251.243 | 4.8512 | 4.8527 | 33.8558 | 33.8607 | 7.1529 | 1.2501 | 14 | 14 | 383
46 | 202.859 | 5.2348 | 5.2364 | 33.8345 | 33.8395 | 7.089 | 1.4209 | 15 | 15 | 384
46 | 176.96 | 5.5908 | 5.5975 | 33.8316 | 33.8362 | 7.0297 | 1.5351 | 16 | 16 | 385
46 | 151.552 | 6.1755 | 6.1775 | 33.8362 | 33.8414 | 6.9339 | 1.6871 | 17 | 17 | 386
46 | 101.254 | 6.2406 | 6.2417 | 32.889 | 32.8949 | 6.967 | 2.5428 | 19 | 19 | 388
46 | 75.93 | 6.5345 | 6.5349 | 32.7301 | 32.7342 | 6.9268 | 2.7359 | 20 | 20 | 389
46 | 50.988 | 6.9508 | 6.9553 | 32.7155 | 32.7197 | 6.8611 | 2.8167 | 21 | 21 | 390
46 | 27.614 | 7.7406 | 7.7422 | 32.7069 | 32.7114 | 6.7386 | 2.834 | 22 | 22 | 391
46 | 11.675 | 8.0583 | 8.0606 | 32.7095 | 32.7141 | 6.6902 | 2.8372 | 23 | 23 | 392
46 | 6.198 | 8.0572 | 8.0583 | 32.7099 | 32.7145 | 6.6904 | 2.8386 | 24 | 24 | 393
61 | 4024.135 | 1.5383 | 1.5386 | 34.6713 | 34.678 | 7.7172 | 1.0118 | 1 | 1 | 484
61 | 3500.216 | 1.5245 | 1.525 | 34.6648 | 34.6712 | 7.7203 | 1.0229 | 2 | 2 | 485
61 | 3501.057 | 1.5249 | 1.5253 | 34.6648 | 34.6712 | 7.7202 | 1.0227 | 3 | 3 | 486
61 | 3001.222 | 1.5782 | 1.5787 | 34.6498 | 34.6559 | 7.7106 | 0.99087 | 4 | 4 | 487
61 | 2501.023 | 1.7224 | 1.723 | 34.6232 | 34.629 | 7.684 | 0.90406 | 5 | 5 | 488
61 | 2001.892 | 1.9376 | 1.9382 | 34.5821 | 34.5878 | 7.6445 | 0.79466 | 6 | 6 | 489
61 | 1502.197 | 2.3476 | 2.3485 | 34.4989 | 34.5044 | 7.5707 | 0.6357 | 7 | 7 | 490
61 | 1251.063 | 2.6636 | 2.6646 | 34.4388 | 34.4442 | 7.5145 | 0.58743 | 8 | 8 | 491
61 | 1001.251 | 3.0219 | 3.0231 | 34.3578 | 34.363 | 7.4523 | 0.59783 | 9 | 9 | 492
61 | 800.214 | 3.3183 | 3.3195 | 34.2824 | 34.2875 | 7.4019 | 0.63456 | 10 | 10 | 493
61 | 600.366 | 3.716 | 3.7171 | 34.1813 | 34.1864 | 7.3353 | 0.68297 | 11 | 11 | 494
61 | 401.146 | 4.1999 | 4.2012 | 34.0341 | 34.0391 | 7.2569 | 0.88082 | 12 | 12 | 495
61 | 298.794 | 4.5171 | 4.5175 | 33.9343 | 33.9397 | 7.2066 | 1.0818 | 13 | 13 | 496
61 | 251.5 | 4.813 | 4.8144 | 33.8868 | 33.8916 | 7.158 | 1.2312 | 14 | 14 | 497
61 | 201.869 | 5.1826 | 5.1838 | 33.8358 | 33.8408 | 7.0977 | 1.4596 | 15 | 15 | 498
61 | 176.828 | 5.4128 | 5.4139 | 33.8352 | 33.8399 | 7.0591 | 1.5457 | 16 | 16 | 499
61 | 151.564 | 5.4873 | 5.4884 | 33.7871 | 33.792 | 7.049 | 1.6581 | 17 | 17 | 500
61 | 126.859 | 5.6253 | 5.6267 | 33.5798 | 33.5847 | 7.0358 | 1.9602 | 18 | 18 | 501
61 | 101.395 | 5.3688 | 5.3691 | 32.875 | 32.8789 | 7.1119 | 2.5299 | 19 | 19 | 502
61 | 76.009 | 5.978 | 5.9777 | 32.6936 | 32.6976 | 7.019 | 2.7477 | 20 | 20 | 503
61 | 50.949 | 6.8953 | 6.9024 | 32.6696 | 32.6739 | 6.8719 | 2.7955 | 21 | 21 | 504
61 | 26.547 | 7.0247 | 7.0262 | 32.6631 | 32.6678 | 6.8518 | 2.8014 | 22 | 22 | 505
61 | 11.641 | 7.041 | 7.0423 | 32.6629 | 32.6673 | 6.8492 | 2.8071 | 23 | 23 | 506
61 | 6.409 | 7.0535 | 7.0546 | 32.6624 | 32.6668 | 6.8472 | 2.8087 | 24 | 24 | 507
78 | 4317.031 | 1.5199 | 1.5202 | 34.6775 | 34.6843 | 7.7205 | 1.0249 | 1 | 1 | 604
78 | 4002.275 | 1.5035 | 1.5039 | 34.6742 | 34.6809 | 7.7239 | 1.0365 | 2 | 2 | 605
78 | 3501.917 | 1.5109 | 1.5115 | 34.6659 | 34.6722 | 7.7229 | 1.033 | 3 | 3 | 606
78 | 3001.236 | 1.5711 | 1.5717 | 34.6506 | 34.6567 | 7.7119 | 0.99611 | 4 | 4 | 607
78 | 2500.819 | 1.7084 | 1.7091 | 34.6234 | 34.6293 | 7.6867 | 0.91238 | 5 | 5 | 608
78 | 2002.265 | 1.9266 | 1.9274 | 34.5807 | 34.5863 | 7.6467 | 0.79291 | 6 | 6 | 609
78 | 1501.176 | 2.3235 | 2.3245 | 34.4994 | 34.5049 | 7.5752 | 0.63264 | 7 | 7 | 610
78 | 1250.903 | 2.5836 | 2.5845 | 34.4403 | 34.4457 | 7.5293 | 0.63108 | 8 | 8 | 611
78 | 1000.85 | 2.8909 | 2.8921 | 34.3751 | 34.3803 | 7.4755 | 0.63917 | 9 | 9 | 612
78 | 801.257 | 3.2401 | 3.2412 | 34.298 | 34.3032 | 7.4154 | 0.65638 | 10 | 10 | 613
78 | 601.704 | 3.6285 | 3.6296 | 34.1991 | 34.2042 | 7.3501 | 0.69071 | 11 | 11 | 614
78 | 402.235 | 3.8855 | 3.8865 | 34.0297 | 34.0344 | 7.3126 | 0.84945 | 12 | 12 | 615
78 | 302.137 | 3.8614 | 3.8624 | 33.8907 | 33.8954 | 7.3238 | 1.0002 | 13 | 13 | 616
78 | 251.173 | 3.9262 | 3.9272 | 33.8215 | 33.8261 | 7.3157 | 1.1549 | 14 | 14 | 617
78 | 201.208 | 4.0197 | 4.0215 | 33.7407 | 33.745 | 7.3031 | 1.3627 | 15 | 15 | 618
78 | 176.22 | 4.3256 | 4.3278 | 33.664 | 33.6683 | 7.253 | 1.6492 | 16 | 16 | 619
78 | 151.922 | 4.5024 | 4.5036 | 33.4443 | 33.4473 | 7.2329 | 1.9787 | 17 | 17 | 620
78 | 126.302 | 4.0963 | 4.0965 | 32.8423 | 32.8459 | 7.3339 | 2.5811 | 18 | 18 | 621
78 | 101.342 | 4.7729 | 4.7794 | 32.7372 | 32.7414 | 7.2203 | 2.6722 | 19 | 19 | 622
78 | 76.518 | 5.1761 | 5.1779 | 32.713 | 32.7168 | 7.1523 | 2.7336 | 20 | 20 | 623
78 | 51.677 | 6.024 | 6.0349 | 32.7063 | 32.711 | 7.0108 | 2.7853 | 21 | 21 | 624
78 | 25.719 | 6.3269 | 6.33 | 32.7028 | 32.7071 | 6.9615 | 2.8111 | 22 | 22 | 625
78 | 11.895 | 6.3781 | 6.379 | 32.7025 | 32.7068 | 6.9532 | 2.8175 | 23 | 23 | 626
78 | 5.981 | 6.3942 | 6.3956 | 32.7025 | 32.7067 | 6.9506 | 2.8189 | 24 | 24 | 627
107 | 1322.816 | 2.8818 | 2.881 | 34.4829 | 34.4881 | 7.4717 | 0.63446 | 1 | 1 | 812
107 | 1251.456 | 3.0697 | 3.0716 | 34.4625 | 34.4674 | 7.4382 | 0.61454 | 2 | 2 | 813
107 | 1251.613 | 3.0717 | 3.0724 | 34.4623 | 34.4673 | 7.4378 | 0.61438 | 3 | 3 | 814
107 | 1001.104 | 3.6378 | 3.6384 | 34.388 | 34.393 | 7.339 | 0.57213 | 4 | 4 | 815
107 | 800.793 | 4.1386 | 4.1398 | 34.3057 | 34.3106 | 7.2544 | 0.56391 | 5 | 5 | 816
107 | 601.432 | 4.7601 | 4.7614 | 34.1797 | 34.1845 | 7.1529 | 0.61086 | 6 | 6 | 817
107 | 401.461 | 5.3215 | 5.3231 | 34.0016 | 34.0063 | 7.0665 | 0.86084 | 7 | 7 | 818
107 | 300.999 | 5.7827 | 5.7856 | 33.9263 | 33.9308 | 6.9937 | 1.1346 | 8 | 8 | 819
107 | 251.741 | 6.221 | 6.2261 | 33.8918 | 33.8966 | 6.924 | 1.3508 | 9 | 9 | 820
107 | 200.454 | 6.7389 | 6.7484 | 33.863 | 33.8672 | 6.8427 | 1.5582 | 10 | 10 | 821
107 | 176.362 | 7.0276 | 7.0302 | 33.852 | 33.8568 | 6.7979 | 1.5304 | 11 | 11 | 822
107 | 151.726 | 7.2478 | 7.2499 | 33.7813 | 33.786 | 6.7669 | 1.6817 | 12 | 12 | 823
107 | 125.921 | 7.461 | 7.468 | 33.6701 | 33.6725 | 6.739 | 1.7091 | 13 | 13 | 824
107 | 101.735 | 7.7031 | 7.7032 | 33.2876 | 33.2892 | 6.7188 | 1.9066 | 14 | 14 | 825
107 | 76.997 | 7.4897 | 7.4922 | 32.6758 | 32.6791 | 6.7786 | 2.6578 | 15 | 15 | 826
107 | 51.647 | 8.266 | 8.2675 | 32.6015 | 32.6056 | 6.6637 | 2.6965 | 16 | 16 | 827
107 | 26.754 | 10.7009 | 10.7052 | 31.8401 | 31.8439 | 6.3467 | 2.8655 | 17 | 17 | 828
107 | 11.461 | 10.8242 | 10.8265 | 31.8205 | 31.8247 | 6.3306 | 2.8821 | 18 | 20 | 829

## hysteresis
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
78 | 4317.031 | 1.5199 | 1.5202 | 34.6775 | 34.6843 | 7.7205 | 1.0249 | 1 | 1 | 604 | 13.8339 | 0.2324 | 3.215 |  | 0.0024013 | 0.4497 | -0.2261
78 | 4002.275 | 1.5035 | 1.5039 | 34.6742 | 34.6809 | 7.7239 | 1.0365 | 2 | 2 | 605 | 13.2623 | 0.23767 | 3.152 |  | 0.0023662 |  |  | Summary of Soc Voffset including the original values in the factory calibration when E=0.036
46 | 3686.661 | 1.5689 | 1.5692 | 34.6626 | 34.6691 | 7.7117 | 0.99288 | 1 | 1 | 370 | 12.6839 | 0.22028 | 2.794 |  | 0.00011931 |  |  |  |  | m | b | Soc | Voffset | R2
46 | 3502.546 | 1.5723 | 1.5727 | 34.6597 | 34.6661 | 7.7112 | 0.99106 | 2 | 2 | 371 | 12.3702 | 0.21794 | 2.696 |  | 0.0016371 |  |  | Bottles used | Original | 0.4278 | -0.2186 | 0.4278 | -0.5109 |  | 2.776
46 | 3502.228 | 1.5725 | 1.5729 | 34.6597 | 34.666 | 7.7112 | 0.9908 | 3 | 3 | 372 | 12.3696 | 0.21836 | 2.701 | 6 | 0.0011053 |  |  | 168 | Using 2012-15 fit | 0.449 | -0.2235 | 0.449 | -0.49777 | 0.988
78 | 3501.917 | 1.5109 | 1.5115 | 34.6659 | 34.6722 | 7.7229 | 1.033 | 3 | 3 | 606 | 12.3903 | 0.23656 | 2.931 |  | 0.0019013 |  |  | 158 | all except cast #1 | 0.4491 | -0.2237 | 0.4491 | -0.49811 | 0.998
61 | 3501.057 | 1.5249 | 1.5253 | 34.6648 | 34.6712 | 7.7202 | 1.0227 | 3 | 3 | 486 | 12.384 | 0.23232 | 2.877 | 2 | 0.0015102 |  |  | 157 | excluding cast #1 and 1 outlier | 0.4493 | -0.2253 | 0.4493 | -0.50145 | 0.9997 | 2.92
61 | 3500.216 | 1.5245 | 1.525 | 34.6648 | 34.6712 | 7.7203 | 1.0229 | 2 | 2 | 485 | 12.3827 | 0.23291 | 2.884 | 2 | 0.0010014 |  |  | 148 | excluding diff >0.01 | 0.4497 | -0.2261 | 0.4497 | -0.50278 | 0.9999
29 | 3276.45 | 1.6715 | 1.6718 | 34.6431 | 34.6486 | 7.6928 | 0.93449 | 1 | 1 | 232 | 11.9645 | 0.19407 | 2.322 |  | 6.6721e-05 |  |  | 137 | excluding diff >0.005 | 0.4494 | -0.2261 | 0.4494 | -0.50312 | 1
46 | 3002.397 | 1.5941 | 1.5945 | 34.6494 | 34.6554 | 7.7075 | 0.9795 | 4 | 4 | 373 | 11.5529 | 0.2125 | 2.455 |  | 0.0018801 |  |  | 129 | excluding diff >0.004 | 0.4492 | -0.2261 | 0.4492 | -0.50334 | 1
29 | 3001.327 | 1.6482 | 1.6485 | 34.6426 | 34.648 | 7.6973 | 0.94797 | 2 | 2 | 233 | 11.5342 | 0.19863 | 2.291 |  | 0.0015751 |  |  | 111 | excluding diff >0.003 | 0.4492 | -0.2264 | 0.4492 | -0.50401 | 1
78 | 3001.236 | 1.5711 | 1.5717 | 34.6506 | 34.6567 | 7.7119 | 0.99611 | 4 | 4 | 607 | 11.5585 | 0.2201 | 2.544 |  | 0.0017521 |  |  |  | used for recal of 2012-26 | 0.4483 | -0.2246 | 0.4483 | -0.501
61 | 3001.222 | 1.5782 | 1.5787 | 34.6498 | 34.6559 | 7.7106 | 0.99087 | 4 | 4 | 487 | 11.5562 | 0.21772 | 2.516 |  | 0.0017755
29 | 3001.159 | 1.6484 | 1.6487 | 34.6425 | 34.648 | 7.6973 | 0.94785 | 3 | 3 | 234 | 11.5339 | 0.19855 | 2.29 | 6 | 0.0016022
29 | 2501.274 | 1.7508 | 1.7511 | 34.6226 | 34.6278 | 7.6785 | 0.89666 | 4 | 4 | 235 | 10.749 | 0.17555 | 1.887 |  | 0.0015767 |  |  | Summary of Hysteresis Study - fits using different E values above and below DO minimum
61 | 2501.023 | 1.7224 | 1.723 | 34.6232 | 34.629 | 7.684 | 0.90406 | 5 | 5 | 488 | 10.7571 | 0.17774 | 1.912 | 6 | 0.0027127 |  |  |  | 125-400db |  | 2000-4400db
78 | 2500.819 | 1.7084 | 1.7091 | 34.6234 | 34.6293 | 7.6867 | 0.91238 | 5 | 5 | 608 | 10.761 | 0.18223 | 1.961 | 6 | 0.0019646 |  |  | E | m | b | m | b | m diff | b diff
46 | 2500.769 | 1.7311 | 1.7315 | 34.6243 | 34.6301 | 7.6822 | 0.90277 | 5 | 5 | 374 | 10.754 | 0.17789 | 1.913 |  | 0.001989 |  |  | 3.5 | 0.4497 | -0.2269 | 0.4712 | -0.2428 | -0.0215 | 0.0159
78 | 2002.265 | 1.9266 | 1.9274 | 34.5807 | 34.5863 | 7.6467 | 0.79291 | 6 | 6 | 609 | 10.0006 | 0.12779 | 1.278 |  | 0.0026795 |  |  | 3.6 | 0.4495 | -0.227 | 0.4612 | -0.2357 | -0.0117 | 0.0087
22 | 2002.255 | 1.9168 | 1.9172 | 34.5893 | 34.594 | 7.6482 | 0.82239 | 1 | 1 | 186 | 10.0027 | 0.14616 | 1.462 | 6 | 0.0024318 |  |  | 3.65 | 0.4494 | -0.227 | 0.4563 | -0.2322 | -0.0069 | 0.0052
61 | 2001.892 | 1.9376 | 1.9382 | 34.5821 | 34.5878 | 7.6445 | 0.79466 | 6 | 6 | 489 | 9.997 | 0.12894 | 1.289 |  | 0.00232 |  |  | 3.7 | 0.4492 | -0.227 | 0.4514 | -0.2287 | -0.0022 | 0.0017
46 | 2001.08 | 1.9486 | 1.9491 | 34.5822 | 34.5879 | 7.6424 | 0.79323 | 6 | 6 | 375 | 9.9929 | 0.12879 | 1.287 |  | 0.0018239 |  |  | 3.705 | 0.4492 | -0.227 | 0.451 | -0.2286 | -0.0018 | 0.0016
29 | 2000.287 | 1.9531 | 1.9537 | 34.5824 | 34.5873 | 7.6415 | 0.80823 | 5 | 5 | 236 | 9.9905 | 0.13543 | 1.353 |  | 0.001933 |  |  | 3.71 | 0.4492 | -0.227 | 0.4504 | -0.228 | -0.0012 | 0.001
29 | 401.365 | 4.552 | 4.553 | 33.9511 | 33.9552 | 7.1997 | 0.96797 | 11 | 11 | 242 | 7.554 | 0.20651 | 1.56 |  | 0.0026831 |  |  | 3.715 | 0.4492 | -0.227 | 0.4499 | -0.2276 | -0.0007 | 0.0006
61 | 401.146 | 4.1999 | 4.2012 | 34.0341 | 34.0391 | 7.2569 | 0.88082 | 12 | 12 | 495 | 7.6171 | 0.16804 | 1.28 |  | 0.0019617 |  |  | 3.72 | 0.4492 | -0.227 | 0.4495 | -0.2273 | -0.0003 | 0.0003
46 | 302.467 | 4.5866 | 4.5879 | 33.89 | 33.8948 | 7.1967 | 1.1016 | 13 | 13 | 382 | 7.451 | 0.26748 | 1.993 |  | 0.0018008 |  |  | 3.725 | 0.4492 | -0.227 | 0.449 | -0.2269 | 0.0002 | -0.0001
15 | 302.342 | 5.8263 | 5.8276 | 33.996 | 34.0052 | 6.9833 | 0.90861 | 8 | 8 | 124 | 7.2202 | 0.1831 | 1.322 |  | 0.00059462 |  |  | 3.73 | 0.4492 | -0.227 | 0.4485 | -0.2266 | 0.0007 | -0.0004
78 | 302.137 | 3.8614 | 3.8624 | 33.8907 | 33.8954 | 7.3238 | 1.0002 | 13 | 13 | 616 | 7.5888 | 0.22019 | 1.671 |  | 0.0034894 |  |  | 3.74 | 0.4491 | -0.227 | 0.4475 | -0.2259 | 0.0016 | -0.0011
29 | 301.367 | 5.146 | 5.1476 | 33.8696 | 33.8739 | 7.1022 | 1.276 | 12 | 12 | 243 | 7.3475 | 0.34488 | 2.534 | 6 | 0.0028266 |  |  | 3.75 | 0.4491 | -0.227 | 0.4466 | -0.2252 | 0.0025 | -0.0018
22 | 300.937 | 5.6799 | 5.6807 | 33.9472 | 33.9514 | 7.0096 | 1.0519 | 8 | 8 | 193 | 7.2472 | 0.25058 | 1.816 |  | 0.0036415 |  |  | 3.8 | 0.449 | -0.227 | 0.4417 | -0.2218 | 0.0073 | -0.0052
61 | 298.794 | 4.5171 | 4.5175 | 33.9343 | 33.9397 | 7.2066 | 1.0818 | 13 | 13 | 496 | 7.4582 | 0.25744 | 1.92 |  | 0.0029411
22 | 251.956 | 6.0844 | 6.0859 | 33.9039 | 33.9078 | 6.9455 | 1.2889 | 9 | 9 | 194 | 7.1312 | 0.3559 | 2.538 |  | 0.0023862
61 | 251.5 | 4.813 | 4.8144 | 33.8868 | 33.8916 | 7.158 | 1.2312 | 14 | 14 | 497 | 7.3586 | 0.32438 | 2.387 |  | 0.0032055
29 | 251.355 | 5.81 | 5.8109 | 33.8664 | 33.8708 | 6.992 | 1.4025 | 13 | 13 | 244 | 7.1803 | 0.40347 | 2.897 |  | 0.0011346
46 | 251.243 | 4.8512 | 4.8527 | 33.8558 | 33.8607 | 7.1529 | 1.2501 | 14 | 14 | 383 | 7.3528 | 0.33212 | 2.442 |  | 0.0039443
78 | 251.173 | 3.9262 | 3.9272 | 33.8215 | 33.8261 | 7.3157 | 1.1549 | 14 | 14 | 617 | 7.5281 | 0.29171 | 2.196 |  | 0.0015599

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
24 | 75.865 | 7.7505 | 6.231 | 33.9906 | 33.9893 | 6.9179 | 1.0309 | 6 | 6 | 88 | 6.9559 | 0.24411 | 1.698 |  | 0.00023143 | 0.465 | -0.2355 |  |  |  | -88 | Summary of Soc Voffset including the original values in the factory calibration
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers