## File List
Number of Files= |  | 9
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 11 | 20 | 2012-12-0001.mrgcor1 & 2012-12-0001.samcor1
2 | 2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 22 | 29 | 2012-12-0003.mrgcor1 & 2012-12-0003.samcor1
3 | 2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 31 | 49 | 2012-12-0015.mrgcor1 & 2012-12-0015.samcor1
4 | 2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 51 | 69 | 2012-12-0022.mrgcor1 & 2012-12-0022.samcor1
5 | 2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 71 | 93 | 2012-12-0029.mrgcor1 & 2012-12-0029.samcor1
6 | 2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 95 | 117 | 2012-12-0046.mrgcor1 & 2012-12-0046.samcor1
7 | 2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 119 | 142 | 2012-12-0061.mrgcor1 & 2012-12-0061.samcor1
8 | 2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 144 | 167 | 2012-12-0078.mrgcor1 & 2012-12-0078.samcor1
9 | 2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 169 | 185 | 2012-12-0107.mrgcor1 & 2012-12-0107.samcor1

## (DOX_CTD - DOX_BOT)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.020846 |  | X | CALC Y
m= | 2.905e-06 | 1.907e-06 |  | Std Dev of Resid |  | 0.018346 |  | 5.9814 | -0.0030272
b= | -0.0030445 | 0.0023354 |  | Resid Sum of Squares |  | 0.036687 |  | 3502.36 | 0.0071298
# of Values Extracted = |  |  | 167
# of Values Used in Fit = |  |  | 111
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 202.433 | 0.062834 | 0.0021089 | 0.062834 | -0.0024565 | -0.06529 |  | 106.599 | -0.024784
1 | NO | 177.047 | 0.060187 | 0.0019628 | 0.060187 | -0.0025302 | -0.062717 |  | 103.39 | -0.060693
1 | NO | 150.761 | 0.058689 | 0.0021111 | 0.010689 | -0.0026066 | -0.013295 |  | 53.4026 | -0.015381
1 | NO | 125.723 | 0.12198 | 0.0018325 | -0.079017 | -0.0026793 | 0.076337 |  | 1317.45 | -0.016112
1 | NO | 101.293 | 1.3448 | 0.019998 | -0.093245 | -0.0027503 | 0.090495 |  | 1251.86 | -0.0056451
1 | NO | 76.5279 | 3.0528 | 0.0071636 | 0.023759 | -0.0028222 | -0.026582 |  | 1251.86 | -0.0056451
1 | NO | 51.6656 | 4.5748 | 0.0027081 | 0.032843 | -0.0028945 | -0.035737 |  | 1001.43 | -0.0083403
1 | NO | 26.5564 | 7.3571 | 0.0055187 | 0.23208 | -0.0029674 | -0.23505 |  | 802.22 | -0.0015187
1 | NO | 11.5966 | 8.2646 | 0.0077588 | -0.0033608 | -0.0030109 | 0.0003499 |  | 601.784 | -0.0021369
1 | NO | 6.3628 | 9.2006 | 0.018324 | 0.081617 | -0.0030261 | -0.084643 |  | 401.671 | -0.0088382
2 | YES | 106.599 | 2.0922 | 0.002509 | -0.024784 | -0.0027349 | 0.022049 |  | 302.342 | -0.012108
2 | YES | 103.39 | 2.1013 | 0.0016321 | -0.060693 | -0.0027442 | 0.057949 |  | 200.199 | 0.014398
2 | NO | 78.0801 | 3.7745 | 0.013375 | -0.048472 | -0.0028177 | 0.045655 |  | 151.95 | -0.0096681
2 | NO | 78.0801 | 3.7745 | 0.013375 | -0.056473 | -0.0028177 | 0.053655 |  | 52.2339 | 0.021963
2 | YES | 53.4026 | 5.4546 | 0.0040161 | -0.015381 | -0.0028894 | 0.012492 |  | 11.4485 | -0.0048299
2 | NO | 27.0918 | 6.2435 | 0.0089363 | -0.0015435 | -0.0029658 | -0.0014223 |  | 2002.25 | -0.035718
2 | NO | 11.4186 | 7.098 | 0.021015 | 0.11904 | -0.0030114 | -0.12205 |  | 1500.73 | -0.034647
2 | NO | 5.7948 | 6.8624 | 0.020127 | -0.0096517 | -0.0030277 | 0.006624 |  | 1251.76 | -0.038303
3 | YES | 1317.45 | 0.48189 | 0.003452 | -0.016112 | 0.00078264 | 0.016895 |  | 1001.2 | -0.010178
3 | YES | 1251.86 | 0.41935 | 0.003165 | -0.0056451 | 0.0005921 | 0.0062372 |  | 801.132 | 0.00089626
3 | YES | 1251.86 | 0.41935 | 0.003165 | -0.0056451 | 0.0005921 | 0.0062372 |  | 601.83 | -0.0020166
3 | YES | 1001.43 | 0.26666 | 0.0021604 | -0.0083403 | -0.00013539 | 0.0082049 |  | 402.232 | -0.019905
3 | YES | 802.22 | 0.20748 | 0.002023 | -0.0015187 | -0.0007141 | 0.00080458 |  | 251.956 | -0.025266
3 | YES | 601.784 | 0.33586 | 0.0018891 | -0.0021369 | -0.0012964 | 0.00084056 |  | 201.944 | -0.0040374
3 | YES | 401.671 | 0.86716 | 0.0025566 | -0.0088382 | -0.0018777 | 0.0069605 |  | 176.355 | -0.014137
3 | YES | 302.342 | 1.3099 | 0.0027863 | -0.012108 | -0.0021662 | 0.0099418 |  | 151.66 | -0.027618
3 | NO | 251.391 | 1.8208 | 0.0069165 | -0.051203 | -0.0023143 | 0.048889 |  | 127.302 | -0.0073528
3 | YES | 200.199 | 3.3764 | 0.0033688 | 0.014398 | -0.002463 | -0.016861 |  | 76.1556 | 0.010191
3 | NO | 176.451 | 3.2252 | 0.0072069 | 0.029178 | -0.002532 | -0.03171 |  | 3001.23 | 0.0046222
3 | YES | 151.95 | 3.1073 | 0.0037625 | -0.0096681 | -0.0026031 | 0.007065 |  | 3001.23 | 0.0056222
3 | NO | 125.765 | 3.6318 | 0.0079332 | -0.0011618 | -0.0026792 | -0.0015174 |  | 2501.27 | 0.0050747
3 | NO | 100.932 | 5.8981 | 0.020326 | 0.15714 | -0.0027513 | -0.15989 |  | 2000.29 | 0.0081742
3 | NO | 75.4452 | 5.9251 | 0.0063509 | -0.072938 | -0.0028254 | 0.070113 |  | 1501.41 | 0.014515
3 | YES | 52.2339 | 6.162 | 0.0042782 | 0.021963 | -0.0028928 | -0.024855 |  | 1251.7 | 0.0012614
3 | NO | 25.525 | 6.7452 | 0.01633 | -0.031809 | -0.0029704 | 0.028839 |  | 1001.35 | 0.0057344
3 | YES | 11.4485 | 6.7322 | 0.0041773 | -0.0048299 | -0.0030113 | 0.0018186 |  | 802.144 | 0.010423
3 | NO | 6.6003 | 6.73 | 0.0060944 | -0.006 | -0.0030254 | 0.0029747 |  | 600.811 | 0.0095809
4 | YES | 2002.25 | 1.4263 | 0.0038756 | -0.035718 | 0.002772 | 0.03849 |  | 401.365 | 0.011913
4 | YES | 1500.73 | 0.65835 | 0.0043422 | -0.034647 | 0.0013151 | 0.035962 |  | 301.367 | 0.011917
4 | YES | 1251.76 | 0.4217 | 0.0033819 | -0.038303 | 0.00059181 | 0.038895 |  | 251.355 | -0.0002284
4 | YES | 1001.2 | 0.24682 | 0.0029547 | -0.010178 | -0.00013606 | 0.010042 |  | 175.892 | 0.0044317
4 | YES | 801.132 | 0.2349 | 0.0024937 | 0.00089626 | -0.00071726 | -0.0016135 |  | 152.017 | 0.0078507
4 | YES | 601.83 | 0.35298 | 0.0018573 | -0.0020166 | -0.0012962 | 0.00072037 |  | 101.54 | 0.0091991
4 | YES | 402.232 | 1.1311 | 0.002883 | -0.019905 | -0.0018761 | 0.018029 |  | 75.454 | 0.017091
4 | NO | 300.937 | 1.783 | 0.0091142 | -0.033046 | -0.0021703 | 0.030875 |  | 51.2657 | 0.0028338
4 | YES | 251.956 | 2.5127 | 0.0036929 | -0.025266 | -0.0023126 | 0.022953 |  | 26.5587 | 0.0013404
4 | YES | 201.944 | 3.032 | 0.0019243 | -0.0040374 | -0.0024579 | 0.0015795 |  | 11.8022 | 0.0090666
4 | YES | 176.355 | 3.1059 | 0.0039647 | -0.014137 | -0.0025322 | 0.011605 |  | 5.9859 | 0.0085688
4 | YES | 151.66 | 3.4614 | 0.0050682 | -0.027618 | -0.002604 | 0.025014 |  | 3502.36 | 0.0052238
4 | YES | 127.302 | 3.8207 | 0.0030035 | -0.0073528 | -0.0026747 | 0.0046781 |  | 3502.36 | 0.00022388
4 | NO | 101.397 | 4.9615 | 0.05593 | -0.20851 | -0.00275 | 0.20576 |  | 3002.14 | 0.0091205
4 | YES | 76.1556 | 6.7162 | 0.0056514 | 0.010191 | -0.0028233 | -0.013014 |  | 2500.77 | 0.009083
4 | NO | 50.5981 | 6.9217 | 0.0070124 | 0.019676 | -0.0028976 | -0.022573 |  | 2001.08 | 0.0069668
4 | NO | 27.9794 | 6.9493 | 0.013347 | 0.027295 | -0.0029633 | -0.030258 |  | 1501.36 | 0.021295
4 | NO | 11.1393 | 6.7434 | 0.0078356 | 0.00043154 | -0.0030122 | -0.0034437 |  | 1251.13 | 0.017112
4 | NO | 6.8665 | 6.737 | 0.0072416 | 0.002985 | -0.0030246 | -0.0060096 |  | 1001.16 | 0.0069004
5 | NO | 3276.45 | 2.31 | 0.0062189 | -0.011992 | 0.0064735 | 0.018465 |  | 802.154 | 0.015855
5 | YES | 3001.23 | 2.2956 | 0.005312 | 0.0046222 | 0.005674 | 0.0010518 |  | 601.861 | -0.0023693
5 | YES | 3001.23 | 2.2956 | 0.005312 | 0.0056222 | 0.005674 | 5.1859e-05 |  | 402.365 | 0.029743
5 | YES | 2501.27 | 1.8921 | 0.0048402 | 0.0050747 | 0.0042216 | -0.00085311 |  | 302.555 | 0.007355
5 | YES | 2000.29 | 1.3612 | 0.0037454 | 0.0081742 | 0.0027663 | -0.0054079 |  | 251.243 | 0.022361
5 | YES | 1501.41 | 0.65451 | 0.0033753 | 0.014515 | 0.001317 | -0.013198 |  | 151.552 | 0.013211
5 | YES | 1251.7 | 0.34026 | 0.0029358 | 0.0012614 | 0.00059164 | -0.00066977 |  | 75.9301 | 0.011145
5 | YES | 1001.35 | 0.25073 | 0.0024639 | 0.0057344 | -0.00013563 | -0.0058701 |  | 50.9875 | 0.0094066
5 | YES | 802.144 | 0.32442 | 0.003151 | 0.010423 | -0.00071432 | -0.011138 |  | 11.6753 | 0.0095396
5 | YES | 600.811 | 0.58458 | 0.0026526 | 0.0095809 | -0.0012992 | -0.01088 |  | 6.1982 | -0.0031867
5 | YES | 401.365 | 1.5719 | 0.0032972 | 0.011913 | -0.0018786 | -0.013791 |  | 3500.59 | -0.0017133
5 | YES | 301.367 | 2.5459 | 0.0033643 | 0.011917 | -0.0021691 | -0.014086 |  | 3500.87 | 0.0049548
5 | YES | 251.355 | 2.8968 | 0.005109 | -0.0002284 | -0.0023144 | -0.002086 |  | 2501.02 | 0.017332
5 | NO | 200.723 | 3.7747 | 0.0025842 | 0.064676 | -0.0024614 | -0.067138 |  | 2001.89 | 0.011531
5 | YES | 175.892 | 4.0504 | 0.0050823 | 0.0044317 | -0.0025336 | -0.0069653 |  | 1502.39 | 0.011248
5 | YES | 152.017 | 4.3399 | 0.0033632 | 0.0078507 | -0.0026029 | -0.010454 |  | 1251.06 | 0.024419
5 | NO | 126.208 | 6.4345 | 0.014297 | 0.16955 | -0.0026779 | -0.17223 |  | 1001.25 | 0.019809
5 | YES | 101.54 | 6.8092 | 0.0041404 | 0.0091991 | -0.0027496 | -0.011949 |  | 800.214 | 0.021403
5 | YES | 75.454 | 6.8701 | 0.0039062 | 0.017091 | -0.0028253 | -0.019917 |  | 600.366 | 0.01768
5 | YES | 51.2657 | 6.9428 | 0.0043577 | 0.0028338 | -0.0028956 | -0.0057295 |  | 401.138 | 0.0060126
5 | YES | 26.5587 | 6.7873 | 0.004111 | 0.0013404 | -0.0029674 | -0.0043078 |  | 298.794 | 0.012664
5 | YES | 11.8022 | 6.7911 | 0.0040584 | 0.0090666 | -0.0030103 | -0.012077 |  | 251.5 | 0.014863
5 | YES | 5.9859 | 6.7926 | 0.0045795 | 0.0085688 | -0.0030272 | -0.011596 |  | 201.869 | 0.013892
6 | NO | 3686.66 | 2.7799 | 0.0062034 | -0.014133 | 0.0076652 | 0.021798 |  | 176.828 | 0.0022614
6 | YES | 3502.36 | 2.7012 | 0.0055509 | 0.0052238 | 0.0071298 | 0.001906 |  | 151.564 | -0.0087926
6 | YES | 3502.36 | 2.7012 | 0.0055509 | 0.00022388 | 0.0071298 | 0.0069059 |  | 26.5474 | 0.012361
6 | YES | 3002.14 | 2.4641 | 0.0051003 | 0.0091205 | 0.0056767 | -0.0034438 |  | 11.6412 | 0.011087
6 | YES | 2500.77 | 1.9221 | 0.0052964 | 0.009083 | 0.0042202 | -0.0048629 |  | 6.4089 | 0.010909
6 | YES | 2001.08 | 1.294 | 0.0042816 | 0.0069668 | 0.0027686 | -0.0041982 |  | 2500.82 | 0.0089959
6 | YES | 1501.36 | 0.52929 | 0.0029959 | 0.021295 | 0.0013169 | -0.019978 |  | 2002.27 | 0.015734
6 | YES | 1251.13 | 0.32311 | 0.0034436 | 0.017112 | 0.00058998 | -0.016522 |  | 1501.18 | 0.018826
6 | YES | 1001.16 | 0.2599 | 0.0031022 | 0.0069004 | -0.00013618 | -0.0070366 |  | 1250.9 | 0.0056888
6 | YES | 802.154 | 0.36185 | 0.0032519 | 0.015855 | -0.00071429 | -0.016569 |  | 1000.85 | 0.0082282
6 | YES | 601.861 | 0.52963 | 0.0020677 | -0.0023693 | -0.0012961 | 0.0010731 |  | 801.257 | 0.021095
6 | YES | 402.365 | 1.1407 | 0.0033028 | 0.029743 | -0.0018757 | -0.031618 |  | 601.704 | 0.0095851
6 | YES | 302.555 | 2.0004 | 0.0030079 | 0.007355 | -0.0021656 | -0.0095206 |  | 402.235 | 0.015909
6 | YES | 251.243 | 2.4644 | 0.0040495 | 0.022361 | -0.0023147 | -0.024676 |  | 302.137 | 0.017917
6 | NO | 202.859 | 2.9784 | 0.0081902 | -0.0016141 | -0.0024552 | -0.00084115 |  | 251.173 | 0.0028174
6 | NO | 176.96 | 3.3153 | 0.011363 | -0.029685 | -0.0025305 | 0.027154 |  | 126.302 | 0.0029502
6 | YES | 151.552 | 3.7322 | 0.0034401 | 0.013211 | -0.0026043 | -0.015816 |  | 51.6772 | -0.021124
6 | NO | 101.254 | 6.4235 | 0.054169 | -0.058473 | -0.0027504 | 0.055723 |  | 25.7193 | -0.0039334
6 | YES | 75.9301 | 6.9652 | 0.0043836 | 0.011145 | -0.002824 | -0.013969 |  | 11.8952 | -0.055062
6 | YES | 50.9875 | 7.1214 | 0.0046351 | 0.0094066 | -0.0028964 | -0.012303 |  | 5.9814 | -0.053606
6 | NO | 27.6139 | 7.0217 | 0.0082218 | 0.023747 | -0.0029643 | -0.026711 |  | 1322.82 | -0.025431
6 | YES | 11.6753 | 6.9645 | 0.0034999 | 0.0095396 | -0.0030106 | -0.01255 |  | 1251.59 | 0.0039075
6 | YES | 6.1982 | 6.9648 | 0.0038246 | -0.0031867 | -0.0030265 | 0.00016016 |  | 1251.59 | -0.035092
7 | NO | 4024.14 | 3.0274 | 0.0076419 | -0.13661 | 0.0086456 | 0.14525 |  | 1001.1 | -0.015087
7 | YES | 3500.59 | 2.8823 | 0.0058304 | -0.0017133 | 0.0071246 | 0.0088379 |  | 800.793 | -0.018041
7 | YES | 3500.87 | 2.8819 | 0.0058644 | 0.0049548 | 0.0071255 | 0.0021706 |  | 601.432 | -0.013938
7 | NO | 3001.22 | 2.524 | 0.0072739 | 0.008054 | 0.005674 | -0.00238 |  | 401.461 | -0.01283
7 | YES | 2501.02 | 1.9293 | 0.0044175 | 0.017332 | 0.0042209 | -0.013111 |  | 300.999 | -0.055394
7 | YES | 2001.89 | 1.3005 | 0.0033429 | 0.011531 | 0.0027709 | -0.0087602 |  | 151.726 | 0.0047178
7 | YES | 1502.39 | 0.54225 | 0.0034388 | 0.011248 | 0.0013199 | -0.0099279 |  | 125.921 | -0.029589
7 | YES | 1251.06 | 0.32742 | 0.0029486 | 0.024419 | 0.00058978 | -0.023829 |  | 51.6466 | -0.027386
7 | YES | 1001.25 | 0.35381 | 0.0027893 | 0.019809 | -0.00013592 | -0.019945 |  | 11.4606 | -0.037116 | -0.00067381 | 0.018456
7 | YES | 800.214 | 0.4774 | 0.0032506 | 0.021403 | -0.00071993 | -0.022122
7 | YES | 600.366 | 0.63268 | 0.0021 | 0.01768 | -0.0013005 | -0.018981
7 | YES | 401.138 | 1.286 | 0.0028291 | 0.0060126 | -0.0018792 | -0.0078918
7 | YES | 298.794 | 1.9327 | 0.0039454 | 0.012664 | -0.0021765 | -0.014841
7 | YES | 251.5 | 2.4019 | 0.0027542 | 0.014863 | -0.0023139 | -0.017177
7 | YES | 201.869 | 3.1099 | 0.0023356 | 0.013892 | -0.0024581 | -0.01635
7 | YES | 176.828 | 3.3583 | 0.0035358 | 0.0022614 | -0.0025309 | -0.0047923
7 | YES | 151.564 | 3.7042 | 0.0029281 | -0.0087926 | -0.0026042 | 0.0061884
7 | NO | 126.859 | 4.6478 | 0.021172 | -0.020237 | -0.002676 | 0.017561
7 | NO | 101.395 | 6.5194 | 0.0086317 | 0.021394 | -0.00275 | -0.024144
7 | NO | 76.0086 | 7.0972 | 0.0060386 | 0.011166 | -0.0028237 | -0.01399
7 | NO | 50.9488 | 7.0676 | 0.0083551 | -0.010398 | -0.0028965 | 0.0075019
7 | YES | 26.5474 | 7.0424 | 0.0043222 | 0.012361 | -0.0029674 | -0.015328
7 | YES | 11.6412 | 7.0431 | 0.0036704 | 0.011087 | -0.0030107 | -0.014098
7 | YES | 6.4089 | 7.0409 | 0.0042983 | 0.010909 | -0.0030259 | -0.013935
8 | NO | 4317.03 | 3.2331 | 0.0069299 | 0.018121 | 0.0094964 | -0.0086241
8 | NO | 4002.27 | 3.1681 | 0.0085231 | 0.016054 | 0.008582 | -0.0074719
8 | NO | 3501.92 | 2.9406 | 0.009179 | 0.009614 | 0.0071285 | -0.0024855
8 | NO | 3001.24 | 2.5513 | 0.0065556 | 0.0073197 | 0.005674 | -0.0016457
8 | YES | 2500.82 | 1.97 | 0.0042323 | 0.0089959 | 0.0042203 | -0.0047756
8 | YES | 2002.27 | 1.2937 | 0.0044819 | 0.015734 | 0.002772 | -0.012962
8 | YES | 1501.18 | 0.52983 | 0.0036002 | 0.018826 | 0.0013164 | -0.017509
8 | YES | 1250.9 | 0.50269 | 0.0031102 | 0.0056888 | 0.00058931 | -0.0050995
8 | YES | 1000.85 | 0.51323 | 0.0029864 | 0.0082282 | -0.00013708 | -0.0083653
8 | YES | 801.257 | 0.5591 | 0.0026683 | 0.021095 | -0.0007169 | -0.021812
8 | YES | 601.704 | 0.66158 | 0.0028697 | 0.0095851 | -0.0012966 | -0.010882
8 | YES | 402.235 | 1.1879 | 0.0026426 | 0.015909 | -0.0018761 | -0.017785
8 | YES | 302.137 | 1.6889 | 0.0018032 | 0.017917 | -0.0021668 | -0.020084
8 | YES | 251.173 | 2.1988 | 0.0034904 | 0.0028174 | -0.0023149 | -0.0051323
8 | NO | 201.208 | 2.8793 | 0.0061589 | -0.014689 | -0.00246 | 0.012229
8 | NO | 176.22 | 3.7989 | 0.006781 | -0.0010705 | -0.0025326 | -0.0014621
8 | NO | 151.922 | 4.8635 | 0.0074055 | -0.036531 | -0.0026032 | 0.033928
8 | YES | 126.302 | 6.9249 | 0.0056086 | 0.0029502 | -0.0026776 | -0.0056278
8 | NO | 101.342 | 7.0887 | 0.0072418 | 0.0026803 | -0.0027501 | -0.0054304
8 | NO | 76.5176 | 7.1932 | 0.0067864 | -0.011834 | -0.0028223 | 0.0090118
8 | YES | 51.6772 | 7.1839 | 0.0039677 | -0.021124 | -0.0028944 | 0.01823
8 | YES | 25.7193 | 7.1891 | 0.0053389 | -0.0039334 | -0.0029698 | 0.0009636
8 | YES | 11.8952 | 7.1859 | 0.0027614 | -0.055062 | -0.00301 | 0.052052
8 | YES | 5.9814 | 7.1824 | 0.0024148 | -0.053606 | -0.0030272 | 0.050579
9 | YES | 1322.82 | 0.51757 | 0.0032946 | -0.025431 | 0.00079824 | 0.02623
9 | YES | 1251.59 | 0.42991 | 0.0031702 | 0.0039075 | 0.00059132 | -0.0033162
9 | YES | 1251.59 | 0.42991 | 0.0031702 | -0.035092 | 0.00059132 | 0.035684
9 | YES | 1001.1 | 0.25091 | 0.0022072 | -0.015087 | -0.00013635 | 0.014951
9 | YES | 800.793 | 0.21196 | 0.0027261 | -0.018041 | -0.00071824 | 0.017323
9 | YES | 601.432 | 0.36606 | 0.0026045 | -0.013938 | -0.0012974 | 0.01264
9 | YES | 401.461 | 1.1842 | 0.0029468 | -0.01283 | -0.0018783 | 0.010952
9 | YES | 300.999 | 2.0466 | 0.0030711 | -0.055394 | -0.0021701 | 0.053224
9 | NO | 200.454 | 3.3028 | 0.0039038 | -0.14624 | -0.0024622 | 0.14377
9 | NO | 176.362 | 3.1839 | 0.014763 | -0.035079 | -0.0025322 | 0.032547
9 | YES | 151.726 | 3.6227 | 0.00489 | 0.0047178 | -0.0026038 | -0.0073216
9 | YES | 125.921 | 3.6784 | 0.0033531 | -0.029589 | -0.0026787 | 0.02691
9 | NO | 101.735 | 4.2717 | 0.0082601 | -0.28225 | -0.002749 | 0.2795
9 | NO | 76.9966 | 6.5761 | 0.0098289 | -0.02093 | -0.0028209 | 0.018109
9 | YES | 51.6466 | 6.5506 | 0.0057725 | -0.027386 | -0.0028945 | 0.024491
9 | NO | 26.4404 | 6.6889 | 0.0073437 | -0.040123 | -0.0029677 | 0.037156
9 | YES | 11.4606 | 6.7059 | 0.0055583 | -0.037116 | -0.0030112 | 0.034105
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
NO
NO
NO
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
NO
NO
YES
NO
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (2)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0071884 |  | X | CALC Y
m= | -0.00064218 | 0.00072287 |  | Std Dev of Resid |  | 0.018474 |  | 0.20748 | 0.00091876
b= | 0.001052 | 0.0026169 |  | Resid Sum of Squares |  | 0.037198 |  | 7.1891 | -0.0035647
# of Values Extracted = |  |  | 167
# of Values Used in Fit = |  |  | 111
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 202.433 | 0.062834 | 0.0021089 | 0.062834 | 0.0010116 | -0.061822 |  | 2.0922 | -0.024784
1 | NO | 177.047 | 0.060187 | 0.0019628 | 0.060187 | 0.0010133 | -0.059173 |  | 2.1013 | -0.060693
1 | NO | 150.761 | 0.058689 | 0.0021111 | 0.010689 | 0.0010143 | -0.0096745 |  | 5.4546 | -0.015381
1 | NO | 125.723 | 0.12198 | 0.0018325 | -0.079017 | 0.00097366 | 0.07999 |  | 0.48189 | -0.016112
1 | NO | 101.293 | 1.3448 | 0.019998 | -0.093245 | 0.00018842 | 0.093433 |  | 0.41935 | -0.0056451
1 | NO | 76.5279 | 3.0528 | 0.0071636 | 0.023759 | -0.00090841 | -0.024668 |  | 0.41935 | -0.0056451
1 | NO | 51.6656 | 4.5748 | 0.0027081 | 0.032843 | -0.0018859 | -0.034728 |  | 0.26666 | -0.0083403
1 | NO | 26.5564 | 7.3571 | 0.0055187 | 0.23208 | -0.0036725 | -0.23575 |  | 0.20748 | -0.0015187
1 | NO | 11.5966 | 8.2646 | 0.0077588 | -0.0033608 | -0.0042554 | -0.00089461 |  | 0.33586 | -0.0021369
1 | NO | 6.3628 | 9.2006 | 0.018324 | 0.081617 | -0.0048564 | -0.086474 |  | 0.86716 | -0.0088382
2 | YES | 106.599 | 2.0922 | 0.002509 | -0.024784 | -0.00029158 | 0.024493 |  | 1.3099 | -0.012108
2 | YES | 103.39 | 2.1013 | 0.0016321 | -0.060693 | -0.00029742 | 0.060395 |  | 3.3764 | 0.014398
2 | NO | 78.0801 | 3.7745 | 0.013375 | -0.048472 | -0.0013719 | 0.0471 |  | 3.1073 | -0.0096681
2 | NO | 78.0801 | 3.7745 | 0.013375 | -0.056473 | -0.0013719 | 0.055101 |  | 6.162 | 0.021963
2 | YES | 53.4026 | 5.4546 | 0.0040161 | -0.015381 | -0.0024508 | 0.01293 |  | 6.7322 | -0.0048299
2 | NO | 27.0918 | 6.2435 | 0.0089363 | -0.0015435 | -0.0029574 | -0.0014139 |  | 1.4263 | -0.035718
2 | NO | 11.4186 | 7.098 | 0.021015 | 0.11904 | -0.0035062 | -0.12254 |  | 0.65835 | -0.034647
2 | NO | 5.7948 | 6.8624 | 0.020127 | -0.0096517 | -0.0033548 | 0.0062968 |  | 0.4217 | -0.038303
3 | YES | 1317.45 | 0.48189 | 0.003452 | -0.016112 | 0.00074254 | 0.016855 |  | 0.24682 | -0.010178
3 | YES | 1251.86 | 0.41935 | 0.003165 | -0.0056451 | 0.0007827 | 0.0064278 |  | 0.2349 | 0.00089626
3 | YES | 1251.86 | 0.41935 | 0.003165 | -0.0056451 | 0.0007827 | 0.0064278 |  | 0.35298 | -0.0020166
3 | YES | 1001.43 | 0.26666 | 0.0021604 | -0.0083403 | 0.00088075 | 0.009221 |  | 1.1311 | -0.019905
3 | YES | 802.22 | 0.20748 | 0.002023 | -0.0015187 | 0.00091876 | 0.0024374 |  | 2.5127 | -0.025266
3 | YES | 601.784 | 0.33586 | 0.0018891 | -0.0021369 | 0.00083631 | 0.0029732 |  | 3.032 | -0.0040374
3 | YES | 401.671 | 0.86716 | 0.0025566 | -0.0088382 | 0.00049512 | 0.0093333 |  | 3.1059 | -0.014137
3 | YES | 302.342 | 1.3099 | 0.0027863 | -0.012108 | 0.00021082 | 0.012319 |  | 3.4614 | -0.027618
3 | NO | 251.391 | 1.8208 | 0.0069165 | -0.051203 | -0.00011728 | 0.051086 |  | 3.8207 | -0.0073528
3 | YES | 200.199 | 3.3764 | 0.0033688 | 0.014398 | -0.0011162 | -0.015515 |  | 6.7162 | 0.010191
3 | NO | 176.451 | 3.2252 | 0.0072069 | 0.029178 | -0.0010191 | -0.030198 |  | 2.2956 | 0.0046222
3 | YES | 151.95 | 3.1073 | 0.0037625 | -0.0096681 | -0.00094346 | 0.0087247 |  | 2.2956 | 0.0056222
3 | NO | 125.765 | 3.6318 | 0.0079332 | -0.0011618 | -0.0012803 | -0.00011847 |  | 1.8921 | 0.0050747
3 | NO | 100.932 | 5.8981 | 0.020326 | 0.15714 | -0.0027356 | -0.15987 |  | 1.3612 | 0.0081742
3 | NO | 75.4452 | 5.9251 | 0.0063509 | -0.072938 | -0.0027529 | 0.070185 |  | 0.65451 | 0.014515
3 | YES | 52.2339 | 6.162 | 0.0042782 | 0.021963 | -0.0029051 | -0.024868 |  | 0.34026 | 0.0012614
3 | NO | 25.525 | 6.7452 | 0.01633 | -0.031809 | -0.0032796 | 0.028529 |  | 0.25073 | 0.0057344
3 | YES | 11.4485 | 6.7322 | 0.0041773 | -0.0048299 | -0.0032712 | 0.0015586 |  | 0.32442 | 0.010423
3 | NO | 6.6003 | 6.73 | 0.0060944 | -0.006 | -0.0032698 | 0.0027302 |  | 0.58458 | 0.0095809
4 | YES | 2002.25 | 1.4263 | 0.0038756 | -0.035718 | 0.00013607 | 0.035854 |  | 1.5719 | 0.011913
4 | YES | 1500.73 | 0.65835 | 0.0043422 | -0.034647 | 0.00062922 | 0.035277 |  | 2.5459 | 0.011917
4 | YES | 1251.76 | 0.4217 | 0.0033819 | -0.038303 | 0.00078119 | 0.039084 |  | 2.8968 | -0.0002284
4 | YES | 1001.2 | 0.24682 | 0.0029547 | -0.010178 | 0.00089349 | 0.011072 |  | 4.0504 | 0.0044317
4 | YES | 801.132 | 0.2349 | 0.0024937 | 0.00089626 | 0.00090115 | 4.8907e-06 |  | 4.3399 | 0.0078507
4 | YES | 601.83 | 0.35298 | 0.0018573 | -0.0020166 | 0.00082532 | 0.0028419 |  | 6.8092 | 0.0091991
4 | YES | 402.232 | 1.1311 | 0.002883 | -0.019905 | 0.00032563 | 0.02023 |  | 6.8701 | 0.017091
4 | NO | 300.937 | 1.783 | 0.0091142 | -0.033046 | -9.2972e-05 | 0.032953 |  | 6.9428 | 0.0028338
4 | YES | 251.956 | 2.5127 | 0.0036929 | -0.025266 | -0.00056162 | 0.024704 |  | 6.7873 | 0.0013404
4 | YES | 201.944 | 3.032 | 0.0019243 | -0.0040374 | -0.00089506 | 0.0031423 |  | 6.7911 | 0.0090666
4 | YES | 176.355 | 3.1059 | 0.0039647 | -0.014137 | -0.00094251 | 0.013194 |  | 6.7926 | 0.0085688
4 | YES | 151.66 | 3.4614 | 0.0050682 | -0.027618 | -0.0011708 | 0.026448 |  | 2.7012 | 0.0052238
4 | YES | 127.302 | 3.8207 | 0.0030035 | -0.0073528 | -0.0014015 | 0.0059513 |  | 2.7012 | 0.00022388
4 | NO | 101.397 | 4.9615 | 0.05593 | -0.20851 | -0.0021342 | 0.20638 |  | 2.4641 | 0.0091205
4 | YES | 76.1556 | 6.7162 | 0.0056514 | 0.010191 | -0.003261 | -0.013452 |  | 1.9221 | 0.009083
4 | NO | 50.5981 | 6.9217 | 0.0070124 | 0.019676 | -0.0033929 | -0.023069 |  | 1.294 | 0.0069668
4 | NO | 27.9794 | 6.9493 | 0.013347 | 0.027295 | -0.0034107 | -0.030705 |  | 0.52929 | 0.021295
4 | NO | 11.1393 | 6.7434 | 0.0078356 | 0.00043154 | -0.0032785 | -0.00371 |  | 0.32311 | 0.017112
4 | NO | 6.8665 | 6.737 | 0.0072416 | 0.002985 | -0.0032743 | -0.0062593 |  | 0.2599 | 0.0069004
5 | NO | 3276.45 | 2.31 | 0.0062189 | -0.011992 | -0.00043144 | 0.01156 |  | 0.36185 | 0.015855
5 | YES | 3001.23 | 2.2956 | 0.005312 | 0.0046222 | -0.0004222 | -0.0050444 |  | 0.52963 | -0.0023693
5 | YES | 3001.23 | 2.2956 | 0.005312 | 0.0056222 | -0.0004222 | -0.0060443 |  | 1.1407 | 0.029743
5 | YES | 2501.27 | 1.8921 | 0.0048402 | 0.0050747 | -0.00016305 | -0.0052378 |  | 2.0004 | 0.007355
5 | YES | 2000.29 | 1.3612 | 0.0037454 | 0.0081742 | 0.00017788 | -0.0079963 |  | 2.4644 | 0.022361
5 | YES | 1501.41 | 0.65451 | 0.0033753 | 0.014515 | 0.00063168 | -0.013883 |  | 3.7322 | 0.013211
5 | YES | 1251.7 | 0.34026 | 0.0029358 | 0.0012614 | 0.00083349 | -0.00042792 |  | 6.9652 | 0.011145
5 | YES | 1001.35 | 0.25073 | 0.0024639 | 0.0057344 | 0.00089098 | -0.0048435 |  | 7.1214 | 0.0094066
5 | YES | 802.144 | 0.32442 | 0.003151 | 0.010423 | 0.00084366 | -0.0095795 |  | 6.9645 | 0.0095396
5 | YES | 600.811 | 0.58458 | 0.0026526 | 0.0095809 | 0.00067659 | -0.0089043 |  | 6.9648 | -0.0031867
5 | YES | 401.365 | 1.5719 | 0.0032972 | 0.011913 | 4.2553e-05 | -0.01187 |  | 2.8823 | -0.0017133
5 | YES | 301.367 | 2.5459 | 0.0033643 | 0.011917 | -0.00058293 | -0.0125 |  | 2.8819 | 0.0049548
5 | YES | 251.355 | 2.8968 | 0.005109 | -0.0002284 | -0.00080824 | -0.00057984 |  | 1.9293 | 0.017332
5 | NO | 200.723 | 3.7747 | 0.0025842 | 0.064676 | -0.001372 | -0.066048 |  | 1.3005 | 0.011531
5 | YES | 175.892 | 4.0504 | 0.0050823 | 0.0044317 | -0.0015491 | -0.0059808 |  | 0.54225 | 0.011248
5 | YES | 152.017 | 4.3399 | 0.0033632 | 0.0078507 | -0.001735 | -0.0095856 |  | 0.32742 | 0.024419
5 | NO | 126.208 | 6.4345 | 0.014297 | 0.16955 | -0.0030801 | -0.17263 |  | 0.35381 | 0.019809
5 | YES | 101.54 | 6.8092 | 0.0041404 | 0.0091991 | -0.0033207 | -0.01252 |  | 0.4774 | 0.021403
5 | YES | 75.454 | 6.8701 | 0.0039062 | 0.017091 | -0.0033598 | -0.020451 |  | 0.63268 | 0.01768
5 | YES | 51.2657 | 6.9428 | 0.0043577 | 0.0028338 | -0.0034065 | -0.0062404 |  | 1.286 | 0.0060126
5 | YES | 26.5587 | 6.7873 | 0.004111 | 0.0013404 | -0.0033067 | -0.0046471 |  | 1.9327 | 0.012664
5 | YES | 11.8022 | 6.7911 | 0.0040584 | 0.0090666 | -0.0033091 | -0.012376 |  | 2.4019 | 0.014863
5 | YES | 5.9859 | 6.7926 | 0.0045795 | 0.0085688 | -0.00331 | -0.011879 |  | 3.1099 | 0.013892
6 | NO | 3686.66 | 2.7799 | 0.0062034 | -0.014133 | -0.00073317 | 0.0134 |  | 3.3583 | 0.0022614
6 | YES | 3502.36 | 2.7012 | 0.0055509 | 0.0052238 | -0.00068266 | -0.0059064 |  | 3.7042 | -0.0087926
6 | YES | 3502.36 | 2.7012 | 0.0055509 | 0.00022388 | -0.00068266 | -0.00090654 |  | 7.0424 | 0.012361
6 | YES | 3002.14 | 2.4641 | 0.0051003 | 0.0091205 | -0.0005304 | -0.0096509 |  | 7.0431 | 0.011087
6 | YES | 2500.77 | 1.9221 | 0.0052964 | 0.009083 | -0.00018232 | -0.0092653 |  | 7.0409 | 0.010909
6 | YES | 2001.08 | 1.294 | 0.0042816 | 0.0069668 | 0.00022104 | -0.0067458 |  | 1.97 | 0.0089959
6 | YES | 1501.36 | 0.52929 | 0.0029959 | 0.021295 | 0.00071209 | -0.020583 |  | 1.2937 | 0.015734
6 | YES | 1251.13 | 0.32311 | 0.0034436 | 0.017112 | 0.0008445 | -0.016267 |  | 0.52983 | 0.018826
6 | YES | 1001.16 | 0.2599 | 0.0031022 | 0.0069004 | 0.00088509 | -0.0060153 |  | 0.50269 | 0.0056888
6 | YES | 802.154 | 0.36185 | 0.0032519 | 0.015855 | 0.00081962 | -0.015035 |  | 0.51323 | 0.0082282
6 | YES | 601.861 | 0.52963 | 0.0020677 | -0.0023693 | 0.00071188 | 0.0030812 |  | 0.5591 | 0.021095
6 | YES | 402.365 | 1.1407 | 0.0033028 | 0.029743 | 0.00031944 | -0.029423 |  | 0.66158 | 0.0095851
6 | YES | 302.555 | 2.0004 | 0.0030079 | 0.007355 | -0.00023259 | -0.0075876 |  | 1.1879 | 0.015909
6 | YES | 251.243 | 2.4644 | 0.0040495 | 0.022361 | -0.00053056 | -0.022892 |  | 1.6889 | 0.017917
6 | NO | 202.859 | 2.9784 | 0.0081902 | -0.0016141 | -0.00086065 | 0.00075344 |  | 2.1988 | 0.0028174
6 | NO | 176.96 | 3.3153 | 0.011363 | -0.029685 | -0.001077 | 0.028608 |  | 6.9249 | 0.0029502
6 | YES | 151.552 | 3.7322 | 0.0034401 | 0.013211 | -0.0013447 | -0.014556 |  | 7.1839 | -0.021124
6 | NO | 101.254 | 6.4235 | 0.054169 | -0.058473 | -0.003073 | 0.0554 |  | 7.1891 | -0.0039334
6 | YES | 75.9301 | 6.9652 | 0.0043836 | 0.011145 | -0.0034209 | -0.014566 |  | 7.1859 | -0.055062
6 | YES | 50.9875 | 7.1214 | 0.0046351 | 0.0094066 | -0.0035212 | -0.012928 |  | 7.1824 | -0.053606
6 | NO | 27.6139 | 7.0217 | 0.0082218 | 0.023747 | -0.0034572 | -0.027204 |  | 0.51757 | -0.025431
6 | YES | 11.6753 | 6.9645 | 0.0034999 | 0.0095396 | -0.0034205 | -0.01296 |  | 0.42991 | 0.0039075
6 | YES | 6.1982 | 6.9648 | 0.0038246 | -0.0031867 | -0.0034206 | -0.00023394 |  | 0.42991 | -0.035092
7 | NO | 4024.14 | 3.0274 | 0.0076419 | -0.13661 | -0.00089212 | 0.13571 |  | 0.25091 | -0.015087
7 | YES | 3500.59 | 2.8823 | 0.0058304 | -0.0017133 | -0.00079894 | 0.00091434 |  | 0.21196 | -0.018041
7 | YES | 3500.87 | 2.8819 | 0.0058644 | 0.0049548 | -0.00079872 | -0.0057535 |  | 0.36606 | -0.013938
7 | NO | 3001.22 | 2.524 | 0.0072739 | 0.008054 | -0.00056889 | -0.0086229 |  | 1.1842 | -0.01283
7 | YES | 2501.02 | 1.9293 | 0.0044175 | 0.017332 | -0.00018697 | -0.017519 |  | 2.0466 | -0.055394
7 | YES | 2001.89 | 1.3005 | 0.0033429 | 0.011531 | 0.00021683 | -0.011314 |  | 3.6227 | 0.0047178
7 | YES | 1502.39 | 0.54225 | 0.0034388 | 0.011248 | 0.00070378 | -0.010544 |  | 3.6784 | -0.029589
7 | YES | 1251.06 | 0.32742 | 0.0029486 | 0.024419 | 0.00084173 | -0.023577 |  | 6.5506 | -0.027386
7 | YES | 1001.25 | 0.35381 | 0.0027893 | 0.019809 | 0.00082479 | -0.018984 |  | 6.7059 | -0.037116
7 | YES | 800.214 | 0.4774 | 0.0032506 | 0.021403 | 0.00074542 | -0.020657
7 | YES | 600.366 | 0.63268 | 0.0021 | 0.01768 | 0.0006457 | -0.017035
7 | YES | 401.138 | 1.286 | 0.0028291 | 0.0060126 | 0.00022615 | -0.0057864
7 | YES | 298.794 | 1.9327 | 0.0039454 | 0.012664 | -0.00018911 | -0.012853
7 | YES | 251.5 | 2.4019 | 0.0027542 | 0.014863 | -0.00049042 | -0.015353
7 | YES | 201.869 | 3.1099 | 0.0023356 | 0.013892 | -0.0009451 | -0.014837
7 | YES | 176.828 | 3.3583 | 0.0035358 | 0.0022614 | -0.0011046 | -0.003366
7 | YES | 151.564 | 3.7042 | 0.0029281 | -0.0087926 | -0.0013268 | 0.0074659
7 | NO | 126.859 | 4.6478 | 0.021172 | -0.020237 | -0.0019327 | 0.018304
7 | NO | 101.395 | 6.5194 | 0.0086317 | 0.021394 | -0.0031346 | -0.024529
7 | NO | 76.0086 | 7.0972 | 0.0060386 | 0.011166 | -0.0035056 | -0.014672
7 | NO | 50.9488 | 7.0676 | 0.0083551 | -0.010398 | -0.0034866 | 0.0069118
7 | YES | 26.5474 | 7.0424 | 0.0043222 | 0.012361 | -0.0034704 | -0.015831
7 | YES | 11.6412 | 7.0431 | 0.0036704 | 0.011087 | -0.0034709 | -0.014558
7 | YES | 6.4089 | 7.0409 | 0.0042983 | 0.010909 | -0.0034695 | -0.014378
8 | NO | 4317.03 | 3.2331 | 0.0069299 | 0.018121 | -0.0010242 | -0.019145
8 | NO | 4002.27 | 3.1681 | 0.0085231 | 0.016054 | -0.00098245 | -0.017036
8 | NO | 3501.92 | 2.9406 | 0.009179 | 0.009614 | -0.00083639 | -0.01045
8 | NO | 3001.24 | 2.5513 | 0.0065556 | 0.0073197 | -0.0005864 | -0.0079061
8 | YES | 2500.82 | 1.97 | 0.0042323 | 0.0089959 | -0.00021309 | -0.009209
8 | YES | 2002.27 | 1.2937 | 0.0044819 | 0.015734 | 0.00022119 | -0.015513
8 | YES | 1501.18 | 0.52983 | 0.0036002 | 0.018826 | 0.00071175 | -0.018114
8 | YES | 1250.9 | 0.50269 | 0.0031102 | 0.0056888 | 0.00072918 | -0.0049596
8 | YES | 1000.85 | 0.51323 | 0.0029864 | 0.0082282 | 0.00072241 | -0.0075058
8 | YES | 801.257 | 0.5591 | 0.0026683 | 0.021095 | 0.00069296 | -0.020403
8 | YES | 601.704 | 0.66158 | 0.0028697 | 0.0095851 | 0.00062714 | -0.0089579
8 | YES | 402.235 | 1.1879 | 0.0026426 | 0.015909 | 0.00028915 | -0.01562
8 | YES | 302.137 | 1.6889 | 0.0018032 | 0.017917 | -3.2588e-05 | -0.01795
8 | YES | 251.173 | 2.1988 | 0.0034904 | 0.0028174 | -0.00036003 | -0.0031774
8 | NO | 201.208 | 2.8793 | 0.0061589 | -0.014689 | -0.00079703 | 0.013892
8 | NO | 176.22 | 3.7989 | 0.006781 | -0.0010705 | -0.0013876 | -0.00031709
8 | NO | 151.922 | 4.8635 | 0.0074055 | -0.036531 | -0.0020712 | 0.03446
8 | YES | 126.302 | 6.9249 | 0.0056086 | 0.0029502 | -0.003395 | -0.0063452
8 | NO | 101.342 | 7.0887 | 0.0072418 | 0.0026803 | -0.0035002 | -0.0061805
8 | NO | 76.5176 | 7.1932 | 0.0067864 | -0.011834 | -0.0035673 | 0.0082668
8 | YES | 51.6772 | 7.1839 | 0.0039677 | -0.021124 | -0.0035613 | 0.017563
8 | YES | 25.7193 | 7.1891 | 0.0053389 | -0.0039334 | -0.0035647 | 0.00036878
8 | YES | 11.8952 | 7.1859 | 0.0027614 | -0.055062 | -0.0035626 | 0.0515
8 | YES | 5.9814 | 7.1824 | 0.0024148 | -0.053606 | -0.0035604 | 0.050046 |  |  |  | -0.00067381 | 0.018456
9 | YES | 1322.82 | 0.51757 | 0.0032946 | -0.025431 | 0.00071963 | 0.026151
9 | YES | 1251.59 | 0.42991 | 0.0031702 | 0.0039075 | 0.00077592 | -0.0031316
9 | YES | 1251.59 | 0.42991 | 0.0031702 | -0.035092 | 0.00077592 | 0.035868
9 | YES | 1001.1 | 0.25091 | 0.0022072 | -0.015087 | 0.00089087 | 0.015978
9 | YES | 800.793 | 0.21196 | 0.0027261 | -0.018041 | 0.00091588 | 0.018957
9 | YES | 601.432 | 0.36606 | 0.0026045 | -0.013938 | 0.00081692 | 0.014755
9 | YES | 401.461 | 1.1842 | 0.0029468 | -0.01283 | 0.00029155 | 0.013121
9 | YES | 300.999 | 2.0466 | 0.0030711 | -0.055394 | -0.00026229 | 0.055132
9 | NO | 200.454 | 3.3028 | 0.0039038 | -0.14624 | -0.001069 | 0.14517
9 | NO | 176.362 | 3.1839 | 0.014763 | -0.035079 | -0.00099264 | 0.034086
9 | YES | 151.726 | 3.6227 | 0.00489 | 0.0047178 | -0.0012744 | -0.0059923
9 | YES | 125.921 | 3.6784 | 0.0033531 | -0.029589 | -0.0013102 | 0.028279
9 | NO | 101.735 | 4.2717 | 0.0082601 | -0.28225 | -0.0016912 | 0.28056
9 | NO | 76.9966 | 6.5761 | 0.0098289 | -0.02093 | -0.003171 | 0.017759
9 | YES | 51.6466 | 6.5506 | 0.0057725 | -0.027386 | -0.0031546 | 0.024231
9 | NO | 26.4404 | 6.6889 | 0.0073437 | -0.040123 | -0.0032434 | 0.03688
9 | YES | 11.4606 | 6.7059 | 0.0055583 | -0.037116 | -0.0032544 | 0.033862
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
NO
NO
NO
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
NO
NO
YES
NO
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (3)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.002222 |  | X | CALC Y
m= | 0.00044893 | 0.00091118 |  | Std Dev of Resid |  | 0.01852 |  | 2 | -0.0024291
b= | -0.0033269 | 0.0056646 |  | Resid Sum of Squares |  | 0.037385 |  | 9 | 0.00071342
# of Values Extracted = |  |  | 167
# of Values Used in Fit = |  |  | 111
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 202.433 | 0.062834 | 0.0021089 | 0.062834 | -0.002878 | -0.065712 |  | 2 | -0.024784
1 | NO | 177.047 | 0.060187 | 0.0019628 | 0.060187 | -0.002878 | -0.063065 |  | 2 | -0.060693
1 | NO | 150.761 | 0.058689 | 0.0021111 | 0.010689 | -0.002878 | -0.013567 |  | 2 | -0.015381
1 | NO | 125.723 | 0.12198 | 0.0018325 | -0.079017 | -0.002878 | 0.076139 |  | 3 | -0.016112
1 | NO | 101.293 | 1.3448 | 0.019998 | -0.093245 | -0.002878 | 0.090367 |  | 3 | -0.0056451
1 | NO | 76.5279 | 3.0528 | 0.0071636 | 0.023759 | -0.002878 | -0.026637 |  | 3 | -0.0056451
1 | NO | 51.6656 | 4.5748 | 0.0027081 | 0.032843 | -0.002878 | -0.035721 |  | 3 | -0.0083403
1 | NO | 26.5564 | 7.3571 | 0.0055187 | 0.23208 | -0.002878 | -0.23496 |  | 3 | -0.0015187
1 | NO | 11.5966 | 8.2646 | 0.0077588 | -0.0033608 | -0.002878 | 0.00048274 |  | 3 | -0.0021369
1 | NO | 6.3628 | 9.2006 | 0.018324 | 0.081617 | -0.002878 | -0.084495 |  | 3 | -0.0088382
2 | YES | 106.599 | 2.0922 | 0.002509 | -0.024784 | -0.0024291 | 0.022355 |  | 3 | -0.012108
2 | YES | 103.39 | 2.1013 | 0.0016321 | -0.060693 | -0.0024291 | 0.058264 |  | 3 | 0.014398
2 | NO | 78.0801 | 3.7745 | 0.013375 | -0.048472 | -0.0024291 | 0.046043 |  | 3 | -0.0096681
2 | NO | 78.0801 | 3.7745 | 0.013375 | -0.056473 | -0.0024291 | 0.054043 |  | 3 | 0.021963
2 | YES | 53.4026 | 5.4546 | 0.0040161 | -0.015381 | -0.0024291 | 0.012952 |  | 3 | -0.0048299
2 | NO | 27.0918 | 6.2435 | 0.0089363 | -0.0015435 | -0.0024291 | -0.00088556 |  | 4 | -0.035718
2 | NO | 11.4186 | 7.098 | 0.021015 | 0.11904 | -0.0024291 | -0.12147 |  | 4 | -0.034647
2 | NO | 5.7948 | 6.8624 | 0.020127 | -0.0096517 | -0.0024291 | 0.0072226 |  | 4 | -0.038303
3 | YES | 1317.45 | 0.48189 | 0.003452 | -0.016112 | -0.0019802 | 0.014132 |  | 4 | -0.010178
3 | YES | 1251.86 | 0.41935 | 0.003165 | -0.0056451 | -0.0019802 | 0.0036649 |  | 4 | 0.00089626
3 | YES | 1251.86 | 0.41935 | 0.003165 | -0.0056451 | -0.0019802 | 0.0036649 |  | 4 | -0.0020166
3 | YES | 1001.43 | 0.26666 | 0.0021604 | -0.0083403 | -0.0019802 | 0.0063601 |  | 4 | -0.019905
3 | YES | 802.22 | 0.20748 | 0.002023 | -0.0015187 | -0.0019802 | -0.00046147 |  | 4 | -0.025266
3 | YES | 601.784 | 0.33586 | 0.0018891 | -0.0021369 | -0.0019802 | 0.00015677 |  | 4 | -0.0040374
3 | YES | 401.671 | 0.86716 | 0.0025566 | -0.0088382 | -0.0019802 | 0.006858 |  | 4 | -0.014137
3 | YES | 302.342 | 1.3099 | 0.0027863 | -0.012108 | -0.0019802 | 0.010128 |  | 4 | -0.027618
3 | NO | 251.391 | 1.8208 | 0.0069165 | -0.051203 | -0.0019802 | 0.049223 |  | 4 | -0.0073528
3 | YES | 200.199 | 3.3764 | 0.0033688 | 0.014398 | -0.0019802 | -0.016378 |  | 4 | 0.010191
3 | NO | 176.451 | 3.2252 | 0.0072069 | 0.029178 | -0.0019802 | -0.031159 |  | 5 | 0.0046222
3 | YES | 151.95 | 3.1073 | 0.0037625 | -0.0096681 | -0.0019802 | 0.007688 |  | 5 | 0.0056222
3 | NO | 125.765 | 3.6318 | 0.0079332 | -0.0011618 | -0.0019802 | -0.00081834 |  | 5 | 0.0050747
3 | NO | 100.932 | 5.8981 | 0.020326 | 0.15714 | -0.0019802 | -0.15912 |  | 5 | 0.0081742
3 | NO | 75.4452 | 5.9251 | 0.0063509 | -0.072938 | -0.0019802 | 0.070958 |  | 5 | 0.014515
3 | YES | 52.2339 | 6.162 | 0.0042782 | 0.021963 | -0.0019802 | -0.023943 |  | 5 | 0.0012614
3 | NO | 25.525 | 6.7452 | 0.01633 | -0.031809 | -0.0019802 | 0.029829 |  | 5 | 0.0057344
3 | YES | 11.4485 | 6.7322 | 0.0041773 | -0.0048299 | -0.0019802 | 0.0028497 |  | 5 | 0.010423
3 | NO | 6.6003 | 6.73 | 0.0060944 | -0.006 | -0.0019802 | 0.0040199 |  | 5 | 0.0095809
4 | YES | 2002.25 | 1.4263 | 0.0038756 | -0.035718 | -0.0015312 | 0.034187 |  | 5 | 0.011913
4 | YES | 1500.73 | 0.65835 | 0.0043422 | -0.034647 | -0.0015312 | 0.033116 |  | 5 | 0.011917
4 | YES | 1251.76 | 0.4217 | 0.0033819 | -0.038303 | -0.0015312 | 0.036772 |  | 5 | -0.0002284
4 | YES | 1001.2 | 0.24682 | 0.0029547 | -0.010178 | -0.0015312 | 0.0086472 |  | 5 | 0.0044317
4 | YES | 801.132 | 0.2349 | 0.0024937 | 0.00089626 | -0.0015312 | -0.0024275 |  | 5 | 0.0078507
4 | YES | 601.83 | 0.35298 | 0.0018573 | -0.0020166 | -0.0015312 | 0.00048538 |  | 5 | 0.0091991
4 | YES | 402.232 | 1.1311 | 0.002883 | -0.019905 | -0.0015312 | 0.018373 |  | 5 | 0.017091
4 | NO | 300.937 | 1.783 | 0.0091142 | -0.033046 | -0.0015312 | 0.031514 |  | 5 | 0.0028338
4 | YES | 251.956 | 2.5127 | 0.0036929 | -0.025266 | -0.0015312 | 0.023734 |  | 5 | 0.0013404
4 | YES | 201.944 | 3.032 | 0.0019243 | -0.0040374 | -0.0015312 | 0.0025062 |  | 5 | 0.0090666
4 | YES | 176.355 | 3.1059 | 0.0039647 | -0.014137 | -0.0015312 | 0.012606 |  | 5 | 0.0085688
4 | YES | 151.66 | 3.4614 | 0.0050682 | -0.027618 | -0.0015312 | 0.026087 |  | 6 | 0.0052238
4 | YES | 127.302 | 3.8207 | 0.0030035 | -0.0073528 | -0.0015312 | 0.0058216 |  | 6 | 0.00022388
4 | NO | 101.397 | 4.9615 | 0.05593 | -0.20851 | -0.0015312 | 0.20698 |  | 6 | 0.0091205
4 | YES | 76.1556 | 6.7162 | 0.0056514 | 0.010191 | -0.0015312 | -0.011722 |  | 6 | 0.009083
4 | NO | 50.5981 | 6.9217 | 0.0070124 | 0.019676 | -0.0015312 | -0.021207 |  | 6 | 0.0069668
4 | NO | 27.9794 | 6.9493 | 0.013347 | 0.027295 | -0.0015312 | -0.028826 |  | 6 | 0.021295
4 | NO | 11.1393 | 6.7434 | 0.0078356 | 0.00043154 | -0.0015312 | -0.0019628 |  | 6 | 0.017112
4 | NO | 6.8665 | 6.737 | 0.0072416 | 0.002985 | -0.0015312 | -0.0045162 |  | 6 | 0.0069004
5 | NO | 3276.45 | 2.31 | 0.0062189 | -0.011992 | -0.0010823 | 0.010909 |  | 6 | 0.015855
5 | YES | 3001.23 | 2.2956 | 0.005312 | 0.0046222 | -0.0010823 | -0.0057045 |  | 6 | -0.0023693
5 | YES | 3001.23 | 2.2956 | 0.005312 | 0.0056222 | -0.0010823 | -0.0067044 |  | 6 | 0.029743
5 | YES | 2501.27 | 1.8921 | 0.0048402 | 0.0050747 | -0.0010823 | -0.006157 |  | 6 | 0.007355
5 | YES | 2000.29 | 1.3612 | 0.0037454 | 0.0081742 | -0.0010823 | -0.0092565 |  | 6 | 0.022361
5 | YES | 1501.41 | 0.65451 | 0.0033753 | 0.014515 | -0.0010823 | -0.015597 |  | 6 | 0.013211
5 | YES | 1251.7 | 0.34026 | 0.0029358 | 0.0012614 | -0.0010823 | -0.0023437 |  | 6 | 0.011145
5 | YES | 1001.35 | 0.25073 | 0.0024639 | 0.0057344 | -0.0010823 | -0.0068167 |  | 6 | 0.0094066
5 | YES | 802.144 | 0.32442 | 0.003151 | 0.010423 | -0.0010823 | -0.011505 |  | 6 | 0.0095396
5 | YES | 600.811 | 0.58458 | 0.0026526 | 0.0095809 | -0.0010823 | -0.010663 |  | 6 | -0.0031867
5 | YES | 401.365 | 1.5719 | 0.0032972 | 0.011913 | -0.0010823 | -0.012995 |  | 7 | -0.0017133
5 | YES | 301.367 | 2.5459 | 0.0033643 | 0.011917 | -0.0010823 | -0.012999 |  | 7 | 0.0049548
5 | YES | 251.355 | 2.8968 | 0.005109 | -0.0002284 | -0.0010823 | -0.00085389 |  | 7 | 0.017332
5 | NO | 200.723 | 3.7747 | 0.0025842 | 0.064676 | -0.0010823 | -0.065759 |  | 7 | 0.011531
5 | YES | 175.892 | 4.0504 | 0.0050823 | 0.0044317 | -0.0010823 | -0.005514 |  | 7 | 0.011248
5 | YES | 152.017 | 4.3399 | 0.0033632 | 0.0078507 | -0.0010823 | -0.0089329 |  | 7 | 0.024419
5 | NO | 126.208 | 6.4345 | 0.014297 | 0.16955 | -0.0010823 | -0.17063 |  | 7 | 0.019809
5 | YES | 101.54 | 6.8092 | 0.0041404 | 0.0091991 | -0.0010823 | -0.010281 |  | 7 | 0.021403
5 | YES | 75.454 | 6.8701 | 0.0039062 | 0.017091 | -0.0010823 | -0.018174 |  | 7 | 0.01768
5 | YES | 51.2657 | 6.9428 | 0.0043577 | 0.0028338 | -0.0010823 | -0.0039161 |  | 7 | 0.0060126
5 | YES | 26.5587 | 6.7873 | 0.004111 | 0.0013404 | -0.0010823 | -0.0024227 |  | 7 | 0.012664
5 | YES | 11.8022 | 6.7911 | 0.0040584 | 0.0090666 | -0.0010823 | -0.010149 |  | 7 | 0.014863
5 | YES | 5.9859 | 6.7926 | 0.0045795 | 0.0085688 | -0.0010823 | -0.0096511 |  | 7 | 0.013892
6 | NO | 3686.66 | 2.7799 | 0.0062034 | -0.014133 | -0.00063337 | 0.013499 |  | 7 | 0.0022614
6 | YES | 3502.36 | 2.7012 | 0.0055509 | 0.0052238 | -0.00063337 | -0.0058571 |  | 7 | -0.0087926
6 | YES | 3502.36 | 2.7012 | 0.0055509 | 0.00022388 | -0.00063337 | -0.00085724 |  | 7 | 0.012361
6 | YES | 3002.14 | 2.4641 | 0.0051003 | 0.0091205 | -0.00063337 | -0.0097538 |  | 7 | 0.011087
6 | YES | 2500.77 | 1.9221 | 0.0052964 | 0.009083 | -0.00063337 | -0.0097164 |  | 7 | 0.010909
6 | YES | 2001.08 | 1.294 | 0.0042816 | 0.0069668 | -0.00063337 | -0.0076002 |  | 8 | 0.0089959
6 | YES | 1501.36 | 0.52929 | 0.0029959 | 0.021295 | -0.00063337 | -0.021928 |  | 8 | 0.015734
6 | YES | 1251.13 | 0.32311 | 0.0034436 | 0.017112 | -0.00063337 | -0.017745 |  | 8 | 0.018826
6 | YES | 1001.16 | 0.2599 | 0.0031022 | 0.0069004 | -0.00063337 | -0.0075338 |  | 8 | 0.0056888
6 | YES | 802.154 | 0.36185 | 0.0032519 | 0.015855 | -0.00063337 | -0.016488 |  | 8 | 0.0082282
6 | YES | 601.861 | 0.52963 | 0.0020677 | -0.0023693 | -0.00063337 | 0.0017359 |  | 8 | 0.021095
6 | YES | 402.365 | 1.1407 | 0.0033028 | 0.029743 | -0.00063337 | -0.030376 |  | 8 | 0.0095851
6 | YES | 302.555 | 2.0004 | 0.0030079 | 0.007355 | -0.00063337 | -0.0079883 |  | 8 | 0.015909
6 | YES | 251.243 | 2.4644 | 0.0040495 | 0.022361 | -0.00063337 | -0.022994 |  | 8 | 0.017917
6 | NO | 202.859 | 2.9784 | 0.0081902 | -0.0016141 | -0.00063337 | 0.00098072 |  | 8 | 0.0028174
6 | NO | 176.96 | 3.3153 | 0.011363 | -0.029685 | -0.00063337 | 0.029051 |  | 8 | 0.0029502
6 | YES | 151.552 | 3.7322 | 0.0034401 | 0.013211 | -0.00063337 | -0.013845 |  | 8 | -0.021124
6 | NO | 101.254 | 6.4235 | 0.054169 | -0.058473 | -0.00063337 | 0.05784 |  | 8 | -0.0039334
6 | YES | 75.9301 | 6.9652 | 0.0043836 | 0.011145 | -0.00063337 | -0.011778 |  | 8 | -0.055062
6 | YES | 50.9875 | 7.1214 | 0.0046351 | 0.0094066 | -0.00063337 | -0.01004 |  | 8 | -0.053606
6 | NO | 27.6139 | 7.0217 | 0.0082218 | 0.023747 | -0.00063337 | -0.02438 |  | 9 | -0.025431
6 | YES | 11.6753 | 6.9645 | 0.0034999 | 0.0095396 | -0.00063337 | -0.010173 |  | 9 | 0.0039075
6 | YES | 6.1982 | 6.9648 | 0.0038246 | -0.0031867 | -0.00063337 | 0.0025533 |  | 9 | -0.035092
7 | NO | 4024.14 | 3.0274 | 0.0076419 | -0.13661 | -0.00018444 | 0.13642 |  | 9 | -0.015087
7 | YES | 3500.59 | 2.8823 | 0.0058304 | -0.0017133 | -0.00018444 | 0.0015288 |  | 9 | -0.018041
7 | YES | 3500.87 | 2.8819 | 0.0058644 | 0.0049548 | -0.00018444 | -0.0051392 |  | 9 | -0.013938
7 | NO | 3001.22 | 2.524 | 0.0072739 | 0.008054 | -0.00018444 | -0.0082385 |  | 9 | -0.01283
7 | YES | 2501.02 | 1.9293 | 0.0044175 | 0.017332 | -0.00018444 | -0.017516 |  | 9 | -0.055394
7 | YES | 2001.89 | 1.3005 | 0.0033429 | 0.011531 | -0.00018444 | -0.011716 |  | 9 | 0.0047178
7 | YES | 1502.39 | 0.54225 | 0.0034388 | 0.011248 | -0.00018444 | -0.011432 |  | 9 | -0.029589
7 | YES | 1251.06 | 0.32742 | 0.0029486 | 0.024419 | -0.00018444 | -0.024604 |  | 9 | -0.027386
7 | YES | 1001.25 | 0.35381 | 0.0027893 | 0.019809 | -0.00018444 | -0.019994 |  | 9 | -0.037116
7 | YES | 800.214 | 0.4774 | 0.0032506 | 0.021403 | -0.00018444 | -0.021587
7 | YES | 600.366 | 0.63268 | 0.0021 | 0.01768 | -0.00018444 | -0.017865
7 | YES | 401.138 | 1.286 | 0.0028291 | 0.0060126 | -0.00018444 | -0.006197
7 | YES | 298.794 | 1.9327 | 0.0039454 | 0.012664 | -0.00018444 | -0.012848
7 | YES | 251.5 | 2.4019 | 0.0027542 | 0.014863 | -0.00018444 | -0.015047
7 | YES | 201.869 | 3.1099 | 0.0023356 | 0.013892 | -0.00018444 | -0.014077
7 | YES | 176.828 | 3.3583 | 0.0035358 | 0.0022614 | -0.00018444 | -0.0024458
7 | YES | 151.564 | 3.7042 | 0.0029281 | -0.0087926 | -0.00018444 | 0.0086082
7 | NO | 126.859 | 4.6478 | 0.021172 | -0.020237 | -0.00018444 | 0.020053
7 | NO | 101.395 | 6.5194 | 0.0086317 | 0.021394 | -0.00018444 | -0.021579
7 | NO | 76.0086 | 7.0972 | 0.0060386 | 0.011166 | -0.00018444 | -0.011351
7 | NO | 50.9488 | 7.0676 | 0.0083551 | -0.010398 | -0.00018444 | 0.010214
7 | YES | 26.5474 | 7.0424 | 0.0043222 | 0.012361 | -0.00018444 | -0.012545
7 | YES | 11.6412 | 7.0431 | 0.0036704 | 0.011087 | -0.00018444 | -0.011271
7 | YES | 6.4089 | 7.0409 | 0.0042983 | 0.010909 | -0.00018444 | -0.011093
8 | NO | 4317.03 | 3.2331 | 0.0069299 | 0.018121 | 0.00026449 | -0.017856
8 | NO | 4002.27 | 3.1681 | 0.0085231 | 0.016054 | 0.00026449 | -0.015789
8 | NO | 3501.92 | 2.9406 | 0.009179 | 0.009614 | 0.00026449 | -0.0093495
8 | NO | 3001.24 | 2.5513 | 0.0065556 | 0.0073197 | 0.00026449 | -0.0070552
8 | YES | 2500.82 | 1.97 | 0.0042323 | 0.0089959 | 0.00026449 | -0.0087314
8 | YES | 2002.27 | 1.2937 | 0.0044819 | 0.015734 | 0.00026449 | -0.01547
8 | YES | 1501.18 | 0.52983 | 0.0036002 | 0.018826 | 0.00026449 | -0.018561
8 | YES | 1250.9 | 0.50269 | 0.0031102 | 0.0056888 | 0.00026449 | -0.0054243
8 | YES | 1000.85 | 0.51323 | 0.0029864 | 0.0082282 | 0.00026449 | -0.0079637
8 | YES | 801.257 | 0.5591 | 0.0026683 | 0.021095 | 0.00026449 | -0.020831
8 | YES | 601.704 | 0.66158 | 0.0028697 | 0.0095851 | 0.00026449 | -0.0093206
8 | YES | 402.235 | 1.1879 | 0.0026426 | 0.015909 | 0.00026449 | -0.015644
8 | YES | 302.137 | 1.6889 | 0.0018032 | 0.017917 | 0.00026449 | -0.017653
8 | YES | 251.173 | 2.1988 | 0.0034904 | 0.0028174 | 0.00026449 | -0.0025529
8 | NO | 201.208 | 2.8793 | 0.0061589 | -0.014689 | 0.00026449 | 0.014953
8 | NO | 176.22 | 3.7989 | 0.006781 | -0.0010705 | 0.00026449 | 0.001335
8 | NO | 151.922 | 4.8635 | 0.0074055 | -0.036531 | 0.00026449 | 0.036796
8 | YES | 126.302 | 6.9249 | 0.0056086 | 0.0029502 | 0.00026449 | -0.0026857
8 | NO | 101.342 | 7.0887 | 0.0072418 | 0.0026803 | 0.00026449 | -0.0024158
8 | NO | 76.5176 | 7.1932 | 0.0067864 | -0.011834 | 0.00026449 | 0.012099
8 | YES | 51.6772 | 7.1839 | 0.0039677 | -0.021124 | 0.00026449 | 0.021389
8 | YES | 25.7193 | 7.1891 | 0.0053389 | -0.0039334 | 0.00026449 | 0.0041979
8 | YES | 11.8952 | 7.1859 | 0.0027614 | -0.055062 | 0.00026449 | 0.055327
8 | YES | 5.9814 | 7.1824 | 0.0024148 | -0.053606 | 0.00026449 | 0.05387
9 | YES | 1322.82 | 0.51757 | 0.0032946 | -0.025431 | 0.00071342 | 0.026145
9 | YES | 1251.59 | 0.42991 | 0.0031702 | 0.0039075 | 0.00071342 | -0.0031941
9 | YES | 1251.59 | 0.42991 | 0.0031702 | -0.035092 | 0.00071342 | 0.035806
9 | YES | 1001.1 | 0.25091 | 0.0022072 | -0.015087 | 0.00071342 | 0.015801
9 | YES | 800.793 | 0.21196 | 0.0027261 | -0.018041 | 0.00071342 | 0.018755
9 | YES | 601.432 | 0.36606 | 0.0026045 | -0.013938 | 0.00071342 | 0.014651
9 | YES | 401.461 | 1.1842 | 0.0029468 | -0.01283 | 0.00071342 | 0.013543
9 | YES | 300.999 | 2.0466 | 0.0030711 | -0.055394 | 0.00071342 | 0.056108
9 | NO | 200.454 | 3.3028 | 0.0039038 | -0.14624 | 0.00071342 | 0.14695
9 | NO | 176.362 | 3.1839 | 0.014763 | -0.035079 | 0.00071342 | 0.035792
9 | YES | 151.726 | 3.6227 | 0.00489 | 0.0047178 | 0.00071342 | -0.0040044
9 | YES | 125.921 | 3.6784 | 0.0033531 | -0.029589 | 0.00071342 | 0.030303
9 | NO | 101.735 | 4.2717 | 0.0082601 | -0.28225 | 0.00071342 | 0.28296
9 | NO | 76.9966 | 6.5761 | 0.0098289 | -0.02093 | 0.00071342 | 0.021643
9 | YES | 51.6466 | 6.5506 | 0.0057725 | -0.027386 | 0.00071342 | 0.028099
9 | NO | 26.4404 | 6.6889 | 0.0073437 | -0.040123 | 0.00071342 | 0.040837
9 | YES | 11.4606 | 6.7059 | 0.0055583 | -0.037116 | 0.00071342 | 0.03783
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
NO
NO
NO
YES
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
NO
NO
YES
NO
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 2 |  |  |  | COMPARE Version 3.3 | Run at 2012/08/28 17:29:03
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | DOX_BOT | DOX_CTD |  | DOX_CTD - DOX_BOT
Average Size: | 2
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | DOX_BOT | DOX_CTD | DOX_CTD | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 202.433 | 202.433 | 0.16199 | -0.00021362 | 0 | 0.062834 | 0.0021089 | 0.062834
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 177.047 | 177.047 | 0.030704 | -3.0518e-05 | 0 | 0.060187 | 0.0019628 | 0.060187
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 150.761 | 150.761 | 0.03812 | -0.00019836 | 0.048 | 0.058689 | 0.0021111 | 0.010689
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 125.723 | 125.723 | 0.036549 | -0.00037384 | 0.201 | 0.12198 | 0.0018325 | -0.079017
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 101.293 | 101.293 | 0.046951 | -1.5259e-05 | 1.438 | 1.3448 | 0.019998 | -0.093245
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 76.528 | 76.5279 | 0.034142 | -9.1553e-05 | 3.029 | 3.0528 | 0.0071636 | 0.023759
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 51.666 | 51.6656 | 0.049496 | -0.00040436 | 4.542 | 4.5748 | 0.0027081 | 0.032843
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 26.556 | 26.5564 | 0.043342 | 0.00038528 | 7.125 | 7.3571 | 0.0055187 | 0.23208
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 11.597 | 11.5966 | 0.12807 | -0.00041103 | 8.268 | 8.2646 | 0.0077588 | -0.0033608
2012-12-0001.mrgcor1 | 2012-12-0001.samcor1 | 241 | 6.363 | 6.3628 | 0.043734 | -0.00021553 | 9.119 | 9.2006 | 0.018324 | 0.081617
2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 241 | 106.599 | 106.599 | 0.38461 | -0.00048828 | 2.117 | 2.0922 | 0.002509 | -0.024784
2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 241 | 103.39 | 103.39 | 0.62135 | -0.00026703 | 2.162 | 2.1013 | 0.0016321 | -0.060693
2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 434 | 78.14 | 78.0801 | 0.38366 | -0.059853 | 3.823 | 3.7745 | 0.013375 | -0.048472
2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 434 | 78.032 | 78.0801 | 0.38366 | 0.048149 | 3.831 | 3.7745 | 0.013375 | -0.056473
2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 241 | 53.403 | 53.4026 | 0.54531 | -0.00038528 | 5.47 | 5.4546 | 0.0040161 | -0.015381
2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 241 | 27.092 | 27.0918 | 0.37641 | -0.0001564 | 6.245 | 6.2435 | 0.0089363 | -0.0015435
2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 241 | 11.419 | 11.4186 | 0.20598 | -0.00040245 | 6.979 | 7.098 | 0.021015 | 0.11904
2012-12-0003.mrgcor1 | 2012-12-0003.samcor1 | 241 | 5.795 | 5.7948 | 0.45789 | -0.00017452 | 6.872 | 6.8624 | 0.020127 | -0.0096517
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 1317.45 | 1317.45 | 0.34048 | -0.00024414 | 0.498 | 0.48189 | 0.003452 | -0.016112
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 293 | 1251.99 | 1251.86 | 0.18469 | -0.13269 | 0.425 | 0.41935 | 0.003165 | -0.0056451
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 293 | 1251.83 | 1251.86 | 0.18469 | 0.029297 | 0.425 | 0.41935 | 0.003165 | -0.0056451
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 1001.43 | 1001.43 | 0.13674 | -6.1035e-05 | 0.275 | 0.26666 | 0.0021604 | -0.0083403
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 802.22 | 802.22 | 0.9667 | 0.00030518 | 0.209 | 0.20748 | 0.002023 | -0.0015187
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 601.784 | 601.784 | 0.34575 | -0.00036621 | 0.338 | 0.33586 | 0.0018891 | -0.0021369
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 401.671 | 401.671 | 0.41949 | 0.00048828 | 0.876 | 0.86716 | 0.0025566 | -0.0088382
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 302.342 | 302.342 | 0.68267 | -0.00036621 | 1.322 | 1.3099 | 0.0027863 | -0.012108
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 251.391 | 251.391 | 0.4261 | -0.00032043 | 1.872 | 1.8208 | 0.0069165 | -0.051203
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 200.199 | 200.199 | 0.74304 | -0.00021362 | 3.362 | 3.3764 | 0.0033688 | 0.014398
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 176.451 | 176.451 | 0.2305 | 0.00013733 | 3.196 | 3.2252 | 0.0072069 | 0.029178
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 151.95 | 151.95 | 0.45773 | 0.00047302 | 3.117 | 3.1073 | 0.0037625 | -0.0096681
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 125.765 | 125.765 | 0.27153 | -0.00019836 | 3.633 | 3.6318 | 0.0079332 | -0.0011618
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 100.932 | 100.932 | 0.2673 | -8.3923e-05 | 5.741 | 5.8981 | 0.020326 | 0.15714
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 75.445 | 75.4452 | 0.47646 | 0.00024414 | 5.998 | 5.9251 | 0.0063509 | -0.072938
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 52.234 | 52.2339 | 0.21381 | -0.0001297 | 6.14 | 6.162 | 0.0042782 | 0.021963
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 25.525 | 25.525 | 0.25193 | 2.0981e-05 | 6.777 | 6.7452 | 0.01633 | -0.031809
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 11.449 | 11.4485 | 0.16478 | -0.00048161 | 6.737 | 6.7322 | 0.0041773 | -0.0048299
2012-12-0015.mrgcor1 | 2012-12-0015.samcor1 | 241 | 6.6 | 6.6003 | 0.1517 | 0.00031567 | 6.736 | 6.73 | 0.0060944 | -0.006
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 2002.26 | 2002.25 | 0.59629 | -0.00024414 | 1.462 | 1.4263 | 0.0038756 | -0.035718
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 1500.73 | 1500.73 | 1.1349 | 0.00036621 | 0.693 | 0.65835 | 0.0043422 | -0.034647
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 1251.76 | 1251.76 | 0.52737 | 0.00048828 | 0.46 | 0.4217 | 0.0033819 | -0.038303
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 1001.2 | 1001.2 | 0.37699 | -0.00030518 | 0.257 | 0.24682 | 0.0029547 | -0.010178
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 801.132 | 801.132 | 0.29589 | -0.00042725 | 0.234 | 0.2349 | 0.0024937 | 0.00089626
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 601.83 | 601.83 | 0.56238 | 0.00024414 | 0.355 | 0.35298 | 0.0018573 | -0.0020166
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 402.232 | 402.232 | 0.39049 | 6.1035e-05 | 1.151 | 1.1311 | 0.002883 | -0.019905
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 300.937 | 300.937 | 0.66819 | 0.00027466 | 1.816 | 1.783 | 0.0091142 | -0.033046
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 251.956 | 251.956 | 0.20978 | -0.00010681 | 2.538 | 2.5127 | 0.0036929 | -0.025266
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 201.944 | 201.944 | 0.59497 | 0.0002594 | 3.036 | 3.032 | 0.0019243 | -0.0040374
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 176.355 | 176.355 | 0.68093 | -0.00038147 | 3.12 | 3.1059 | 0.0039647 | -0.014137
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 151.66 | 151.66 | 0.52055 | 0.00018311 | 3.489 | 3.4614 | 0.0050682 | -0.027618
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 127.302 | 127.302 | 0.29222 | 0.00032043 | 3.828 | 3.8207 | 0.0030035 | -0.0073528
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 101.397 | 101.397 | 0.42019 | 0.00030518 | 5.17 | 4.9615 | 0.05593 | -0.20851
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 76.156 | 76.1556 | 0.34971 | -0.00043487 | 6.706 | 6.7162 | 0.0056514 | 0.010191
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 435 | 50.624 | 50.5981 | 0.2888 | -0.025879 | 6.902 | 6.9217 | 0.0070124 | 0.019676
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 27.979 | 27.9794 | 0.46524 | 0.0004139 | 6.922 | 6.9493 | 0.013347 | 0.027295
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 241 | 11.139 | 11.1393 | 0.19188 | 0.00027371 | 6.743 | 6.7434 | 0.0078356 | 0.00043154
2012-12-0022.mrgcor1 | 2012-12-0022.samcor1 | 339 | 7.112 | 6.8665 | 0.32019 | -0.24547 | 6.734 | 6.737 | 0.0072416 | 0.002985
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 3276.45 | 3276.45 | 0.22979 | 0.00048828 | 2.322 | 2.31 | 0.0062189 | -0.011992
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 434 | 3001.33 | 3001.23 | 0.32903 | -0.093018 | 2.291 | 2.2956 | 0.005312 | 0.0046222
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 434 | 3001.16 | 3001.23 | 0.32903 | 0.074951 | 2.29 | 2.2956 | 0.005312 | 0.0056222
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 2501.27 | 2501.27 | 0.25263 | 0.00048828 | 1.887 | 1.8921 | 0.0048402 | 0.0050747
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 2000.29 | 2000.29 | 0.24233 | -0.00024414 | 1.353 | 1.3612 | 0.0037454 | 0.0081742
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 1501.41 | 1501.41 | 0.22577 | 0.00012207 | 0.64 | 0.65451 | 0.0033753 | 0.014515
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 1251.7 | 1251.7 | 0.18595 | -0.00048828 | 0.339 | 0.34026 | 0.0029358 | 0.0012614
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 1001.35 | 1001.35 | 0.11672 | 0.00030518 | 0.245 | 0.25073 | 0.0024639 | 0.0057344
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 802.144 | 802.144 | 0.1325 | 0.00042725 | 0.314 | 0.32442 | 0.003151 | 0.010423
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 600.811 | 600.811 | 0.16057 | 0.00030518 | 0.575 | 0.58458 | 0.0026526 | 0.0095809
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 401.365 | 401.365 | 0.09003 | 0.00042725 | 1.56 | 1.5719 | 0.0032972 | 0.011913
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 301.367 | 301.367 | 0.23054 | -0.00030518 | 2.534 | 2.5459 | 0.0033643 | 0.011917
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 251.355 | 251.355 | 0.19872 | -0.00047302 | 2.897 | 2.8968 | 0.005109 | -0.0002284
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 200.723 | 200.723 | 0.1601 | 0.00032043 | 3.71 | 3.7747 | 0.0025842 | 0.064676
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 175.892 | 175.892 | 0.25681 | 0.00019836 | 4.046 | 4.0504 | 0.0050823 | 0.0044317
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 152.017 | 152.017 | 0.10924 | 0.00047302 | 4.332 | 4.3399 | 0.0033632 | 0.0078507
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 126.208 | 126.208 | 0.12997 | -0.00023651 | 6.265 | 6.4345 | 0.014297 | 0.16955
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 101.54 | 101.54 | 0.10846 | 0.00014496 | 6.8 | 6.8092 | 0.0041404 | 0.0091991
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 75.454 | 75.454 | 0.23799 | 3.8147e-05 | 6.853 | 6.8701 | 0.0039062 | 0.017091
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 51.266 | 51.2657 | 0.16407 | -0.00032806 | 6.94 | 6.9428 | 0.0043577 | 0.0028338
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 26.559 | 26.5587 | 0.21085 | -0.00028992 | 6.786 | 6.7873 | 0.004111 | 0.0013404
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 11.802 | 11.8022 | 0.11953 | 0.00022411 | 6.782 | 6.7911 | 0.0040584 | 0.0090666
2012-12-0029.mrgcor1 | 2012-12-0029.samcor1 | 241 | 5.986 | 5.9859 | 0.096546 | -9.5367e-05 | 6.784 | 6.7926 | 0.0045795 | 0.0085688
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 3686.66 | 3686.66 | 0.39633 | 0.00024414 | 2.794 | 2.7799 | 0.0062034 | -0.014133
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 402 | 3502.55 | 3502.36 | 0.83617 | -0.19043 | 2.696 | 2.7012 | 0.0055509 | 0.0052238
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 402 | 3502.23 | 3502.36 | 0.83617 | 0.12744 | 2.701 | 2.7012 | 0.0055509 | 0.00022388
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 216 | 3002.4 | 3002.14 | 0.84462 | -0.25269 | 2.455 | 2.4641 | 0.0051003 | 0.0091205
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 2500.77 | 2500.77 | 0.53751 | -0.00048828 | 1.913 | 1.9221 | 0.0052964 | 0.009083
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 2001.08 | 2001.08 | 0.46306 | 0 | 1.287 | 1.294 | 0.0042816 | 0.0069668
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 1501.35 | 1501.36 | 0.80537 | 0.00048828 | 0.508 | 0.52929 | 0.0029959 | 0.021295
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 1251.14 | 1251.13 | 0.53517 | -0.00012207 | 0.306 | 0.32311 | 0.0034436 | 0.017112
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 1001.16 | 1001.16 | 0.29333 | -0.00018311 | 0.253 | 0.2599 | 0.0031022 | 0.0069004
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 802.154 | 802.154 | 0.59643 | 0.00012207 | 0.346 | 0.36185 | 0.0032519 | 0.015855
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 601.861 | 601.861 | 0.64129 | 6.1035e-05 | 0.532 | 0.52963 | 0.0020677 | -0.0023693
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 402.365 | 402.365 | 0.82407 | 0.00045776 | 1.111 | 1.1407 | 0.0033028 | 0.029743
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 231 | 302.467 | 302.555 | 1.1709 | 0.087891 | 1.993 | 2.0004 | 0.0030079 | 0.007355
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 251.243 | 251.243 | 0.50074 | 0.00015259 | 2.442 | 2.4644 | 0.0040495 | 0.022361
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 202.859 | 202.859 | 0.67247 | 0.00013733 | 2.98 | 2.9784 | 0.0081902 | -0.0016141
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 176.96 | 176.96 | 0.62993 | 0.00030518 | 3.345 | 3.3153 | 0.011363 | -0.029685
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 151.552 | 151.552 | 0.8543 | -0.00021362 | 3.719 | 3.7322 | 0.0034401 | 0.013211
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 101.254 | 101.254 | 0.34108 | -0.00016785 | 6.482 | 6.4235 | 0.054169 | -0.058473
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 75.93 | 75.9301 | 0.61569 | 0.00012207 | 6.954 | 6.9652 | 0.0043836 | 0.011145
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 50.988 | 50.9875 | 0.35327 | -0.00046539 | 7.112 | 7.1214 | 0.0046351 | 0.0094066
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 27.614 | 27.6139 | 0.737 | -5.3406e-05 | 6.998 | 7.0217 | 0.0082218 | 0.023747
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 11.675 | 11.6753 | 0.54129 | 0.00025272 | 6.955 | 6.9645 | 0.0034999 | 0.0095396
2012-12-0046.mrgcor1 | 2012-12-0046.samcor1 | 241 | 6.198 | 6.1982 | 0.44871 | 0.00014925 | 6.968 | 6.9648 | 0.0038246 | -0.0031867
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 4024.14 | 4024.14 | 0.8971 | 0.00048828 | 3.164 | 3.0274 | 0.0076419 | -0.13661
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 265 | 3500.22 | 3500.59 | 0.68124 | 0.36939 | 2.884 | 2.8823 | 0.0058304 | -0.0017133
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 311 | 3501.06 | 3500.87 | 0.92608 | -0.18945 | 2.877 | 2.8819 | 0.0058644 | 0.0049548
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 3001.22 | 3001.22 | 0.89778 | -0.00024414 | 2.516 | 2.524 | 0.0072739 | 0.008054
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 2501.02 | 2501.02 | 0.76909 | 0 | 1.912 | 1.9293 | 0.0044175 | 0.017332
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 2001.89 | 2001.89 | 0.61364 | -0.00024414 | 1.289 | 1.3005 | 0.0033429 | 0.011531
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 222 | 1502.2 | 1502.39 | 1.1257 | 0.19104 | 0.531 | 0.54225 | 0.0034388 | 0.011248
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 1251.06 | 1251.06 | 0.4692 | 0.00024414 | 0.303 | 0.32742 | 0.0029486 | 0.024419
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 1001.25 | 1001.25 | 0.47484 | 0.00024414 | 0.334 | 0.35381 | 0.0027893 | 0.019809
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 800.214 | 800.214 | 0.75188 | -0.00018311 | 0.456 | 0.4774 | 0.0032506 | 0.021403
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 600.366 | 600.366 | 0.53044 | 0.00018311 | 0.615 | 0.63268 | 0.0021 | 0.01768
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 240 | 401.146 | 401.138 | 1.0973 | -0.0081482 | 1.28 | 1.286 | 0.0028291 | 0.0060126
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 298.794 | 298.794 | 0.63539 | -0.00042725 | 1.92 | 1.9327 | 0.0039454 | 0.012664
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 251.5 | 251.5 | 0.1798 | 0.00032043 | 2.387 | 2.4019 | 0.0027542 | 0.014863
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 201.869 | 201.869 | 0.37361 | 3.0518e-05 | 3.096 | 3.1099 | 0.0023356 | 0.013892
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 176.828 | 176.828 | 0.24967 | -0.00039673 | 3.356 | 3.3583 | 0.0035358 | 0.0022614
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 151.564 | 151.564 | 0.19157 | 0.0002594 | 3.713 | 3.7042 | 0.0029281 | -0.0087926
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 126.859 | 126.859 | 1.0618 | 9.9182e-05 | 4.668 | 4.6478 | 0.021172 | -0.020237
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 101.395 | 101.395 | 0.5585 | -0.00034332 | 6.498 | 6.5194 | 0.0086317 | 0.021394
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 76.009 | 76.0086 | 0.59238 | -0.00040436 | 7.086 | 7.0972 | 0.0060386 | 0.011166
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 50.949 | 50.9488 | 0.29986 | -0.00021744 | 7.078 | 7.0676 | 0.0083551 | -0.010398
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 26.547 | 26.5474 | 0.5503 | 0.00041771 | 7.03 | 7.0424 | 0.0043222 | 0.012361
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 11.641 | 11.6412 | 0.36756 | 0.00020313 | 7.032 | 7.0431 | 0.0036704 | 0.011087
2012-12-0061.mrgcor1 | 2012-12-0061.samcor1 | 241 | 6.409 | 6.4089 | 0.47099 | -8.297e-05 | 7.03 | 7.0409 | 0.0042983 | 0.010909
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 4317.03 | 4317.03 | 0.35008 | 0.00048828 | 3.215 | 3.2331 | 0.0069299 | 0.018121
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 4002.27 | 4002.27 | 0.80724 | -0.00024414 | 3.152 | 3.1681 | 0.0085231 | 0.016054
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 3501.92 | 3501.92 | 0.98652 | 0.00048828 | 2.931 | 2.9406 | 0.009179 | 0.009614
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 3001.24 | 3001.24 | 0.51147 | -0.00024414 | 2.544 | 2.5513 | 0.0065556 | 0.0073197
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 2500.82 | 2500.82 | 0.21133 | 0.00024414 | 1.961 | 1.97 | 0.0042323 | 0.0089959
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 2002.27 | 2002.27 | 0.58703 | 0 | 1.278 | 1.2937 | 0.0044819 | 0.015734
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 1501.18 | 1501.18 | 0.69468 | 0.00024414 | 0.511 | 0.52983 | 0.0036002 | 0.018826
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 1250.9 | 1250.9 | 0.357 | 0.00024414 | 0.497 | 0.50269 | 0.0031102 | 0.0056888
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 1000.85 | 1000.85 | 0.33986 | 0.00042725 | 0.505 | 0.51323 | 0.0029864 | 0.0082282
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 801.257 | 801.257 | 0.6888 | -0.00030518 | 0.538 | 0.5591 | 0.0026683 | 0.021095
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 601.704 | 601.704 | 0.28041 | -0.00030518 | 0.652 | 0.66158 | 0.0028697 | 0.0095851
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 402.235 | 402.235 | 0.38881 | -0.00033569 | 1.172 | 1.1879 | 0.0026426 | 0.015909
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 302.137 | 302.137 | 0.34302 | -0.00015259 | 1.671 | 1.6889 | 0.0018032 | 0.017917
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 251.173 | 251.173 | 0.32835 | 0.00028992 | 2.196 | 2.1988 | 0.0034904 | 0.0028174
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 201.208 | 201.208 | 0.30964 | 0.00015259 | 2.894 | 2.8793 | 0.0061589 | -0.014689
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 176.22 | 176.22 | 0.36456 | -0.00028992 | 3.8 | 3.7989 | 0.006781 | -0.0010705
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 151.922 | 151.922 | 0.27552 | 0.00012207 | 4.9 | 4.8635 | 0.0074055 | -0.036531
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 126.302 | 126.302 | 0.25671 | -0.00035858 | 6.922 | 6.9249 | 0.0056086 | 0.0029502
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 101.342 | 101.342 | 0.10475 | -0.00049591 | 7.086 | 7.0887 | 0.0072418 | 0.0026803
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 76.518 | 76.5176 | 0.40493 | -0.00035095 | 7.205 | 7.1932 | 0.0067864 | -0.011834
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 51.677 | 51.6772 | 0.32778 | 0.00019074 | 7.205 | 7.1839 | 0.0039677 | -0.021124
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 25.719 | 25.7193 | 0.17255 | 0.00032425 | 7.193 | 7.1891 | 0.0053389 | -0.0039334
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 11.895 | 11.8952 | 0.13005 | 0.00017357 | 7.241 | 7.1859 | 0.0027614 | -0.055062
2012-12-0078.mrgcor1 | 2012-12-0078.samcor1 | 241 | 5.981 | 5.9814 | 0.14565 | 0.00039005 | 7.236 | 7.1824 | 0.0024148 | -0.053606
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 1322.82 | 1322.82 | 0.84199 | -0.00024414 | 0.543 | 0.51757 | 0.0032946 | -0.025431
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 292 | 1251.46 | 1251.59 | 0.20697 | 0.12964 | 0.426 | 0.42991 | 0.0031702 | 0.0039075
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 292 | 1251.61 | 1251.59 | 0.20697 | -0.027344 | 0.465 | 0.42991 | 0.0031702 | -0.035092
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 1001.1 | 1001.1 | 0.83661 | 0.00042725 | 0.266 | 0.25091 | 0.0022072 | -0.015087
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 800.793 | 800.793 | 0.19936 | 0.00042725 | 0.23 | 0.21196 | 0.0027261 | -0.018041
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 601.432 | 601.432 | 0.52369 | 0.00012207 | 0.38 | 0.36606 | 0.0026045 | -0.013938
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 401.461 | 401.461 | 0.60514 | -0.00048828 | 1.197 | 1.1842 | 0.0029468 | -0.01283
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 300.999 | 300.999 | 0.64992 | 0.00036621 | 2.102 | 2.0466 | 0.0030711 | -0.055394
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 200.454 | 200.454 | 0.4328 | -3.0518e-05 | 3.449 | 3.3028 | 0.0039038 | -0.14624
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 176.362 | 176.362 | 0.59087 | -0.00033569 | 3.219 | 3.1839 | 0.014763 | -0.035079
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 151.726 | 151.726 | 0.34654 | 9.1553e-05 | 3.618 | 3.6227 | 0.00489 | 0.0047178
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 125.921 | 125.921 | 0.42118 | 0.00020599 | 3.708 | 3.6784 | 0.0033531 | -0.029589
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 101.735 | 101.735 | 0.57168 | 9.1553e-05 | 4.554 | 4.2717 | 0.0082601 | -0.28225
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 76.997 | 76.9966 | 0.2592 | -0.00037384 | 6.597 | 6.5761 | 0.0098289 | -0.02093
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 51.647 | 51.6466 | 0.54905 | -0.0004158 | 6.578 | 6.5506 | 0.0057725 | -0.027386
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 349 | 26.754 | 26.4404 | 0.46778 | -0.31362 | 6.729 | 6.6889 | 0.0073437 | -0.040123
2012-12-0107.mrgcor1 | 2012-12-0107.samcor1 | 241 | 11.461 | 11.4606 | 0.23323 | -0.00043583 | 6.743 | 6.7059 | 0.0055583 | -0.037116
Create Fit Sheet