## Notes
Analysis: | Melissa Hennekes, 10-July-12 to 11-July-12 using Water Properties' Guildline 8400B Salinometer Serial Number 68572.
Chief Scientist: | Marie Robert
2012-08-15 |  | GG
Flags/Comments changes
cast 19, sample 181 - outlier in COMPARE in well-mixed surface waters; already flagged because of missing liner; added note, left flag unchanged.
cast 53, sample 435 - Surface sample. The value looks too high, nothing like it until almost 100m down. Flag 3.
cast 107, sample 818 - Looks like a mis-sample; probably from 600m bottle; DO sample looks ok so bottle likely closed correctly
cast 197, sample 825 - outlier in COMPARE but local gradient very high, so no flag or comment added.
2012-11-06 |  | KS
Salinometer Calibration Information
Sample ID | Event | Station | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio
calibration#1 |  |  | K15 | 2012-07-10 09:21 | 24C | 0.99978 | 2.2397e-06 | 1.0133e-06 | 0.99978
Included sample #605 in duplicate analysis
Separated Loop samples into its own QF worksheet
Deep Water Reference Data (2012-01 Feb Line P Sample #: 320/321 Station: P24 Cast: 36 Depth: 2000 m)
Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Julian Day collected | Julian Day analyzed | Days stored
DWR2012-01 | 2012-07-10 09:26 | 24C | 0.98954 | 3.6692e-06 | 9.5367e-07 | 0.98954 | 34.5885 | 0.00014345 | 45 | 192 | 147
DWR2012-01 | 2012-07-10 10:50 | 24C | 0.98954 | 2.0838e-06 | 9.5367e-07 | 0.98954 | 34.5886 | 8.1621e-05 | 45 | 192 | 147
DWR2012-01 | 2012-07-10 11:54 | 24C | 0.98939 | 3.4669e-06 | 1.0133e-06 | 0.98939 | 34.5827 | 0.00013345 | 45 | 192 | 147
DWR2012-12 | 2012-07-10 11:57 | 24C | 0.98961 | 2.4909e-06 | 9.5367e-07 | 0.98961 | 34.5916 | 9.869e-05 | 45 | 192 | 147
DWS2012-12 | 2012-07-10 14:21 | 24C | 0.98955 | 2.4918e-06 | 9.5367e-07 | 0.98955 | 34.5892 | 9.8074e-05 | 45 | 192 | 147
DWS2012-01 | 2012-07-10 15:14 | 24C | 0.98949 | 5.2332e-06 | 9.5367e-07 | 0.98949 | 34.5866 | 0.00020838 | 45 | 192 | 147
DWR2012-01 | 2012-07-11 09:11 | 24C | 0.98941 | 3.2141e-06 | 9.5367e-07 | 0.98941 | 34.5836 | 0.00012667 | 45 | 193 | 148
DWR2012-12 | 2012-07-11 09:16 | 24C | 0.98954 | 2.4912e-06 | 9.5367e-07 | 0.98954 | 34.5886 | 9.869e-05 | 45 | 193 | 148
DWR2012-01 | 2012-07-11 09:58 | 24C | 0.98942 | 2.4909e-06 | 9.5367e-07 | 0.98942 | 34.584 | 9.7904e-05 | 45 | 193 | 148
DWS2012-12 | 2012-07-11 13:16 | 24C | 0.98956 | 2.2401e-06 | 9.5367e-07 | 0.98956 | 34.5894 | 8.7082e-05 | 45 | 193 | 148
DWR2012-01 | 2012-07-11 14:48 | 24C | 0.98999 | 3.2144e-06 | 9.5367e-07 | 0.98999 | 34.6063 | 0.0001263 | 45 | 193 | 148
2015-03-30 |  | KS
Amended header in the QF worksheet to show that silicone inserts were used in place of the typical nylon inserts supplied by Guildline Instruments.

## QF 2012-12 SAL FINAL
Header: | Salinity samples were collected in 200 mL type ll glass bottles with reusable silicone inserts supplied by Xiamen Cuten Accessories Industry, Ltd. And
screw caps supplied by Ocean Scientific International Limited. They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer
standardized with IAPSO standard seawater ~2 months after collection. For details including a duplicate analysis, see document QF2012-12SAL.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Salinity:Bottle | Flag:Salinity Bottle | Comments
2012-12 | 2 | P1 | 13 | 30.7415
2012-12 | 3 | P2 | 21 | 31.7705
2012-12 | 6 | P3 | 29 | 31.6641
2012-12 | 15 | P4 | 118 | 34.4636 | 6
2012-12 | 15 | P4 | 120 | 34.4054 | 6
2012-12 | 15 | P4 | 121 | 34.3127
2012-12 | 15 | P4 | 135 | 31.7809
2012-12 | 16 | P4 | 136 | 34.4766
2012-12 | 16 | P4 | 139 | 34.3924
2012-12 | 16 | P4 | 142 | 34.1418
2012-12 | 16 | P4 | 145 | 31.7648
2012-12 | 19 | P5 | 181 | 31.8631 | 4 | SAL: no liner in bottle; outlier in comparison to CTD in well-mixed surface waters
2012-12 | 20 | P6 | 183 | 31.7385
2012-12 | 21 | P7 | 185 | 31.8061
2012-12 | 22 | P8 | 186 | 34.5973 | 6
2012-12 | 22 | P8 | 188 | 34.4711
2012-12 | 22 | P8 | 189 | 34.3878 | 6
2012-12 | 22 | P8 | 190 | 34.2988
2012-12 | 22 | P8 | 204 | 31.9919
2012-12 | 24 | P9 | 209 | 31.751
2012-12 | 25 | P10 | 211 | 32.6493
2012-12 | 26 | P11 | 213 | 32.6485
2012-12 | 29 | P12 | 233 | 34.6495 | 6
2012-12 | 29 | P12 | 236 | 34.5904
2012-12 | 29 | P12 | 239 | 34.3703 | 6
2012-12 | 29 | P12 | 240 | 34.2665
2012-12 | 29 | P12 | 254 | 32.6579
2012-12 | 34 | P12 | 279 | 34.5898
2012-12 | 34 | P12 | 282 | 34.367
2012-12 | 34 | P12 | 285 | 34.0605
2012-12 | 34 | P12 | 288 | 32.6613
2012-12 | 36 | P12 | 291 | 34.2706
2012-12 | 36 | P12 | 292 | 34.1423
2012-12 | 36 | P12 | 293 | 33.9788
2012-12 | 36 | P12 | 294 | 33.8777
2012-12 | 38 | P13 | 327 | 32.6591
2012-12 | 39 | P14 | 329 | 32.6512
2012-12 | 40 | P15 | 331 | 32.7039
2012-12 | 46 | P16 | 371 | 34.6689
2012-12 | 46 | P16 | 373 | 34.6584 | 6
2012-12 | 46 | P16 | 375 | 34.5907
2012-12 | 46 | P16 | 378 | 34.3845 | 6
2012-12 | 46 | P16 | 379 | 34.297
2012-12 | 46 | P16 | 393 | 32.7175
2012-12 | 51 | P16 | 418 | 34.6592
2012-12 | 51 | P16 | 419 | 34.5945
2012-12 | 51 | P16 | 420 | 34.5512
2012-12 | 51 | P16 | 421 | 34.4939
2012-12 | 51 | P16 | 422 | 34.4418
2012-12 | 51 | P16 | 423 | 34.379
2012-12 | 52 | P16 | 430 | 32.7456
2012-12 | 53 | P17 | 435 | 32.7312 | 3 | SAL: Surface sample. The value looks too high. Nothing like it until almost 100m level
2012-12 | 54 | P18 | 437 | 32.6725
2012-12 | 55 | P19 | 439 | 32.6704
2012-12 | 61 | P20 | 485 | 34.6725
2012-12 | 61 | P20 | 487 | 34.6563 | 6
2012-12 | 61 | P20 | 489 | 34.5906
2012-12 | 61 | P20 | 492 | 34.3649 | 6
2012-12 | 61 | P20 | 493 | 34.2905
2012-12 | 61 | P20 | 507 | 32.6681
2012-12 | 65 | P20 | 526 | 34.6549
2012-12 | 65 | P20 | 527 | 34.5892
2012-12 | 65 | P20 | 528 | 34.5097
2012-12 | 65 | P20 | 529 | 34.4364
2012-12 | 65 | P20 | 530 | 34.375
2012-12 | 65 | P20 | 531 | 34.2869
2012-12 | 65 | P20 | 532 | 34.1787
2012-12 | 67 | P21 | 542 | 32.6693
2012-12 | 70 | P22 | 544 | 32.6526
2012-12 | 71 | P23 | 545 | 34.5882
2012-12 | 71 | P23 | 546 | 34.588
2012-12 | 71 | P23 | 547 | 34.5887
2012-12 | 71 | P23 | 548 | 34.5895
2012-12 | 71 | P23 | 549 | 34.5891
2012-12 | 71 | P23 | 550 | 34.5869
2012-12 | 71 | P23 | 551 | 34.586
2012-12 | 71 | P23 | 552 | 34.5889
2012-12 | 71 | P23 | 553 | 34.5884
2012-12 | 71 | P23 | 554 | 34.588
2012-12 | 71 | P23 | 555 | 34.5883
2012-12 | 71 | P23 | 556 | 34.5875
2012-12 | 71 | P23 | 557 | 34.5881
2012-12 | 71 | P23 | 558 | 34.5849
2012-12 | 71 | P23 | 559 | 34.5881
2012-12 | 71 | P23 | 560 | 34.5871
2012-12 | 71 | P23 | 561 | 34.5883
2012-12 | 71 | P23 | 562 | 34.585
2012-12 | 71 | P23 | 563 | 34.588
2012-12 | 71 | P23 | 564 | 34.5847
2012-12 | 71 | P23 | 565 | 34.5878
2012-12 | 71 | P23 | 566 | 34.5877
2012-12 | 71 | P23 | 567 | 32.6748
2012-12 | 71 | P23 | 568 | 32.6407
2012-12 | 72 | P24 | 570 | 32.6661
2012-12 | 73 | P25 | 575 | 32.658
2012-12 | 74 | P35 | 577 | 32.6924
2012-12 | 78 | P26 | 604 | 34.6879
2012-12 | 78 | P26 | 605 | 34.6835
2012-12 | 78 | P26 | 605 | 34.6829 | 6
2012-12 | 78 | P26 | 606 | 34.6736
2012-12 | 78 | P26 | 607 | 34.6598
2012-12 | 78 | P26 | 608 | 34.6326
2012-12 | 78 | P26 | 609 | 34.5888
2012-12 | 78 | P26 | 610 | 34.5072
2012-12 | 78 | P26 | 611 | 34.4487
2012-12 | 78 | P26 | 612 | 34.3829
2012-12 | 78 | P26 | 613 | 34.3049
2012-12 | 78 | P26 | 614 | 34.206
2012-12 | 78 | P26 | 615 | 34.0363
2012-12 | 78 | P26 | 616 | 33.8997
2012-12 | 78 | P26 | 617 | 33.8266
2012-12 | 78 | P26 | 618 | 33.7459
2012-12 | 78 | P26 | 619 | 33.6724
2012-12 | 78 | P26 | 620 | 33.4393
2012-12 | 78 | P26 | 621 | 32.8525
2012-12 | 78 | P26 | 622 | 32.75
2012-12 | 78 | P26 | 623 | 32.7226
2012-12 | 78 | P26 | 624 | 32.7179
2012-12 | 78 | P26 | 625 | 32.713
2012-12 | 78 | P26 | 626 | 32.7149
2012-12 | 78 | P26 | 627 | 32.7121
2012-12 | 81 | P26 | 628 | 34.5823
2012-12 | 81 | P26 | 631 | 34.3835
2012-12 | 81 | P26 | 634 | 34.1081
2012-12 | 81 | P26 | 637 | 32.7125
2012-12 | 90 | P26 | 705 | 34.6865
2012-12 | 90 | P26 | 706 | 34.6726
2012-12 | 90 | P26 | 707 | 34.6778
2012-12 | 90 | P26 | 708 | 34.6334 | 3 | SAL: cap liner crimped, salt crystals on cap
2012-12 | 90 | P26 | 709 | 34.5831
2012-12 | 90 | P26 | 710 | 34.5723
2012-12 | 96 | P26 | 721 | 34.4314
2012-12 | 96 | P26 | 722 | 34.3742
2012-12 | 96 | P26 | 723 | 34.2936
2012-12 | 96 | P26 | 724 | 34.1847
2012-12 | 96 | P26 | 725 | 34.1128
2012-12 | 96 | P26 | 726 | 33.9404
2012-12 | 107 | P4 | 812 | 34.4884
2012-12 | 107 | P4 | 813 | 34.4681 | 6
2012-12 | 107 | P4 | 814 | 34.4674
2012-12 | 107 | P4 | 815 | 34.3928 | 6
2012-12 | 107 | P4 | 816 | 34.3097
2012-12 | 107 | P4 | 817 | 34.1843
2012-12 | 107 | P4 | 818 | -99 | 5 | SAL: Looks like a mis-sample; probably from 600m bottle; DO sample looks ok so bottle likely closed correctly
2012-12 | 107 | P4 | 819 | 33.9305
2012-12 | 107 | P4 | 820 | 33.8967
2012-12 | 107 | P4 | 821 | 33.866
2012-12 | 107 | P4 | 822 | 33.8553
2012-12 | 107 | P4 | 823 | 33.7915
2012-12 | 107 | P4 | 824 | 33.6765
2012-12 | 107 | P4 | 825 | 33.2259
2012-12 | 107 | P4 | 826 | 32.679
2012-12 | 107 | P4 | 827 | 32.6071
2012-12 | 107 | P4 | 828 | 31.8461
2012-12 | 107 | P4 | 829 | 31.8275
2012-12 | 107 | P4 | 830 | 31.827
2012-12 | 108 | P4 | 834 | 34.4595
2012-12 | 108 | P4 | 835 | 34.3908
2012-12 | 108 | P4 | 836 | 34.3092
2012-12 | 108 | P4 | 837 | 34.189

## QF 2012-12 SAL LOOP FINAL
Header: | Salinity samples were collected in 200 mL type II glass bottles with disposable plastic inserts and screw cap
supplied by Ocean Scientific International Limited. They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer
standardized with IAPSO standard seawater ~2 months after collection. For details including a duplicate analysis, see document QFcruise-##SAL.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Salinity:Bottle | Flag:Salinity Bottle | Comments
2012-12 |  |  | loop JF3 | 30.2828
2012-12 |  |  | loop6 | 32.3335
2012-12 |  |  | loop7 | 31.8302
2012-12 |  |  | loopJF1 | 30.3852
2012-12 |  |  | loopJF1 | 30.3987
2012-12 |  |  | loopJF4 | 31.0426
2012-12 |  |  | loopNofP16 | 32.6867
2012-12 |  |  | loopP11 | 32.6456
2012-12 |  |  | loopP14 | 32.6467
2012-12 |  |  | loopP20 | 32.6592
2012-12 |  |  | loopP21 | 32.6704
2012-12 |  |  | loopP22 | 32.6481
2012-12 |  |  | loopP24-23 | 32.6594

## Duplicates
Precision Statement for Replicate Samples Drawn From a Single Niskin Bottle:
Mean Difference for salinity for the range 34.3639 to 34.6835 was 0.0010, k = 26 with 0 outlier(s) removed
Precision Analysis and the Determination of Outliers
Precision was determined by analyzing replicate samples drawn from one Niskin
Outliers are discarded on the basis of Chauvenet's criteria. The statistic is calculated using the following formula:
The maximum deviation, dmax, is compared with the individual residuals from the original concentrations (see Table 1, columns L, M and N).
If a replicate's residual is greater than dmax, this reading can be rejected based upon Chauvenet's criterion
Table 1. Replicate Analysis for Samples Drawn From a Single Niskin Bottle | [14 empty] | Table 2. Critical values for Chauvenet statistic
Event | Station | Sample_ Number | SAL1 | SAL2 | SALAvg | Diff | Residual1 (Xi-Xm) | Residual2 (Xi-Xm) | Residual3 (Xi-Xm) | Σ(Xi-Xm)2 | n | Reject? | SAL1 Comment | SAL2 Comment | n | Chauvenet
15 | P4 | 118 | 34.4634 | 34.4638 | 34.4636 | 0.00037445 | -0.00018722 | 0.00018722 |  | 7.0106e-08 | 2 |  |  |  | 3 | 1.37
15 | P4 | 120 | 34.4068 | 34.404 | 34.4054 | 0.0027889 | 0.0013944 | -0.0013944 |  | 3.889e-06 | 2 |  |  |  | 4 | 1.53
22 | P8 | 186 | 34.5975 | 34.5972 | 34.5973 | 0.00037468 | 0.00018734 | -0.00018734 |  | 7.0192e-08 | 2 |  |  |  | 5 | 1.64
22 | P8 | 189 | 34.3891 | 34.3865 | 34.3878 | 0.0025875 | 0.0012937 | -0.0012937 |  | 3.3475e-06 | 2 |  |  |  | 6 | 1.73
29 | P12 | 233 | 34.6491 | 34.6498 | 34.6495 | 0.00061134 | -0.00030567 | 0.00030567 |  | 1.8687e-07 | 2 |  |  |  | 7 | 1.8
29 | P12 | 239 | 34.3705 | 34.37 | 34.3703 | 0.0004702 | 0.0002351 | -0.0002351 |  | 1.1054e-07 | 2 |  |  |  | 8 | 1.87
46 | P16 | 373 | 34.6583 | 34.6585 | 34.6584 | 0.00011712 | -5.856e-05 | 5.856e-05 |  | 6.8585e-09 | 2 |  |  |  | 9 | 1.91
46 | P16 | 378 | 34.385 | 34.3841 | 34.3845 | 0.00098023 | 0.00049011 | -0.00049011 |  | 4.8043e-07 | 2 |  |  |  | 10 | 1.97
61 | P20 | 487 | 34.6562 | 34.6565 | 34.6563 | 0.0002155 | -0.00010775 | 0.00010775 |  | 2.322e-08 | 2 |  |  |  | 20 | 2.24
61 | P20 | 492 | 34.3639 | 34.3658 | 34.3649 | 0.0018807 | -0.00094037 | 0.00094037 |  | 1.7686e-06 | 2 |  |  |  | 30 | 2.39
78 | P26 | 605 | 34.6835 | 34.6824 | 34.6829 | 0.0011198 | 0.0005599 | -0.0005599 |  | 6.2697e-07 | 2 |  |  |  | 50 | 2.57
107 | P4 | 813 | 34.4686 | 34.4676 | 34.4681 | 0.0010602 | 0.00053009 | -0.00053009 |  | 5.6199e-07 | 2 |  |  |  | 80 | 2.74
107 | P4 | 815 | 34.3927 | 34.3928 | 34.3928 | 1.8716e-05 | -9.3582e-06 | 9.3582e-06 |  | 1.7515e-10 | 2 |  |  |  | 100 | 2.81
[15 empty] | 150 | 2.93
[15 empty] | 200 | 3.02
[9 empty] | Sum | 1.1142e-05
[9 empty] | Count |  | 26
[9 empty] | σ | 0.0006676
[9 empty] | dmax/σ for n (Table 2) | 2.39
[9 empty] | Thus dmax = 2.57 x σ | 0.0015956
[9 empty] | Mean Diff | 0.00096918
[9 empty] | MIN | 34.3639
[9 empty] | MAX | 34.6835
0 Outlier(s) Removed
[9 empty] | Mean Diff = | 0.0010187

## Flags
Institute of Ocean Sciences - Quality Flags for Water Sampling | Last Updated: October 2011
Double click document to use.

## Salinity Data Processing
Sample ID | Event | Station | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | New Calculated Salinity | Flag | Comments
2012-12#76 | 2 | P1 | 13 | 2012-07-10 14:15 | 24C | 0.89073 | 4.199e-06 | 9.5367e-07 | 0.89073 | 30.7415 | 0.00016215 | #VALUE!
2012-12#83 | 3 | P2 | 21 | 2012-07-10 14:55 | 24C | 0.91733 | 2.2391e-06 | 9.5367e-07 | 0.91733 | 31.7705 | 8.5313e-05 | #VALUE!
2012-12#110 | 6 | P3 | 29 | 2012-07-11 10:25 | 24C | 0.91459 | 5.5696e-06 | 9.5367e-07 | 0.91459 | 31.6641 | 0.00021508 | #VALUE!
2012-12#53 | 15 | P4 | 118 | 2012-07-10 13:12 | 24C | 0.98635 | 2.2391e-06 | 8.9407e-07 | 0.98635 | 34.4634 | 8.7876e-05 | #VALUE!
2012-12#116 | 15 | P4 | 118 | 2012-07-11 10:40 | 24C | 0.98636 | 2.7425e-06 | 9.5367e-07 | 0.98636 | 34.4638 | 0.00010865 | #VALUE!
2012-12#59 | 15 | P4 | 120 | 2012-07-10 13:29 | 24C | 0.98491 | 8.5621e-06 | 1.0729e-06 | 0.98491 | 34.4068 | 0.00033444 | #VALUE!
2012-12#71 | 15 | P4 | 120 | 2012-07-10 14:01 | 24C | 0.98484 | 1.0053e-05 | 9.5367e-07 | 0.98484 | 34.404 | 0.00039415 | #VALUE!
2012-12#65 | 15 | P4 | 121 | 2012-07-10 13:45 | 24C | 0.98251 | 1.7705e-06 | 9.5367e-07 | 0.98251 | 34.3127 | 6.9481e-05 | #VALUE!
2012-12#77 | 15 | P4 | 135 | 2012-07-10 14:17 | 24C | 0.9176 | 2.4776e-06 | 9.5367e-07 | 0.9176 | 31.7809 | 9.752e-05 | #VALUE!
2012-12#84 | 16 | P4 | 136 | 2012-07-10 14:58 | 24C | 0.98669 | 2.7425e-06 | 8.3447e-07 | 0.98669 | 34.4766 | 0.00010656 | #VALUE!
2012-12#111 | 16 | P4 | 139 | 2012-07-11 10:28 | 24C | 0.98454 | 4.6249e-06 | 9.5367e-07 | 0.98454 | 34.3924 | 0.00018166 | #VALUE!
2012-12#117 | 16 | P4 | 142 | 2012-07-11 10:43 | 24C | 0.97815 | 2.8766e-06 | 9.5367e-07 | 0.97815 | 34.1418 | 0.000113 | #VALUE!
2012-12#54 | 16 | P4 | 145 | 2012-07-10 13:15 | 24C | 0.91719 | 4.9345e-06 | 8.9407e-07 | 0.91719 | 31.7648 | 0.00019339 | #VALUE!
2012-12#60 | 19 | P5 | 181 | 2012-07-10 13:32 | 24C | 0.91972 | 6.856e-06 | 9.5367e-07 | 0.91972 | 31.8631 | 0.00026274 | #VALUE! | 4 | SAL: no liner in bottle; outlier in comparison to CTD in well-mixed surface waters
2012-12#66 | 20 | P6 | 183 | 2012-07-10 13:48 | 24C | 0.91651 | 5.2339e-06 | 8.9407e-07 | 0.91651 | 31.7385 | 0.00020275 | #VALUE!
2012-12#72 | 21 | P7 | 185 | 2012-07-10 14:04 | 24C | 0.91825 | 2.4912e-06 | 9.5367e-07 | 0.91825 | 31.8061 | 9.6511e-05 | #VALUE!
2012-12#79 | 22 | P8 | 186 | 2012-07-10 14:44 | 24C | 0.98976 | 2.4747e-06 | 9.5367e-07 | 0.98977 | 34.5975 | 9.4885e-05 | #VALUE!
2012-12#93 | 22 | P8 | 186 | 2012-07-11 09:18 | 24C | 0.98975 | 2.7264e-06 | 9.5367e-07 | 0.98976 | 34.5972 | 0.00010656 | #VALUE!
2012-12#112 | 22 | P8 | 188 | 2012-07-11 10:30 | 24C | 0.98654 | 2.4775e-06 | 9.5367e-07 | 0.98655 | 34.4711 | 9.68e-05 | #VALUE!
2012-12#118 | 22 | P8 | 189 | 2012-07-11 10:46 | 24C | 0.98446 | 2.7425e-06 | 9.5367e-07 | 0.98446 | 34.3891 | 0.00010761 | #VALUE!
2012-12#133 | 22 | P8 | 189 | 2012-07-11 11:25 | 24C | 0.98439 | 2.7425e-06 | 9.5367e-07 | 0.98439 | 34.3865 | 0.00010764 | #VALUE!
2012-12#139 | 22 | P8 | 190 | 2012-07-11 11:40 | 24C | 0.98215 | 4.7146e-06 | 9.5367e-07 | 0.98216 | 34.2988 | 0.00018787 | #VALUE!
2012-12#127 | 22 | P8 | 204 | 2012-07-11 11:09 | 24C | 0.92304 | 2.2262e-06 | 9.5367e-07 | 0.92304 | 31.9919 | 8.6589e-05 | #VALUE!
2012-12#121 | 24 | P9 | 209 | 2012-07-11 10:54 | 24C | 0.91683 | 2.7425e-06 | 9.5367e-07 | 0.91683 | 31.751 | 0.0001071 | #VALUE!
2012-12#115 | 25 | P10 | 211 | 2012-07-11 10:37 | 24C | 0.93995 | 3.1414e-06 | 9.5367e-07 | 0.93995 | 32.6493 | 0.00012313 | #VALUE!
2012-12#102 | 26 | P11 | 213 | 2012-07-11 09:42 | 24C | 0.93993 | 2.7425e-06 | 9.5367e-07 | 0.93993 | 32.6485 | 0.00010762 | #VALUE!
2012-12#132 | 29 | P12 | 233 | 2012-07-11 11:22 | 24C | 0.99108 | 2.7425e-06 | 9.5367e-07 | 0.99108 | 34.6491 | 0.00010657 | #VALUE!
2012-12#138 | 29 | P12 | 233 | 2012-07-11 11:37 | 24C | 0.99109 | 2.4908e-06 | 9.5367e-07 | 0.99109 | 34.6497 | 9.68e-05 | #VALUE!
2012-12#126 | 29 | P12 | 236 | 2012-07-11 11:07 | 24C | 0.98958 | 2.8901e-06 | 9.5367e-07 | 0.98958 | 34.5904 | 0.00011299 | #VALUE!
2012-12#114 | 29 | P12 | 239 | 2012-07-11 10:35 | 24C | 0.98398 | 2.4909e-06 | 9.5367e-07 | 0.98398 | 34.3705 | 9.6791e-05 | #VALUE!
2012-12#120 | 29 | P12 | 239 | 2012-07-11 10:51 | 24C | 0.98397 | 3.4657e-06 | 9.5367e-07 | 0.98397 | 34.37 | 0.00013629 | #VALUE!
2012-12#107 | 29 | P12 | 240 | 2012-07-11 09:56 | 24C | 0.98133 | 2.7425e-06 | 9.5367e-07 | 0.98133 | 34.2665 | 0.00010657 | #VALUE!
2012-12#101 | 29 | P12 | 254 | 2012-07-11 09:39 | 24C | 0.94017 | 3.8651e-06 | 9.5367e-07 | 0.94017 | 32.6579 | 0.00015111 | #VALUE!
2012-12#96 | 34 | P12 | 279 | 2012-07-11 09:26 | 24C | 0.98957 | 1.1203e-06 | 9.5367e-07 | 0.98957 | 34.5898 | 4.3505e-05 | #VALUE!
2012-12#137 | 34 | P12 | 282 | 2012-07-11 11:35 | 24C | 0.98389 | 2.4908e-06 | 9.5367e-07 | 0.9839 | 34.367 | 9.7029e-05 | #VALUE!
2012-12#131 | 34 | P12 | 285 | 2012-07-11 11:20 | 24C | 0.97608 | 1.3712e-06 | 9.5367e-07 | 0.97608 | 34.0605 | 5.4326e-05 | #VALUE!
2012-12#97 | 34 | P12 | 288 | 2012-07-11 09:29 | 24C | 0.94026 | 3.2141e-06 | 9.5367e-07 | 0.94026 | 32.6613 | 0.00012313 | #VALUE!
2012-12#106 | 36 | P12 | 291 | 2012-07-11 09:53 | 24C | 0.98144 | 2.4908e-06 | 9.5367e-07 | 0.98144 | 34.2706 | 9.5813e-05 | #VALUE!
2012-12#113 | 36 | P12 | 292 | 2012-07-11 10:33 | 24C | 0.97817 | 2.7267e-06 | 9.5367e-07 | 0.97817 | 34.1423 | 0.00010658 | #VALUE!
2012-12#119 | 36 | P12 | 293 | 2012-07-11 10:48 | 24C | 0.97399 | 2.2259e-06 | 9.5367e-07 | 0.97399 | 33.9788 | 8.7866e-05 | #VALUE!
2012-12#125 | 36 | P12 | 294 | 2012-07-11 11:04 | 24C | 0.97141 | 2.2397e-06 | 9.5367e-07 | 0.97141 | 33.8777 | 8.5352e-05 | #VALUE!
2012-12#100 | 38 | P13 | 327 | 2012-07-11 09:37 | 24C | 0.9402 | 1.3549e-06 | 9.5367e-07 | 0.9402 | 32.6591 | 5.2235e-05 | #VALUE!
2012-12#94 | 39 | P14 | 329 | 2012-07-11 09:21 | 24C | 0.94 | 2.4909e-06 | 9.5367e-07 | 0.94 | 32.6512 | 9.6599e-05 | #VALUE!
2012-12#143 | 40 | P15 | 331 | 2012-07-11 11:50 | 24C | 0.94135 | 2.2401e-06 | 9.5367e-07 | 0.94136 | 32.7039 | 8.7024e-05 | #VALUE!
2012-12#149 | 46 | P16 | 371 | 2012-07-11 12:05 | 24C | 0.99158 | 2.8543e-06 | 9.5367e-07 | 0.99158 | 34.6689 | 0.00011184 | #VALUE!
2012-12#156 | 46 | P16 | 373 | 2012-07-11 12:58 | 24C | 0.99131 | 2.4912e-06 | 9.5367e-07 | 0.99131 | 34.6583 | 9.6822e-05 | #VALUE!
2012-12#161 | 46 | P16 | 373 | 2012-07-11 13:11 | 24C | 0.99132 | 2.7432e-06 | 9.5367e-07 | 0.99132 | 34.6585 | 0.00010761 | #VALUE!
2012-12#171 | 46 | P16 | 375 | 2012-07-11 13:36 | 24C | 0.98959 | 2.4909e-06 | 9.5367e-07 | 0.98959 | 34.5907 | 9.6822e-05 | #VALUE!
2012-12#177 | 46 | P16 | 378 | 2012-07-11 13:51 | 24C | 0.98435 | 2.7425e-06 | 9.5367e-07 | 0.98435 | 34.385 | 0.00010658 | #VALUE!
2012-12#183 | 46 | P16 | 378 | 2012-07-11 14:04 | 24C | 0.98433 | 2.4911e-06 | 9.5367e-07 | 0.98433 | 34.384 | 9.87e-05 | #VALUE!
2012-12#189 | 46 | P16 | 379 | 2012-07-11 14:19 | 24C | 0.98211 | 2.2397e-06 | 9.5367e-07 | 0.98211 | 34.297 | 8.8748e-05 | #VALUE!
2012-12#144 | 46 | P16 | 393 | 2012-07-11 11:53 | 24C | 0.9417 | 2.7425e-06 | 9.5367e-07 | 0.9417 | 32.7175 | 0.00010868 | #VALUE!
2012-12#172 | 51 | P16 | 418 | 2012-07-11 13:38 | 24C | 0.99133 | 1.1396e-06 | 9.5367e-07 | 0.99134 | 34.6592 | 4.265e-05 | #VALUE!
2012-12#178 | 51 | P16 | 419 | 2012-07-11 13:53 | 24C | 0.98969 | 4.5125e-06 | 9.5367e-07 | 0.98969 | 34.5945 | 0.00017724 | #VALUE!
2012-12#184 | 51 | P16 | 420 | 2012-07-11 14:06 | 24C | 0.98858 | 2.7425e-06 | 9.5367e-07 | 0.98859 | 34.5512 | 0.00010765 | #VALUE!
2012-12#190 | 51 | P16 | 421 | 2012-07-11 14:21 | 24C | 0.98713 | 2.7425e-06 | 9.5367e-07 | 0.98713 | 34.4939 | 0.00010866 | #VALUE!
2012-12#145 | 51 | P16 | 422 | 2012-07-11 11:55 | 24C | 0.9858 | 2.2394e-06 | 9.5367e-07 | 0.9858 | 34.4418 | 8.7072e-05 | #VALUE!
2012-12#151 | 51 | P16 | 423 | 2012-07-11 12:46 | 24C | 0.9842 | 2.4908e-06 | 9.5367e-07 | 0.9842 | 34.379 | 9.87e-05 | #VALUE!
2012-12#150 | 52 | P16 | 430 | 2012-07-11 12:07 | 24C | 0.94242 | 3.2012e-06 | 9.5367e-07 | 0.94242 | 32.7456 | 0.00012421 | #VALUE!
2012-12#157 | 53 | P17 | 435 | 2012-07-11 13:01 | 24C | 0.94205 | 3.5403e-06 | 9.5367e-07 | 0.94206 | 32.7312 | 0.00013759 | #VALUE! | 3 | SAL: Surface sample. The value looks too high. Nothing like it until almost 100m level
2012-12#158 | 54 | P18 | 437 | 2012-07-11 13:03 | 24C | 0.94055 | 4.2166e-06 | 9.5367e-07 | 0.94055 | 32.6725 | 0.0001656 | #VALUE!
2012-12#162 | 55 | P19 | 439 | 2012-07-11 13:13 | 24C | 0.94049 | 2.4908e-06 | 9.5367e-07 | 0.94049 | 32.6704 | 9.6599e-05 | #VALUE!
2012-12#164 | 61 | P20 | 485 | 2012-07-11 13:18 | 24C | 0.99167 | 3.36e-06 | 9.5367e-07 | 0.99167 | 34.6725 | 0.00013052 | #VALUE!
2012-12#173 | 61 | P20 | 487 | 2012-07-11 13:41 | 24C | 0.99126 | 2.4908e-06 | 9.5367e-07 | 0.99126 | 34.6562 | 9.87e-05 | #VALUE!
2012-12#179 | 61 | P20 | 487 | 2012-07-11 13:55 | 24C | 0.99126 | 2.2391e-06 | 9.5367e-07 | 0.99126 | 34.6565 | 8.7009e-05 | #VALUE!
2012-12#185 | 61 | P20 | 489 | 2012-07-11 14:09 | 24C | 0.98959 | 4.6245e-06 | 9.5367e-07 | 0.98959 | 34.5906 | 0.00018129 | #VALUE!
2012-12#14 | 61 | P20 | 492 | 2012-07-10 10:03 | 24C | 0.98382 | 2.5034e-06 | 1.0133e-06 | 0.98382 | 34.3639 | 0.00010014 | #VALUE!
2012-12#191 | 61 | P20 | 492 | 2012-07-11 14:23 | 24C | 0.98386 | 3.4657e-06 | 9.5367e-07 | 0.98386 | 34.3658 | 0.00013861 | #VALUE!
2012-12#2 | 61 | P20 | 493 | 2012-07-10 09:30 | 24C | 0.98194 | 4.6229e-06 | 9.5367e-07 | 0.98195 | 34.2905 | 0.00018126 | #VALUE!
2012-12#8 | 61 | P20 | 507 | 2012-07-10 09:46 | 24C | 0.94043 | 3.9739e-06 | 1.0729e-06 | 0.94044 | 32.6681 | 0.00015499 | #VALUE!
2012-12#44 | 65 | P20 | 526 | 2012-07-10 11:46 | 24C | 0.99122 | 3.1418e-06 | 1.0133e-06 | 0.99123 | 34.6549 | 0.00012381 | #VALUE!
2012-12#3 | 65 | P20 | 527 | 2012-07-10 09:33 | 24C | 0.98955 | 5.2335e-06 | 9.5367e-07 | 0.98955 | 34.5892 | 0.00020463 | #VALUE!
2012-12#9 | 65 | P20 | 528 | 2012-07-10 09:49 | 24C | 0.98753 | 2.2394e-06 | 1.0729e-06 | 0.98753 | 34.5097 | 8.5352e-05 | #VALUE!
2012-12#15 | 65 | P20 | 529 | 2012-07-10 10:06 | 24C | 0.98566 | 2.8772e-06 | 9.5367e-07 | 0.98566 | 34.4364 | 0.00011236 | #VALUE!
2012-12#21 | 65 | P20 | 530 | 2012-07-10 10:44 | 24C | 0.9841 | 6.7177e-06 | 9.5367e-07 | 0.9841 | 34.375 | 0.00026444 | #VALUE!
2012-12#28 | 65 | P20 | 531 | 2012-07-10 11:04 | 24C | 0.98185 | 3.8437e-06 | 9.5367e-07 | 0.98185 | 34.2869 | 0.00015043 | #VALUE!
2012-12#34 | 65 | P20 | 532 | 2012-07-10 11:19 | 24C | 0.97909 | 2.4908e-06 | 9.5367e-07 | 0.97909 | 34.1787 | 9.9553e-05 | #VALUE!
2012-12#27 | 67 | P21 | 542 | 2012-07-10 11:01 | 24C | 0.94047 | 2.2398e-06 | 9.5367e-07 | 0.94047 | 32.6693 | 8.8728e-05 | #VALUE!
2012-12#39 | 70 | P22 | 544 | 2012-07-10 11:32 | 24C | 0.94004 | 2.0808e-06 | 1.0133e-06 | 0.94004 | 32.6526 | 8.1174e-05 | #VALUE!
2012-12#165 | 71 | P23 | 545 | 2012-07-11 13:22 | 24C | 0.98953 | 2.8897e-06 | 9.5367e-07 | 0.98953 | 34.5882 | 0.00011229 | #VALUE!
2012-12#168 | 71 | P23 | 546 | 2012-07-11 13:29 | 24C | 0.98952 | 3.8647e-06 | 9.5367e-07 | 0.98952 | 34.588 | 0.00015017 | #VALUE!
2012-12#174 | 71 | P23 | 547 | 2012-07-11 13:43 | 24C | 0.98954 | 2.0947e-06 | 9.5367e-07 | 0.98954 | 34.5887 | 8.2986e-05 | #VALUE!
2012-12#180 | 71 | P23 | 548 | 2012-07-11 13:57 | 24C | 0.98956 | 2.2405e-06 | 9.5367e-07 | 0.98956 | 34.5895 | 8.7041e-05 | #VALUE!
2012-12#186 | 71 | P23 | 549 | 2012-07-11 14:12 | 24C | 0.98955 | 3.4669e-06 | 9.5367e-07 | 0.98955 | 34.5891 | 0.00013628 | #VALUE!
2012-12#192 | 71 | P23 | 550 | 2012-07-11 14:26 | 24C | 0.98949 | 1.3713e-06 | 9.5367e-07 | 0.98949 | 34.5869 | 5.4326e-05 | #VALUE!
2012-12#195 | 71 | P23 | 551 | 2012-07-11 14:34 | 24C | 0.98947 | 1.1203e-06 | 9.5367e-07 | 0.98947 | 34.586 | 4.4374e-05 | #VALUE!
2012-12#198 | 71 | P23 | 552 | 2012-07-11 14:41 | 24C | 0.98954 | 1.3713e-06 | 9.5367e-07 | 0.98954 | 34.5889 | 5.4326e-05 | #VALUE!
2012-12#166 | 71 | P23 | 553 | 2012-07-11 13:24 | 24C | 0.98953 | 3.4352e-06 | 9.5367e-07 | 0.98953 | 34.5884 | 0.00013731 | #VALUE!
2012-12#169 | 71 | P23 | 554 | 2012-07-11 13:31 | 24C | 0.98952 | 2.89e-06 | 9.5367e-07 | 0.98952 | 34.588 | 0.00011232 | #VALUE!
2012-12#175 | 71 | P23 | 555 | 2012-07-11 13:46 | 24C | 0.98953 | 3.1311e-06 | 9.5367e-07 | 0.98953 | 34.5883 | 0.00012275 | #VALUE!
2012-12#181 | 71 | P23 | 556 | 2012-07-11 14:00 | 24C | 0.98951 | 3.1417e-06 | 9.5367e-07 | 0.98951 | 34.5875 | 0.00012446 | #VALUE!
2012-12#187 | 71 | P23 | 557 | 2012-07-11 14:14 | 24C | 0.98952 | 3.1417e-06 | 9.5367e-07 | 0.98952 | 34.5881 | 0.00012208 | #VALUE!
2012-12#193 | 71 | P23 | 558 | 2012-07-11 14:28 | 24C | 0.98944 | 3.4387e-06 | 9.5367e-07 | 0.98944 | 34.5849 | 0.00013471 | #VALUE!
2012-12#196 | 71 | P23 | 559 | 2012-07-11 14:36 | 24C | 0.98952 | 1.3713e-06 | 9.5367e-07 | 0.98952 | 34.5881 | 5.3286e-05 | #VALUE!
2012-12#199 | 71 | P23 | 560 | 2012-07-11 14:43 | 24C | 0.9895 | 2.2391e-06 | 9.5367e-07 | 0.9895 | 34.5871 | 8.7009e-05 | #VALUE!
2012-12#167 | 71 | P23 | 561 | 2012-07-11 13:26 | 24C | 0.98953 | 2.4753e-06 | 9.5367e-07 | 0.98953 | 34.5883 | 9.7831e-05 | #VALUE!
2012-12#170 | 71 | P23 | 562 | 2012-07-11 13:34 | 24C | 0.98945 | 3.1313e-06 | 9.5367e-07 | 0.98945 | 34.585 | 0.00012207 | #VALUE!
2012-12#176 | 71 | P23 | 563 | 2012-07-11 13:48 | 24C | 0.98952 | 4.0093e-06 | 9.5367e-07 | 0.98952 | 34.588 | 0.00015654 | #VALUE!
2012-12#182 | 71 | P23 | 564 | 2012-07-11 14:02 | 24C | 0.98944 | 2.4776e-06 | 9.5367e-07 | 0.98944 | 34.5847 | 9.6822e-05 | #VALUE!
2012-12#188 | 71 | P23 | 565 | 2012-07-11 14:17 | 24C | 0.98952 | 4.1892e-06 | 9.5367e-07 | 0.98952 | 34.5878 | 0.00016346 | #VALUE!
2012-12#194 | 71 | P23 | 566 | 2012-07-11 14:31 | 24C | 0.98951 | 2.4909e-06 | 9.5367e-07 | 0.98951 | 34.5877 | 9.869e-05 | #VALUE!
2012-12#197 | 71 | P23 | 567 | 2012-07-11 14:38 | 24C | 0.94061 | 2.226e-06 | 9.5367e-07 | 0.94061 | 32.6748 | 8.7067e-05 | #VALUE!
2012-12#200 | 71 | P23 | 568 | 2012-07-11 14:45 | 24C | 0.93973 | 4.2256e-06 | 9.5367e-07 | 0.93973 | 32.6407 | 0.00016481 | #VALUE!
2012-12#40 | 72 | P24 | 570 | 2012-07-10 11:35 | 24C | 0.94038 | 4.3844e-06 | 9.5367e-07 | 0.94038 | 32.6661 | 0.00017288 | #VALUE!
2012-12#45 | 73 | P25 | 575 | 2012-07-10 11:48 | 24C | 0.94017 | 2.7425e-06 | 1.0729e-06 | 0.94018 | 32.658 | 0.00010761 | #VALUE!
2012-12#4 | 74 | P35 | 577 | 2012-07-10 09:35 | 24C | 0.94106 | 6.7622e-06 | 1.0729e-06 | 0.94106 | 32.6924 | 0.00025936 | #VALUE!
2012-12#49 | 78 | P26 | 604 | 2012-07-10 13:00 | 24C | 0.99206 | 2.4908e-06 | 1.0729e-06 | 0.99207 | 34.6879 | 9.6609e-05 | #VALUE!
2012-12#55 | 78 | P26 | 605 | 2012-07-10 13:18 | 24C | 0.99195 | 3.125e-06 | 9.5367e-07 | 0.99195 | 34.6835 | 0.00012304 | #VALUE!
2012-12#61 | 78 | P26 | 605 | 2012-07-10 13:34 | 24C | 0.99192 | 3.6239e-06 | 8.3447e-07 | 0.99192 | 34.6824 | 0.00013994 | #VALUE!
2012-12#67 | 78 | P26 | 606 | 2012-07-10 13:50 | 24C | 0.9917 | 3.6104e-06 | 1.0729e-06 | 0.9917 | 34.6736 | 0.00013946 | #VALUE!
2012-12#73 | 78 | P26 | 607 | 2012-07-10 14:06 | 24C | 0.99135 | 4.4783e-06 | 8.9407e-07 | 0.99135 | 34.6598 | 0.00017401 | #VALUE!
2012-12#80 | 78 | P26 | 608 | 2012-07-10 14:47 | 24C | 0.99066 | 7.7238e-06 | 1.0133e-06 | 0.99066 | 34.6326 | 0.00030306 | #VALUE!
2012-12#98 | 78 | P26 | 609 | 2012-07-11 09:32 | 24C | 0.98954 | 1.3712e-06 | 9.5367e-07 | 0.98954 | 34.5888 | 5.4326e-05 | #VALUE!
2012-12#104 | 78 | P26 | 610 | 2012-07-11 09:48 | 24C | 0.98746 | 2.4776e-06 | 9.5367e-07 | 0.98746 | 34.5072 | 9.6864e-05 | #VALUE!
2012-12#50 | 78 | P26 | 611 | 2012-07-10 13:04 | 24C | 0.98598 | 2.7263e-06 | 9.5367e-07 | 0.98598 | 34.4487 | 0.00010658 | #VALUE!
2012-12#56 | 78 | P26 | 612 | 2012-07-10 13:21 | 24C | 0.9843 | 5.9248e-06 | 1.0133e-06 | 0.9843 | 34.3829 | 0.00023277 | #VALUE!
2012-12#62 | 78 | P26 | 613 | 2012-07-10 13:37 | 24C | 0.98231 | 5.1678e-06 | 9.5367e-07 | 0.98231 | 34.3049 | 0.00020497 | #VALUE!
2012-12#68 | 78 | P26 | 614 | 2012-07-10 13:53 | 24C | 0.97979 | 3.3454e-06 | 9.5367e-07 | 0.97979 | 34.206 | 0.00013393 | #VALUE!
2012-12#74 | 78 | P26 | 615 | 2012-07-10 14:09 | 24C | 0.97546 | 1.8038e-06 | 9.5367e-07 | 0.97546 | 34.0363 | 6.8797e-05 | #VALUE!
2012-12#81 | 78 | P26 | 616 | 2012-07-10 14:50 | 24C | 0.97197 | 1.793e-06 | 9.5367e-07 | 0.97197 | 33.8997 | 6.8803e-05 | #VALUE!
2012-12#99 | 78 | P26 | 617 | 2012-07-11 09:35 | 24C | 0.97011 | 3.1415e-06 | 9.5367e-07 | 0.97011 | 33.8266 | 0.00012208 | #VALUE!
2012-12#105 | 78 | P26 | 618 | 2012-07-11 09:50 | 24C | 0.96804 | 2.2391e-06 | 9.5367e-07 | 0.96805 | 33.7459 | 8.7013e-05 | #VALUE!
2012-12#57 | 78 | P26 | 619 | 2012-07-10 13:23 | 24C | 0.96617 | 2.7425e-06 | 9.5367e-07 | 0.96617 | 33.6724 | 0.00010761 | #VALUE!
2012-12#51 | 78 | P26 | 620 | 2012-07-10 13:07 | 24C | 0.9602 | 2.7425e-06 | 1.0133e-06 | 0.96021 | 33.4393 | 0.00010765 | #VALUE!
2012-12#35 | 78 | P26 | 621 | 2012-07-10 11:22 | 24C | 0.94517 | 6.7985e-06 | 9.5367e-07 | 0.94517 | 32.8525 | 0.00026434 | #VALUE!
2012-12#63 | 78 | P26 | 622 | 2012-07-10 13:40 | 24C | 0.94254 | 1.1631e-05 | 9.5367e-07 | 0.94254 | 32.75 | 0.0004522 | #VALUE!
2012-12#69 | 78 | P26 | 623 | 2012-07-10 13:55 | 24C | 0.94183 | 5.4433e-06 | 9.5367e-07 | 0.94183 | 32.7226 | 0.00021203 | #VALUE!
2012-12#75 | 78 | P26 | 624 | 2012-07-10 14:12 | 24C | 0.94171 | 5.8837e-06 | 9.5367e-07 | 0.94171 | 32.7179 | 0.00022759 | #VALUE!
2012-12#82 | 78 | P26 | 625 | 2012-07-10 14:53 | 24C | 0.94159 | 1.3712e-06 | 1.0729e-06 | 0.94159 | 32.713 | 5.4326e-05 | #VALUE!
2012-12#95 | 78 | P26 | 626 | 2012-07-11 09:23 | 24C | 0.94164 | 2.4916e-06 | 9.5367e-07 | 0.94164 | 32.7149 | 9.7862e-05 | #VALUE!
2012-12#103 | 78 | P26 | 627 | 2012-07-11 09:45 | 24C | 0.94156 | 2.4908e-06 | 9.5367e-07 | 0.94156 | 32.7121 | 9.7001e-05 | #VALUE!
2012-12#22 | 81 | P26 | 628 | 2012-07-10 10:47 | 24C | 0.98938 | 4.4373e-06 | 1.0729e-06 | 0.98938 | 34.5823 | 0.0001738 | #VALUE!
2012-12#29 | 81 | P26 | 631 | 2012-07-10 11:07 | 24C | 0.98431 | 8.126e-06 | 9.5367e-07 | 0.98431 | 34.3835 | 0.00031854 | #VALUE!
2012-12#16 | 81 | P26 | 634 | 2012-07-10 10:08 | 24C | 0.97729 | 7.6306e-06 | 9.5367e-07 | 0.97729 | 34.1081 | 0.00029917 | #VALUE!
2012-12#10 | 81 | P26 | 637 | 2012-07-10 09:52 | 24C | 0.94157 | 8.3782e-06 | 1.0729e-06 | 0.94158 | 32.7125 | 0.00032481 | #VALUE!
2012-12#41 | 90 | P26 | 705 | 2012-07-10 11:37 | 24C | 0.99203 | 8.107e-06 | 9.5367e-07 | 0.99203 | 34.6865 | 0.00031969 | #VALUE!
2012-12#46 | 90 | P26 | 706 | 2012-07-10 11:51 | 24C | 0.99168 | 5.9867e-06 | 1.0133e-06 | 0.99168 | 34.6726 | 0.0002335 | #VALUE!
2012-12#5 | 90 | P26 | 707 | 2012-07-10 09:37 | 24C | 0.99181 | 6.8434e-06 | 1.0133e-06 | 0.99181 | 34.6778 | 0.00026786 | #VALUE!
2012-12#11 | 90 | P26 | 708 | 2012-07-10 09:55 | 24C | 0.99068 | 7.3842e-06 | 1.0133e-06 | 0.99068 | 34.6334 | 0.00029065 | #VALUE! | 3 | cap liner crimped, salt crystals on cap
2012-12#17 | 90 | P26 | 709 | 2012-07-10 10:11 | 24C | 0.9894 | 6.5497e-06 | 1.0133e-06 | 0.9894 | 34.5831 | 0.00025579 | #VALUE!
2012-12#24 | 90 | P26 | 710 | 2012-07-10 10:53 | 24C | 0.98912 | 4.1139e-06 | 1.0729e-06 | 0.98912 | 34.5723 | 0.00015882 | #VALUE!
2012-12#6 | 96 | P26 | 721 | 2012-07-10 09:40 | 24C | 0.98553 | 3.131e-06 | 9.5367e-07 | 0.98554 | 34.4314 | 0.00012104 | #VALUE!
2012-12#12 | 96 | P26 | 722 | 2012-07-10 09:57 | 24C | 0.98408 | 2.7425e-06 | 1.0133e-06 | 0.98408 | 34.3742 | 0.00010553 | #VALUE!
2012-12#18 | 96 | P26 | 723 | 2012-07-10 10:14 | 24C | 0.98202 | 9.1386e-06 | 9.5367e-07 | 0.98202 | 34.2936 | 0.00035848 | #VALUE!
2012-12#25 | 96 | P26 | 724 | 2012-07-10 10:56 | 24C | 0.97925 | 6.25e-06 | 9.5367e-07 | 0.97925 | 34.1847 | 0.00024285 | #VALUE!
2012-12#31 | 96 | P26 | 725 | 2012-07-10 11:12 | 24C | 0.97741 | 3.131e-06 | 1.0133e-06 | 0.97741 | 34.1128 | 0.00012618 | #VALUE!
2012-12#37 | 96 | P26 | 726 | 2012-07-10 11:27 | 24C | 0.97301 | 6.3107e-06 | 9.5367e-07 | 0.97301 | 33.9404 | 0.00024841 | #VALUE!
2012-12#122 | 107 | P4 | 812 | 2012-07-11 10:57 | 24C | 0.98699 | 2.4908e-06 | 9.5367e-07 | 0.98699 | 34.4884 | 9.7659e-05 | #VALUE!
2012-12#128 | 107 | P4 | 813 | 2012-07-11 11:12 | 24C | 0.98648 | 2.2398e-06 | 9.5367e-07 | 0.98648 | 34.4686 | 8.8738e-05 | #VALUE!
2012-12#134 | 107 | P4 | 813 | 2012-07-11 11:27 | 24C | 0.98646 | 2.4744e-06 | 9.5367e-07 | 0.98646 | 34.4676 | 9.6791e-05 | #VALUE!
2012-12#140 | 107 | P4 | 814 | 2012-07-11 11:42 | 24C | 0.98645 | 4.3337e-06 | 9.5367e-07 | 0.98645 | 34.4674 | 0.00016698 | #VALUE!
2012-12#146 | 107 | P4 | 815 | 2012-07-11 11:58 | 24C | 0.98455 | 3.4661e-06 | 9.5367e-07 | 0.98455 | 34.3927 | 0.00013816 | #VALUE!
2012-12#152 | 107 | P4 | 815 | 2012-07-11 12:48 | 24C | 0.98455 | 3.5403e-06 | 9.5367e-07 | 0.98455 | 34.3928 | 0.00014029 | #VALUE!
2012-12#154 | 107 | P4 | 816 | 2012-07-11 12:54 | 24C | 0.98243 | 1.4046e-06 | 9.5367e-07 | 0.98243 | 34.3097 | 5.4326e-05 | #VALUE!
2012-12#159 | 107 | P4 | 817 | 2012-07-11 13:05 | 24C | 0.97924 | 4.2739e-06 | 9.5367e-07 | 0.97924 | 34.1843 | 0.00016866 | #VALUE!
2012-12#123 | 107 | P4 | 818 | 2012-07-11 10:59 | 24C | 0.97925 | 3.1418e-06 | 9.5367e-07 | 0.97926 | 34.185 | 0.00012379 | #VALUE! | 5 | SAL: Looks like a mis-sample; probably from 600m bottle; DO sample looks ok so bottle likely closed correctly
2012-12#129 | 107 | P4 | 819 | 2012-07-11 11:14 | 24C | 0.97276 | 3.2145e-06 | 9.5367e-07 | 0.97276 | 33.9305 | 0.00012844 | #VALUE!
2012-12#135 | 107 | P4 | 820 | 2012-07-11 11:30 | 24C | 0.9719 | 4.2633e-06 | 9.5367e-07 | 0.9719 | 33.8967 | 0.00016798 | #VALUE!
2012-12#141 | 107 | P4 | 821 | 2012-07-11 11:45 | 24C | 0.97111 | 4.7434e-06 | 9.5367e-07 | 0.97111 | 33.866 | 0.00018622 | #VALUE!
2012-12#147 | 107 | P4 | 822 | 2012-07-11 12:00 | 24C | 0.97084 | 2.2405e-06 | 9.5367e-07 | 0.97084 | 33.8553 | 8.7077e-05 | #VALUE!
2012-12#153 | 107 | P4 | 823 | 2012-07-11 12:51 | 24C | 0.96921 | 5.5698e-06 | 9.5367e-07 | 0.96921 | 33.7915 | 0.00021817 | #VALUE!
2012-12#155 | 107 | P4 | 824 | 2012-07-11 12:56 | 24C | 0.96627 | 3.1147e-06 | 9.5367e-07 | 0.96627 | 33.6765 | 0.00012035 | #VALUE!
2012-12#160 | 107 | P4 | 825 | 2012-07-11 13:08 | 24C | 0.95474 | 4.8368e-06 | 9.5367e-07 | 0.95474 | 33.2259 | 0.00018993 | #VALUE!
2012-12#124 | 107 | P4 | 826 | 2012-07-11 11:02 | 24C | 0.94071 | 2.8897e-06 | 9.5367e-07 | 0.94072 | 32.679 | 0.00011145 | #VALUE!
2012-12#130 | 107 | P4 | 827 | 2012-07-11 11:17 | 24C | 0.93887 | 3.1415e-06 | 9.5367e-07 | 0.93887 | 32.6071 | 0.00012103 | #VALUE!
2012-12#136 | 107 | P4 | 828 | 2012-07-11 11:32 | 24C | 0.91928 | 2.4912e-06 | 9.5367e-07 | 0.91928 | 31.8461 | 9.7226e-05 | #VALUE!
2012-12#142 | 107 | P4 | 829 | 2012-07-11 11:48 | 24C | 0.9188 | 1.3712e-06 | 9.5367e-07 | 0.9188 | 31.8275 | 5.3817e-05 | #VALUE!
2012-12#148 | 107 | P4 | 830 | 2012-07-11 12:03 | 24C | 0.91879 | 2.2401e-06 | 9.5367e-07 | 0.91879 | 31.827 | 8.6157e-05 | #VALUE!
2012-12#19 | 108 | P4 | 834 | 2012-07-10 10:16 | 24C | 0.98625 | 1.1648e-05 | 1.0729e-06 | 0.98625 | 34.4595 | 0.00046054 | #VALUE!
2012-12#38 | 108 | P4 | 835 | 2012-07-10 11:29 | 24C | 0.9845 | 3.2141e-06 | 1.0133e-06 | 0.9845 | 34.3908 | 0.00012599 | #VALUE!
2012-12#32 | 108 | P4 | 836 | 2012-07-10 11:14 | 24C | 0.98242 | 3.1314e-06 | 9.5367e-07 | 0.98242 | 34.3092 | 0.00012107 | #VALUE!
2012-12#26 | 108 | P4 | 837 | 2012-07-10 10:58 | 24C | 0.97936 | 2.4747e-06 | 9.5367e-07 | 0.97936 | 34.189 | 9.87e-05 | #VALUE!
2012-12#1 |  |  | DWR2012-01 | 2012-07-10 09:26 | 24C | 0.98954 | 3.6692e-06 | 9.5367e-07 | 0.98954 | 34.5885 | 0.00014345 | #VALUE!
2012-12#23 |  |  | DWR2012-01 | 2012-07-10 10:50 | 24C | 0.98954 | 2.0838e-06 | 9.5367e-07 | 0.98954 | 34.5886 | 8.1621e-05 | #VALUE!
2012-12#47 |  |  | DWR2012-01 | 2012-07-10 11:54 | 24C | 0.98939 | 3.4669e-06 | 1.0133e-06 | 0.98939 | 34.5827 | 0.00013345 | #VALUE!
2012-12#91 |  |  | DWR2012-01 | 2012-07-11 09:11 | 24C | 0.98941 | 3.2141e-06 | 9.5367e-07 | 0.98941 | 34.5836 | 0.00012667 | #VALUE!
2012-12#108 |  |  | DWR2012-01 | 2012-07-11 09:58 | 24C | 0.98942 | 2.4909e-06 | 9.5367e-07 | 0.98942 | 34.584 | 9.7904e-05 | #VALUE!
2012-12#201 |  |  | DWR2012-01 | 2012-07-11 14:48 | 24C | 0.98999 | 3.2144e-06 | 9.5367e-07 | 0.98999 | 34.6063 | 0.0001263 | #VALUE!
2012-12#48 |  |  | DWR2012-12 | 2012-07-10 11:57 | 24C | 0.98961 | 2.4909e-06 | 9.5367e-07 | 0.98961 | 34.5916 | 9.869e-05 | #VALUE!
2012-12#92 |  |  | DWR2012-12 | 2012-07-11 09:16 | 24C | 0.98954 | 2.4912e-06 | 9.5367e-07 | 0.98954 | 34.5886 | 9.869e-05 | #VALUE!
2012-12#90 |  |  | DWS2012-01 | 2012-07-10 15:14 | 24C | 0.98949 | 5.2332e-06 | 9.5367e-07 | 0.98949 | 34.5866 | 0.00020838 | #VALUE!
2012-12#78 |  |  | DWS2012-12 | 2012-07-10 14:21 | 24C | 0.98955 | 2.4918e-06 | 9.5367e-07 | 0.98955 | 34.5892 | 9.8074e-05 | #VALUE!
2012-12#163 |  |  | DWS2012-12 | 2012-07-11 13:16 | 24C | 0.98956 | 2.2401e-06 | 9.5367e-07 | 0.98956 | 34.5894 | 8.7082e-05 | #VALUE!
Calibration#1 |  |  | K15 | 2012-07-10 09:21 | 24C | 0.99978 | 2.2397e-06 | 1.0133e-06 | 0.99978 |  |  | #VALUE!
2012-12#64 |  |  | loop JF3 | 2012-07-10 13:42 | 24C | 0.87883 | 6.5601e-06 | 9.5367e-07 | 0.87883 | 30.2828 | 0.00025163 | #VALUE!
2012-12#7 |  |  | loop6 | 2012-07-10 09:43 | 24C | 0.93183 | 5.3086e-06 | 1.0133e-06 | 0.93184 | 32.3335 | 0.00020578 | #VALUE!
2012-12#13 |  |  | loop7 | 2012-07-10 10:00 | 24C | 0.91887 | 1.0735e-05 | 1.0133e-06 | 0.91887 | 31.8302 | 0.00041627 | #VALUE!
2012-12#52 |  |  | loopJF1 | 2012-07-10 13:10 | 24C | 0.88149 | 8.5835e-06 | 8.9407e-07 | 0.88149 | 30.3852 | 0.00033015 | #VALUE!
2012-12#58 |  |  | loopJF1 | 2012-07-10 13:26 | 24C | 0.88184 | 6.5475e-06 | 9.5367e-07 | 0.88184 | 30.3987 | 0.00025356 | #VALUE!
2012-12#70 |  |  | loopJF4 | 2012-07-10 13:58 | 24C | 0.89853 | 1.0243e-05 | 9.5367e-07 | 0.89853 | 31.0426 | 0.00039328 | #VALUE!
2012-12#36 |  |  | loopNofP16 | 2012-07-10 11:24 | 24C | 0.94091 | 3.2148e-06 | 1.0729e-06 | 0.94091 | 32.6867 | 0.00012685 | #VALUE!
2012-12#109 |  |  | loopP11 | 2012-07-11 10:23 | 24C | 0.93986 | 2.4912e-06 | 9.5367e-07 | 0.93986 | 32.6456 | 9.6976e-05 | #VALUE!
2012-12#43 |  |  | loopP14 | 2012-07-10 11:43 | 24C | 0.93988 | 4.2152e-06 | 9.5367e-07 | 0.93989 | 32.6467 | 0.00016282 | #VALUE!
2012-12#42 |  |  | loopP20 | 2012-07-10 11:40 | 24C | 0.94021 | 3.1414e-06 | 1.0133e-06 | 0.94021 | 32.6592 | 0.00012514 | #VALUE!
2012-12#20 |  |  | loopP21 | 2012-07-10 10:41 | 24C | 0.94049 | 9.8061e-06 | 9.5367e-07 | 0.9405 | 32.6704 | 0.00038098 | #VALUE!
2012-12#33 |  |  | loopP22 | 2012-07-10 11:17 | 24C | 0.93992 | 1.1049e-05 | 9.5367e-07 | 0.93992 | 32.6481 | 0.00043227 | #VALUE!
2012-12#30 |  |  | loopP24-23 | 2012-07-10 11:09 | 24C | 0.94021 | 2.4908e-06 | 9.5367e-07 | 0.94021 | 32.6594 | 9.8098e-05 | #VALUE!
2012-12#86 |  |  | ss319 | 2012-07-10 15:03 | 24C | 0.98939 | 4.3346e-06 | 9.5367e-07 | 0.98939 | 34.5827 | 0.00017155 | #VALUE!
2012-12#87 |  |  | ss319 | 2012-07-10 15:05 | 24C | 0.98939 | 4.5379e-06 | 9.5367e-07 | 0.98939 | 34.5827 | 0.00018014 | #VALUE!
2012-12#88 |  |  | ss319 | 2012-07-10 15:08 | 24C | 0.98954 | 6.3011e-06 | 8.9407e-07 | 0.98954 | 34.5888 | 0.00025007 | #VALUE!
2012-12#89 |  |  | ss319 | 2012-07-10 15:11 | 24C | 0.98946 | 6.0549e-06 | 8.9407e-07 | 0.98946 | 34.5855 | 0.0002378 | #VALUE!
2012-12#85 |  |  | storage study319 | 2012-07-10 15:01 | 24C | 0.98946 | 1.2515e-05 | 8.3447e-07 | 0.98946 | 34.5855 | 0.00049231 | #VALUE!

## RAW Salinometer Output
Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
K15 | 2012-07-10 09:21 | 24C | 0.99978 | 2.2397e-06 | 1.0133e-06 | 0.99978
DWR2012-01 | 2012-07-10 09:26 | 24C | 0.98954 | 3.6692e-06 | 9.5367e-07 | 0.98954 | 34.5885 | 0.00014345
493 | 2012-07-10 09:30 | 24C | 0.98194 | 4.6229e-06 | 9.5367e-07 | 0.98195 | 34.2905 | 0.00018126
527 | 2012-07-10 09:33 | 24C | 0.98955 | 5.2335e-06 | 9.5367e-07 | 0.98955 | 34.5892 | 0.00020463
577 | 2012-07-10 09:35 | 24C | 0.94106 | 6.7622e-06 | 1.0729e-06 | 0.94106 | 32.6924 | 0.00025936
707 | 2012-07-10 09:37 | 24C | 0.99181 | 6.8434e-06 | 1.0133e-06 | 0.99181 | 34.6778 | 0.00026786
721 | 2012-07-10 09:40 | 24C | 0.98553 | 3.131e-06 | 9.5367e-07 | 0.98554 | 34.4314 | 0.00012104
loop6 | 2012-07-10 09:43 | 24C | 0.93183 | 5.3086e-06 | 1.0133e-06 | 0.93184 | 32.3335 | 0.00020578
507 | 2012-07-10 09:46 | 24C | 0.94043 | 3.9739e-06 | 1.0729e-06 | 0.94044 | 32.6681 | 0.00015499
528 | 2012-07-10 09:49 | 24C | 0.98753 | 2.2394e-06 | 1.0729e-06 | 0.98753 | 34.5097 | 8.5352e-05
637 | 2012-07-10 09:52 | 24C | 0.94157 | 8.3782e-06 | 1.0729e-06 | 0.94158 | 32.7125 | 0.00032481
708 | 2012-07-10 09:55 | 24C | 0.99068 | 7.3842e-06 | 1.0133e-06 | 0.99068 | 34.6334 | 0.00029065
722 | 2012-07-10 09:57 | 24C | 0.98408 | 2.7425e-06 | 1.0133e-06 | 0.98408 | 34.3742 | 0.00010553
loop7 | 2012-07-10 10:00 | 24C | 0.91887 | 1.0735e-05 | 1.0133e-06 | 0.91887 | 31.8302 | 0.00041627
492 | 2012-07-10 10:03 | 24C | 0.98382 | 2.5034e-06 | 1.0133e-06 | 0.98382 | 34.3639 | 0.00010014
529 | 2012-07-10 10:06 | 24C | 0.98566 | 2.8772e-06 | 9.5367e-07 | 0.98566 | 34.4364 | 0.00011236
634 | 2012-07-10 10:08 | 24C | 0.97729 | 7.6306e-06 | 9.5367e-07 | 0.97729 | 34.1081 | 0.00029917
709 | 2012-07-10 10:11 | 24C | 0.9894 | 6.5497e-06 | 1.0133e-06 | 0.9894 | 34.5831 | 0.00025579
723 | 2012-07-10 10:14 | 24C | 0.98202 | 9.1386e-06 | 9.5367e-07 | 0.98202 | 34.2936 | 0.00035848
834 | 2012-07-10 10:16 | 24C | 0.98625 | 1.1648e-05 | 1.0729e-06 | 0.98625 | 34.4595 | 0.00046054
loopP21 | 2012-07-10 10:41 | 24C | 0.94049 | 9.8061e-06 | 9.5367e-07 | 0.9405 | 32.6704 | 0.00038098
530 | 2012-07-10 10:44 | 24C | 0.9841 | 6.7177e-06 | 9.5367e-07 | 0.9841 | 34.375 | 0.00026444
628 | 2012-07-10 10:47 | 24C | 0.98938 | 4.4373e-06 | 1.0729e-06 | 0.98938 | 34.5823 | 0.0001738
DWR2012-01 | 2012-07-10 10:50 | 24C | 0.98954 | 2.0838e-06 | 9.5367e-07 | 0.98954 | 34.5886 | 8.1621e-05
710 | 2012-07-10 10:53 | 24C | 0.98912 | 4.1139e-06 | 1.0729e-06 | 0.98912 | 34.5723 | 0.00015882
724 | 2012-07-10 10:56 | 24C | 0.97925 | 6.25e-06 | 9.5367e-07 | 0.97925 | 34.1847 | 0.00024285
837 | 2012-07-10 10:58 | 24C | 0.97936 | 2.4747e-06 | 9.5367e-07 | 0.97936 | 34.189 | 9.87e-05
542 | 2012-07-10 11:01 | 24C | 0.94047 | 2.2398e-06 | 9.5367e-07 | 0.94047 | 32.6693 | 8.8728e-05
531 | 2012-07-10 11:04 | 24C | 0.98185 | 3.8437e-06 | 9.5367e-07 | 0.98185 | 34.2869 | 0.00015043
631 | 2012-07-10 11:07 | 24C | 0.98431 | 8.126e-06 | 9.5367e-07 | 0.98431 | 34.3835 | 0.00031854
loopP24-23 | 2012-07-10 11:09 | 24C | 0.94021 | 2.4908e-06 | 9.5367e-07 | 0.94021 | 32.6594 | 9.8098e-05
725 | 2012-07-10 11:12 | 24C | 0.97741 | 3.131e-06 | 1.0133e-06 | 0.97741 | 34.1128 | 0.00012618
836 | 2012-07-10 11:14 | 24C | 0.98242 | 3.1314e-06 | 9.5367e-07 | 0.98242 | 34.3092 | 0.00012107
loopP22 | 2012-07-10 11:17 | 24C | 0.93992 | 1.1049e-05 | 9.5367e-07 | 0.93992 | 32.6481 | 0.00043227
532 | 2012-07-10 11:19 | 24C | 0.97909 | 2.4908e-06 | 9.5367e-07 | 0.97909 | 34.1787 | 9.9553e-05
621 | 2012-07-10 11:22 | 24C | 0.94517 | 6.7985e-06 | 9.5367e-07 | 0.94517 | 32.8525 | 0.00026434
loopNofP16 | 2012-07-10 11:24 | 24C | 0.94091 | 3.2148e-06 | 1.0729e-06 | 0.94091 | 32.6867 | 0.00012685
726 | 2012-07-10 11:27 | 24C | 0.97301 | 6.3107e-06 | 9.5367e-07 | 0.97301 | 33.9404 | 0.00024841
835 | 2012-07-10 11:29 | 24C | 0.9845 | 3.2141e-06 | 1.0133e-06 | 0.9845 | 34.3908 | 0.00012599
544 | 2012-07-10 11:32 | 24C | 0.94004 | 2.0808e-06 | 1.0133e-06 | 0.94004 | 32.6526 | 8.1174e-05
570 | 2012-07-10 11:35 | 24C | 0.94038 | 4.3844e-06 | 9.5367e-07 | 0.94038 | 32.6661 | 0.00017288
705 | 2012-07-10 11:37 | 24C | 0.99203 | 8.107e-06 | 9.5367e-07 | 0.99203 | 34.6865 | 0.00031969
loopP20 | 2012-07-10 11:40 | 24C | 0.94021 | 3.1414e-06 | 1.0133e-06 | 0.94021 | 32.6592 | 0.00012514
loopP14 | 2012-07-10 11:43 | 24C | 0.93988 | 4.2152e-06 | 9.5367e-07 | 0.93989 | 32.6467 | 0.00016282
526 | 2012-07-10 11:46 | 24C | 0.99122 | 3.1418e-06 | 1.0133e-06 | 0.99123 | 34.6549 | 0.00012381
575 | 2012-07-10 11:48 | 24C | 0.94017 | 2.7425e-06 | 1.0729e-06 | 0.94018 | 32.658 | 0.00010761
706 | 2012-07-10 11:51 | 24C | 0.99168 | 5.9867e-06 | 1.0133e-06 | 0.99168 | 34.6726 | 0.0002335
DWR2012-01 | 2012-07-10 11:54 | 24C | 0.98939 | 3.4669e-06 | 1.0133e-06 | 0.98939 | 34.5827 | 0.00013345
DWR2012-12 | 2012-07-10 11:57 | 24C | 0.98961 | 2.4909e-06 | 9.5367e-07 | 0.98961 | 34.5916 | 9.869e-05
604 | 2012-07-10 13:00 | 24C | 0.99206 | 2.4908e-06 | 1.0729e-06 | 0.99207 | 34.6879 | 9.6609e-05
611 | 2012-07-10 13:04 | 24C | 0.98598 | 2.7263e-06 | 9.5367e-07 | 0.98598 | 34.4487 | 0.00010658
620 | 2012-07-10 13:07 | 24C | 0.9602 | 2.7425e-06 | 1.0133e-06 | 0.96021 | 33.4393 | 0.00010765
loopJF1 | 2012-07-10 13:10 | 24C | 0.88149 | 8.5835e-06 | 8.9407e-07 | 0.88149 | 30.3852 | 0.00033015
118 | 2012-07-10 13:12 | 24C | 0.98635 | 2.2391e-06 | 8.9407e-07 | 0.98635 | 34.4634 | 8.7876e-05
145 | 2012-07-10 13:15 | 24C | 0.91719 | 4.9345e-06 | 8.9407e-07 | 0.91719 | 31.7648 | 0.00019339
605 | 2012-07-10 13:18 | 24C | 0.99195 | 3.125e-06 | 9.5367e-07 | 0.99195 | 34.6835 | 0.00012304
612 | 2012-07-10 13:21 | 24C | 0.9843 | 5.9248e-06 | 1.0133e-06 | 0.9843 | 34.3829 | 0.00023277
619 | 2012-07-10 13:23 | 24C | 0.96617 | 2.7425e-06 | 9.5367e-07 | 0.96617 | 33.6724 | 0.00010761
loopJF1 | 2012-07-10 13:26 | 24C | 0.88184 | 6.5475e-06 | 9.5367e-07 | 0.88184 | 30.3987 | 0.00025356
120 | 2012-07-10 13:29 | 24C | 0.98491 | 8.5621e-06 | 1.0729e-06 | 0.98491 | 34.4068 | 0.00033444
181 | 2012-07-10 13:32 | 24C | 0.91972 | 6.856e-06 | 9.5367e-07 | 0.91972 | 31.8631 | 0.00026274
605 | 2012-07-10 13:34 | 24C | 0.99192 | 3.6239e-06 | 8.3447e-07 | 0.99192 | 34.6824 | 0.00013994
613 | 2012-07-10 13:37 | 24C | 0.98231 | 5.1678e-06 | 9.5367e-07 | 0.98231 | 34.3049 | 0.00020497
622 | 2012-07-10 13:40 | 24C | 0.94254 | 1.1631e-05 | 9.5367e-07 | 0.94254 | 32.75 | 0.0004522
loop JF3 | 2012-07-10 13:42 | 24C | 0.87883 | 6.5601e-06 | 9.5367e-07 | 0.87883 | 30.2828 | 0.00025163
121 | 2012-07-10 13:45 | 24C | 0.98251 | 1.7705e-06 | 9.5367e-07 | 0.98251 | 34.3127 | 6.9481e-05
183 | 2012-07-10 13:48 | 24C | 0.91651 | 5.2339e-06 | 8.9407e-07 | 0.91651 | 31.7385 | 0.00020275
606 | 2012-07-10 13:50 | 24C | 0.9917 | 3.6104e-06 | 1.0729e-06 | 0.9917 | 34.6736 | 0.00013946
614 | 2012-07-10 13:53 | 24C | 0.97979 | 3.3454e-06 | 9.5367e-07 | 0.97979 | 34.206 | 0.00013393
623 | 2012-07-10 13:55 | 24C | 0.94183 | 5.4433e-06 | 9.5367e-07 | 0.94183 | 32.7226 | 0.00021203
loopJF4 | 2012-07-10 13:58 | 24C | 0.89853 | 1.0243e-05 | 9.5367e-07 | 0.89853 | 31.0426 | 0.00039328
120 | 2012-07-10 14:01 | 24C | 0.98484 | 1.0053e-05 | 9.5367e-07 | 0.98484 | 34.404 | 0.00039415 | 120
185 | 2012-07-10 14:04 | 24C | 0.91825 | 2.4912e-06 | 9.5367e-07 | 0.91825 | 31.8061 | 9.6511e-05
607 | 2012-07-10 14:06 | 24C | 0.99135 | 4.4783e-06 | 8.9407e-07 | 0.99135 | 34.6598 | 0.00017401
615 | 2012-07-10 14:09 | 24C | 0.97546 | 1.8038e-06 | 9.5367e-07 | 0.97546 | 34.0363 | 6.8797e-05
624 | 2012-07-10 14:12 | 24C | 0.94171 | 5.8837e-06 | 9.5367e-07 | 0.94171 | 32.7179 | 0.00022759
13 | 2012-07-10 14:15 | 24C | 0.89073 | 4.199e-06 | 9.5367e-07 | 0.89073 | 30.7415 | 0.00016215
135 | 2012-07-10 14:17 | 24C | 0.9176 | 2.4776e-06 | 9.5367e-07 | 0.9176 | 31.7809 | 9.752e-05
DWS2012-12 | 2012-07-10 14:21 | 24C | 0.98955 | 2.4918e-06 | 9.5367e-07 | 0.98955 | 34.5892 | 9.8074e-05
186 | 2012-07-10 14:44 | 24C | 0.98976 | 2.4747e-06 | 9.5367e-07 | 0.98977 | 34.5975 | 9.4885e-05
608 | 2012-07-10 14:47 | 24C | 0.99066 | 7.7238e-06 | 1.0133e-06 | 0.99066 | 34.6326 | 0.00030306
616 | 2012-07-10 14:50 | 24C | 0.97197 | 1.793e-06 | 9.5367e-07 | 0.97197 | 33.8997 | 6.8803e-05
625 | 2012-07-10 14:53 | 24C | 0.94159 | 1.3712e-06 | 1.0729e-06 | 0.94159 | 32.713 | 5.4326e-05
21 | 2012-07-10 14:55 | 24C | 0.91733 | 2.2391e-06 | 9.5367e-07 | 0.91733 | 31.7705 | 8.5313e-05
136 | 2012-07-10 14:58 | 24C | 0.98669 | 2.7425e-06 | 8.3447e-07 | 0.98669 | 34.4766 | 0.00010656
storage study319 | 2012-07-10 15:01 | 24C | 0.98946 | 1.2515e-05 | 8.3447e-07 | 0.98946 | 34.5855 | 0.00049231
ss319 | 2012-07-10 15:03 | 24C | 0.98939 | 4.3346e-06 | 9.5367e-07 | 0.98939 | 34.5827 | 0.00017155
ss319 | 2012-07-10 15:05 | 24C | 0.98939 | 4.5379e-06 | 9.5367e-07 | 0.98939 | 34.5827 | 0.00018014
ss319 | 2012-07-10 15:08 | 24C | 0.98954 | 6.3011e-06 | 8.9407e-07 | 0.98954 | 34.5888 | 0.00025007
ss319 | 2012-07-10 15:11 | 24C | 0.98946 | 6.0549e-06 | 8.9407e-07 | 0.98946 | 34.5855 | 0.0002378
DWS2012-01 | 2012-07-10 15:14 | 24C | 0.98949 | 5.2332e-06 | 9.5367e-07 | 0.98949 | 34.5866 | 0.00020838
DWR2012-01 | 2012-07-11 09:11 | 24C | 0.98941 | 3.2141e-06 | 9.5367e-07 | 0.98941 | 34.5836 | 0.00012667
DWR2012-12 | 2012-07-11 09:16 | 24C | 0.98954 | 2.4912e-06 | 9.5367e-07 | 0.98954 | 34.5886 | 9.869e-05
186 | 2012-07-11 09:18 | 24C | 0.98975 | 2.7264e-06 | 9.5367e-07 | 0.98976 | 34.5972 | 0.00010656
329 | 2012-07-11 09:21 | 24C | 0.94 | 2.4909e-06 | 9.5367e-07 | 0.94 | 32.6512 | 9.6599e-05
626 | 2012-07-11 09:23 | 24C | 0.94164 | 2.4916e-06 | 9.5367e-07 | 0.94164 | 32.7149 | 9.7862e-05
279 | 2012-07-11 09:26 | 24C | 0.98957 | 1.1203e-06 | 9.5367e-07 | 0.98957 | 34.5898 | 4.3505e-05
288 | 2012-07-11 09:29 | 24C | 0.94026 | 3.2141e-06 | 9.5367e-07 | 0.94026 | 32.6613 | 0.00012313
609 | 2012-07-11 09:32 | 24C | 0.98954 | 1.3712e-06 | 9.5367e-07 | 0.98954 | 34.5888 | 5.4326e-05
617 | 2012-07-11 09:35 | 24C | 0.97011 | 3.1415e-06 | 9.5367e-07 | 0.97011 | 33.8266 | 0.00012208
327 | 2012-07-11 09:37 | 24C | 0.9402 | 1.3549e-06 | 9.5367e-07 | 0.9402 | 32.6591 | 5.2235e-05
254 | 2012-07-11 09:39 | 24C | 0.94017 | 3.8651e-06 | 9.5367e-07 | 0.94017 | 32.6579 | 0.00015111
213 | 2012-07-11 09:42 | 24C | 0.93993 | 2.7425e-06 | 9.5367e-07 | 0.93993 | 32.6485 | 0.00010762
627 | 2012-07-11 09:45 | 24C | 0.94156 | 2.4908e-06 | 9.5367e-07 | 0.94156 | 32.7121 | 9.7001e-05
610 | 2012-07-11 09:48 | 24C | 0.98746 | 2.4776e-06 | 9.5367e-07 | 0.98746 | 34.5072 | 9.6864e-05
618 | 2012-07-11 09:50 | 24C | 0.96804 | 2.2391e-06 | 9.5367e-07 | 0.96805 | 33.7459 | 8.7013e-05
291 | 2012-07-11 09:53 | 24C | 0.98144 | 2.4908e-06 | 9.5367e-07 | 0.98144 | 34.2706 | 9.5813e-05
240 | 2012-07-11 09:56 | 24C | 0.98133 | 2.7425e-06 | 9.5367e-07 | 0.98133 | 34.2665 | 0.00010657
DWR2012-01 | 2012-07-11 09:58 | 24C | 0.98942 | 2.4909e-06 | 9.5367e-07 | 0.98942 | 34.584 | 9.7904e-05
loopP11 | 2012-07-11 10:23 | 24C | 0.93986 | 2.4912e-06 | 9.5367e-07 | 0.93986 | 32.6456 | 9.6976e-05
29 | 2012-07-11 10:25 | 24C | 0.91459 | 5.5696e-06 | 9.5367e-07 | 0.91459 | 31.6641 | 0.00021508
139 | 2012-07-11 10:28 | 24C | 0.98454 | 4.6249e-06 | 9.5367e-07 | 0.98454 | 34.3924 | 0.00018166
188 | 2012-07-11 10:30 | 24C | 0.98654 | 2.4775e-06 | 9.5367e-07 | 0.98655 | 34.4711 | 9.68e-05
292 | 2012-07-11 10:33 | 24C | 0.97817 | 2.7267e-06 | 9.5367e-07 | 0.97817 | 34.1423 | 0.00010658
239 | 2012-07-11 10:35 | 24C | 0.98398 | 2.4909e-06 | 9.5367e-07 | 0.98398 | 34.3705 | 9.6791e-05
211 | 2012-07-11 10:37 | 24C | 0.93995 | 3.1414e-06 | 9.5367e-07 | 0.93995 | 32.6493 | 0.00012313
118 | 2012-07-11 10:40 | 24C | 0.98636 | 2.7425e-06 | 9.5367e-07 | 0.98636 | 34.4638 | 0.00010865
142 | 2012-07-11 10:43 | 24C | 0.97815 | 2.8766e-06 | 9.5367e-07 | 0.97815 | 34.1418 | 0.000113
189 | 2012-07-11 10:46 | 24C | 0.98446 | 2.7425e-06 | 9.5367e-07 | 0.98446 | 34.3891 | 0.00010761
293 | 2012-07-11 10:48 | 24C | 0.97399 | 2.2259e-06 | 9.5367e-07 | 0.97399 | 33.9788 | 8.7866e-05
239 | 2012-07-11 10:51 | 24C | 0.98397 | 3.4657e-06 | 9.5367e-07 | 0.98397 | 34.37 | 0.00013629
209 | 2012-07-11 10:54 | 24C | 0.91683 | 2.7425e-06 | 9.5367e-07 | 0.91683 | 31.751 | 0.0001071
812 | 2012-07-11 10:57 | 24C | 0.98699 | 2.4908e-06 | 9.5367e-07 | 0.98699 | 34.4884 | 9.7659e-05
818 | 2012-07-11 10:59 | 24C | 0.97925 | 3.1418e-06 | 9.5367e-07 | 0.97926 | 34.185 | 0.00012379
826 | 2012-07-11 11:02 | 24C | 0.94071 | 2.8897e-06 | 9.5367e-07 | 0.94072 | 32.679 | 0.00011145
294 | 2012-07-11 11:04 | 24C | 0.97141 | 2.2397e-06 | 9.5367e-07 | 0.97141 | 33.8777 | 8.5352e-05
236 | 2012-07-11 11:07 | 24C | 0.98958 | 2.8901e-06 | 9.5367e-07 | 0.98958 | 34.5904 | 0.00011299
204 | 2012-07-11 11:09 | 24C | 0.92304 | 2.2262e-06 | 9.5367e-07 | 0.92304 | 31.9919 | 8.6589e-05
813 | 2012-07-11 11:12 | 24C | 0.98648 | 2.2398e-06 | 9.5367e-07 | 0.98648 | 34.4686 | 8.8738e-05
819 | 2012-07-11 11:14 | 24C | 0.97276 | 3.2145e-06 | 9.5367e-07 | 0.97276 | 33.9305 | 0.00012844
827 | 2012-07-11 11:17 | 24C | 0.93887 | 3.1415e-06 | 9.5367e-07 | 0.93887 | 32.6071 | 0.00012103
285 | 2012-07-11 11:20 | 24C | 0.97608 | 1.3712e-06 | 9.5367e-07 | 0.97608 | 34.0605 | 5.4326e-05
233 | 2012-07-11 11:22 | 24C | 0.99108 | 2.7425e-06 | 9.5367e-07 | 0.99108 | 34.6491 | 0.00010657
189 | 2012-07-11 11:25 | 24C | 0.98439 | 2.7425e-06 | 9.5367e-07 | 0.98439 | 34.3865 | 0.00010764
813 | 2012-07-11 11:27 | 24C | 0.98646 | 2.4744e-06 | 9.5367e-07 | 0.98646 | 34.4676 | 9.6791e-05
820 | 2012-07-11 11:30 | 24C | 0.9719 | 4.2633e-06 | 9.5367e-07 | 0.9719 | 33.8967 | 0.00016798
828 | 2012-07-11 11:32 | 24C | 0.91928 | 2.4912e-06 | 9.5367e-07 | 0.91928 | 31.8461 | 9.7226e-05
282 | 2012-07-11 11:35 | 24C | 0.98389 | 2.4908e-06 | 9.5367e-07 | 0.9839 | 34.367 | 9.7029e-05
233 | 2012-07-11 11:37 | 24C | 0.99109 | 2.4908e-06 | 9.5367e-07 | 0.99109 | 34.6497 | 9.68e-05
190 | 2012-07-11 11:40 | 24C | 0.98215 | 4.7146e-06 | 9.5367e-07 | 0.98216 | 34.2988 | 0.00018787
814 | 2012-07-11 11:42 | 24C | 0.98645 | 4.3337e-06 | 9.5367e-07 | 0.98645 | 34.4674 | 0.00016698
821 | 2012-07-11 11:45 | 24C | 0.97111 | 4.7434e-06 | 9.5367e-07 | 0.97111 | 33.866 | 0.00018622
829 | 2012-07-11 11:48 | 24C | 0.9188 | 1.3712e-06 | 9.5367e-07 | 0.9188 | 31.8275 | 5.3817e-05
331 | 2012-07-11 11:50 | 24C | 0.94135 | 2.2401e-06 | 9.5367e-07 | 0.94136 | 32.7039 | 8.7024e-05
393 | 2012-07-11 11:53 | 24C | 0.9417 | 2.7425e-06 | 9.5367e-07 | 0.9417 | 32.7175 | 0.00010868
422 | 2012-07-11 11:55 | 24C | 0.9858 | 2.2394e-06 | 9.5367e-07 | 0.9858 | 34.4418 | 8.7072e-05
815 | 2012-07-11 11:58 | 24C | 0.98455 | 3.4661e-06 | 9.5367e-07 | 0.98455 | 34.3927 | 0.00013816
822 | 2012-07-11 12:00 | 24C | 0.97084 | 2.2405e-06 | 9.5367e-07 | 0.97084 | 33.8553 | 8.7077e-05
830 | 2012-07-11 12:03 | 24C | 0.91879 | 2.2401e-06 | 9.5367e-07 | 0.91879 | 31.827 | 8.6157e-05
371 | 2012-07-11 12:05 | 24C | 0.99158 | 2.8543e-06 | 9.5367e-07 | 0.99158 | 34.6689 | 0.00011184
430 | 2012-07-11 12:07 | 24C | 0.94242 | 3.2012e-06 | 9.5367e-07 | 0.94242 | 32.7456 | 0.00012421
423 | 2012-07-11 12:46 | 24C | 0.9842 | 2.4908e-06 | 9.5367e-07 | 0.9842 | 34.379 | 9.87e-05
815 | 2012-07-11 12:48 | 24C | 0.98455 | 3.5403e-06 | 9.5367e-07 | 0.98455 | 34.3928 | 0.00014029
823 | 2012-07-11 12:51 | 24C | 0.96921 | 5.5698e-06 | 9.5367e-07 | 0.96921 | 33.7915 | 0.00021817
816 | 2012-07-11 12:54 | 24C | 0.98243 | 1.4046e-06 | 9.5367e-07 | 0.98243 | 34.3097 | 5.4326e-05
824 | 2012-07-11 12:56 | 24C | 0.96627 | 3.1147e-06 | 9.5367e-07 | 0.96627 | 33.6765 | 0.00012035
373 | 2012-07-11 12:58 | 24C | 0.99131 | 2.4912e-06 | 9.5367e-07 | 0.99131 | 34.6583 | 9.6822e-05
435 | 2012-07-11 13:01 | 24C | 0.94205 | 3.5403e-06 | 9.5367e-07 | 0.94206 | 32.7312 | 0.00013759
437 | 2012-07-11 13:03 | 24C | 0.94055 | 4.2166e-06 | 9.5367e-07 | 0.94055 | 32.6725 | 0.0001656
817 | 2012-07-11 13:05 | 24C | 0.97924 | 4.2739e-06 | 9.5367e-07 | 0.97924 | 34.1843 | 0.00016866
825 | 2012-07-11 13:08 | 24C | 0.95474 | 4.8368e-06 | 9.5367e-07 | 0.95474 | 33.2259 | 0.00018993
373 | 2012-07-11 13:11 | 24C | 0.99132 | 2.7432e-06 | 9.5367e-07 | 0.99132 | 34.6585 | 0.00010761
439 | 2012-07-11 13:13 | 24C | 0.94049 | 2.4908e-06 | 9.5367e-07 | 0.94049 | 32.6704 | 9.6599e-05
DWS2012-12 | 2012-07-11 13:16 | 24C | 0.98956 | 2.2401e-06 | 9.5367e-07 | 0.98956 | 34.5894 | 8.7082e-05
485 | 2012-07-11 13:18 | 24C | 0.99167 | 3.36e-06 | 9.5367e-07 | 0.99167 | 34.6725 | 0.00013052
545 | 2012-07-11 13:22 | 24C | 0.98953 | 2.8897e-06 | 9.5367e-07 | 0.98953 | 34.5882 | 0.00011229
553 | 2012-07-11 13:24 | 24C | 0.98953 | 3.4352e-06 | 9.5367e-07 | 0.98953 | 34.5884 | 0.00013731
561 | 2012-07-11 13:26 | 24C | 0.98953 | 2.4753e-06 | 9.5367e-07 | 0.98953 | 34.5883 | 9.7831e-05
546 | 2012-07-11 13:29 | 24C | 0.98952 | 3.8647e-06 | 9.5367e-07 | 0.98952 | 34.588 | 0.00015017
554 | 2012-07-11 13:31 | 24C | 0.98952 | 2.89e-06 | 9.5367e-07 | 0.98952 | 34.588 | 0.00011232
562 | 2012-07-11 13:34 | 24C | 0.98945 | 3.1313e-06 | 9.5367e-07 | 0.98945 | 34.585 | 0.00012207
375 | 2012-07-11 13:36 | 24C | 0.98959 | 2.4909e-06 | 9.5367e-07 | 0.98959 | 34.5907 | 9.6822e-05
418 | 2012-07-11 13:38 | 24C | 0.99133 | 1.1396e-06 | 9.5367e-07 | 0.99134 | 34.6592 | 4.265e-05
487 | 2012-07-11 13:41 | 24C | 0.99126 | 2.4908e-06 | 9.5367e-07 | 0.99126 | 34.6562 | 9.87e-05
547 | 2012-07-11 13:43 | 24C | 0.98954 | 2.0947e-06 | 9.5367e-07 | 0.98954 | 34.5887 | 8.2986e-05
555 | 2012-07-11 13:46 | 24C | 0.98953 | 3.1311e-06 | 9.5367e-07 | 0.98953 | 34.5883 | 0.00012275
563 | 2012-07-11 13:48 | 24C | 0.98952 | 4.0093e-06 | 9.5367e-07 | 0.98952 | 34.588 | 0.00015654
378 | 2012-07-11 13:51 | 24C | 0.98435 | 2.7425e-06 | 9.5367e-07 | 0.98435 | 34.385 | 0.00010658
419 | 2012-07-11 13:53 | 24C | 0.98969 | 4.5125e-06 | 9.5367e-07 | 0.98969 | 34.5945 | 0.00017724
487 | 2012-07-11 13:55 | 24C | 0.99126 | 2.2391e-06 | 9.5367e-07 | 0.99126 | 34.6565 | 8.7009e-05
548 | 2012-07-11 13:57 | 24C | 0.98956 | 2.2405e-06 | 9.5367e-07 | 0.98956 | 34.5895 | 8.7041e-05
556 | 2012-07-11 14:00 | 24C | 0.98951 | 3.1417e-06 | 9.5367e-07 | 0.98951 | 34.5875 | 0.00012446
564 | 2012-07-11 14:02 | 24C | 0.98944 | 2.4776e-06 | 9.5367e-07 | 0.98944 | 34.5847 | 9.6822e-05
378 | 2012-07-11 14:04 | 24C | 0.98433 | 2.4911e-06 | 9.5367e-07 | 0.98433 | 34.384 | 9.87e-05
420 | 2012-07-11 14:06 | 24C | 0.98858 | 2.7425e-06 | 9.5367e-07 | 0.98859 | 34.5512 | 0.00010765
489 | 2012-07-11 14:09 | 24C | 0.98959 | 4.6245e-06 | 9.5367e-07 | 0.98959 | 34.5906 | 0.00018129
549 | 2012-07-11 14:12 | 24C | 0.98955 | 3.4669e-06 | 9.5367e-07 | 0.98955 | 34.5891 | 0.00013628
557 | 2012-07-11 14:14 | 24C | 0.98952 | 3.1417e-06 | 9.5367e-07 | 0.98952 | 34.5881 | 0.00012208
565 | 2012-07-11 14:17 | 24C | 0.98952 | 4.1892e-06 | 9.5367e-07 | 0.98952 | 34.5878 | 0.00016346
379 | 2012-07-11 14:19 | 24C | 0.98211 | 2.2397e-06 | 9.5367e-07 | 0.98211 | 34.297 | 8.8748e-05
421 | 2012-07-11 14:21 | 24C | 0.98713 | 2.7425e-06 | 9.5367e-07 | 0.98713 | 34.4939 | 0.00010866
492 | 2012-07-11 14:23 | 24C | 0.98386 | 3.4657e-06 | 9.5367e-07 | 0.98386 | 34.3658 | 0.00013861
550 | 2012-07-11 14:26 | 24C | 0.98949 | 1.3713e-06 | 9.5367e-07 | 0.98949 | 34.5869 | 5.4326e-05
558 | 2012-07-11 14:28 | 24C | 0.98944 | 3.4387e-06 | 9.5367e-07 | 0.98944 | 34.5849 | 0.00013471
566 | 2012-07-11 14:31 | 24C | 0.98951 | 2.4909e-06 | 9.5367e-07 | 0.98951 | 34.5877 | 9.869e-05
551 | 2012-07-11 14:34 | 24C | 0.98947 | 1.1203e-06 | 9.5367e-07 | 0.98947 | 34.586 | 4.4374e-05
559 | 2012-07-11 14:36 | 24C | 0.98952 | 1.3713e-06 | 9.5367e-07 | 0.98952 | 34.5881 | 5.3286e-05
567 | 2012-07-11 14:38 | 24C | 0.94061 | 2.226e-06 | 9.5367e-07 | 0.94061 | 32.6748 | 8.7067e-05
552 | 2012-07-11 14:41 | 24C | 0.98954 | 1.3713e-06 | 9.5367e-07 | 0.98954 | 34.5889 | 5.4326e-05
560 | 2012-07-11 14:43 | 24C | 0.9895 | 2.2391e-06 | 9.5367e-07 | 0.9895 | 34.5871 | 8.7009e-05
568 | 2012-07-11 14:45 | 24C | 0.93973 | 4.2256e-06 | 9.5367e-07 | 0.93973 | 32.6407 | 0.00016481
DWR2012-01 | 2012-07-11 14:48 | 24C | 0.98999 | 3.2144e-06 | 9.5367e-07 | 0.98999 | 34.6063 | 0.0001263