## Notes
Analysis: | Kenny Scozzafava, 03-Oct-12 to 04-Oct-12 and 07-Nov-12 using Water Properties' Guildline 8400B Salinometer Serial Number 68572. Data processing by Kenny Scozzafava
Chief Scientist: | Marie Robert
Cruise Dates: | 14 Aug 2012 to 30 Aug 2012
2012-11-08 |  | KS
Salinometer Calibration Information
Bottle Label | Date Time | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 | 2012-10-02 09:58 | 24C | 0.99978 | 2.7425e-06 | 2.0266e-06 | 0.99978
2012-57#1 | 2012-10-02 09:59 | 24C | 0.99979 | 2.89e-06 | 1.967e-06 | 0.9998 | 34.992 | 0.00011401 | STANDARD AS SAMPLE
2012-57#2 | 2012-10-02 10:01 | 24C | 0.9998 | 2.8792e-06 | 1.967e-06 | 0.9998 | 34.9922 | 0.00011402 | STANDARD AS SAMPLE
K15 | 2012-10-04 10:11 | 24C | 0.99979 | 9.3671e-06 | 4.9472e-06 | 0.99979
K15 | 2012-10-04 10:12 | 24C | 0.99979 | 2.7426e-06 | -1.967e-06 | 0.99979
std | 2012-10-04 10:13 | 24C | 0.9998 | 2.4784e-06 | -1.9073e-06 | 0.99979 | 34.9919 | 9.8098e-05 | Ran standard as a sample
Calibration#1 | 2012-11-07 09:45 | 24C | 0.99978 | 2.2391e-06 | -1.0133e-06 | 0.99978
2012-59#1 | 2012-11-07 09:46 | 24C | 0.99979 | 3.2141e-06 | -1.0133e-06 | 0.99979 | 34.9916 | 0.0001264 | Std as sample
Deep Water Reference Data (2012-12 Line P Sample #: Station: P25)
Bottle Label | Date Time | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Julian Day subsampled into 200 mL sample bottles | Julian Day analyzed | Days stored
2012-57#104 | 2012-10-03 10:44 | 24C | 0.98987 | 3.2141e-06 | 2.0862e-06 | 0.98987 | 34.6018 | 0.00012382 | 243 | 277 | 34
2012-57#128 | 2012-10-03 12:12 | 24C | 0.98979 | 2.4912e-06 | 2.0266e-06 | 0.98979 | 34.5986 | 9.9553e-05 | 243 | 277 | 34
2012-57#163 | 2012-10-03 15:54 | 24C | 0.98966 | 2.7425e-06 | 2.0862e-06 | 0.98966 | 34.5934 | 0.00010456 | 243 | 277 | 34
2012-13#2 | 2012-10-04 10:19 | 24C | 0.98971 | 3.595e-06 | -1.9073e-06 | 0.98971 | 34.5953 | 0.00014165 | 243 | 278 | 35
2012-13#3 | 2012-10-04 10:56 | 24C | 0.98979 | 2.8795e-06 | -1.9073e-06 | 0.98979 | 34.5986 | 0.00011403 | 243 | 278 | 35
2012-13#44 | 2012-10-04 15:13 | 24C | 0.98977 | 2.5747e-06 | -2.0266e-06 | 0.98977 | 34.5977 | 0.00010258 | 243 | 278 | 35
2012-13#60 | 2012-10-04 15:55 | 24C | 0.98973 | 3.466e-06 | -2.0266e-06 | 0.98973 | 34.5961 | 0.00013386 | 243 | 278 | 35
2012-59#3 | 2012-11-07 09:54 | 24C | 0.98992 | 2.2405e-06 | -9.5367e-07 | 0.98992 | 34.6037 | 9.0453e-05 | 243 | 312 | 69
2012-59#36 | 2012-11-07 12:00 | 24C | 0.9898 | 3.5792e-06 | -9.5367e-07 | 0.9898 | 34.5988 | 0.00013904 | 243 | 312 | 69
2012-59#62 | 2012-11-07 14:04 | 24C | 0.98977 | 2.2391e-06 | -1.0133e-06 | 0.98977 | 34.5978 | 8.7013e-05 | 243 | 312 | 69
2012-59#73 | 2012-11-07 14:36 | 24C | 0.99014 | 6.664e-08 | -1.0729e-06 | 0.99014 | 34.6122 | 0 | 243 | 312 | 69
2012-59#96 | 2012-11-07 16:01 | 24C | 0.98981 | 2.2401e-06 | -9.5367e-07 | 0.98981 | 34.5991 | 8.7013e-05 | 243 | 312 | 69
2012-11-20 |  | KS
Sample 532 and 533 - There does not seem to be a record for these samples in the rosette log nor the cruise log
I found a case of 24 more 2012-13 salinity samples hidden in the corner of the salinity lab a month after the first analysis. Samples run but they may have deteriorated slightly from the first round of analysis.
Events 72 and 74 - most samples were full to the brim. I had to empty some of the sample and recap in order to mix and invert sample before being run.
It is unclear whether or not this has any affect on the concentration.
There were many samples with inserts that were pushed far into the neck of the sample bottles. I assume most have evaporated more relative to those that
were sealed properly so they are flagged as a 4. There is evidence that is true given that the #302 rep was not sealed and was markedly higher than its partner
Sample 302 - Because the replicate measurement (34.3868) had a known problem, it was not reported.
Only the first value (34.3784) was reported in the QF worksheet. This sample was left out of the duplicate analysis.
2013-01-18 |  | GG
Samples #532 and 533 come from cast #78 which was a TRM cast, so there is no regular CTD cast corresponding to it.
Salinity samples that came from TRM casts are those from casts:14, 28, 33, 46, 59, 82, 86.
2013-01-29 |  | KS
Cut TRM casts from main QF spreadsheet and placed into their own: "QF 2012-13 TRM FINAL"
This includes casts 14, 28, 33, 46, 59, 78, 82 and 86
2013-04-02 |  | KS
Sample 302 - Adjusted comment for easier interpretation by users.
2013-04-04 |  | KS
Added flag 2 and comment "SAL: analyzed a month later than other samples" to samples 13 to 134. This includes samples 69 and 71, which were missing this comment in the *.che file.
Added flag 2 and comment "SAL: analyzed a month later than other samples" to samples 102 and 103 on QF 2012-13 TRM worksheet
2013-04-22 |  | KS
The precision statement erroneously stated that mean difference was 0.0018 when in fact it was calculated to be 0.0015 PSU.
Deleted the mL/L left over from the template borrowed from the DO data
2015-03-30 |  | KS
Amended header in the QF worksheet to show that silicone inserts were used in place of the typical nylon inserts supplied by Guildline Instruments.
2022-01-31 |  | GG
297 |  | Flag 3 added to Nutrients and Salinity - already flagged 4 for oxygen - suspect Niskin closed early as correspondence to other Niskin at 3000db poor.

## QF 2012-13 SAL FINAL
Header: | Salinity samples were collected in 200 mL type ll glass bottles with reusable silicone inserts supplied by Xiamen Cuten Accessories Industry, Ltd. And
screw caps supplied by Ocean Scientific International Limited. They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer
standardized with IAPSO standard seawater 1.6 months after collection. For details including a duplicate analysis, see document QF2012-13SAL.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Salinity:Bottle | Flag:Salinity Bottle | Comments
2012-13 | 2 | P1 | 13 | 29.6665 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 3 | P2 | 20 | 31.3746 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 9 | P4 | 60 | 34.4829 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 9 | P4 | 61 | 34.4399 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 10 | P4 | 69 | 34.4693 | 26 | SAL: analyzed a month later than other samples.
2012-13 | 10 | P4 | 71 | 34.3875 | 26 | SAL: analyzed a month later than other samples.
2012-13 | 10 | P4 | 72 | 34.3072 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 13 | P4 | 90 | 34.4965 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 13 | P4 | 93 | 34.3991 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 13 | P4 | 96 | 34.1076 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 13 | P4 | 99 | 31.9715 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 17 | P6 | 128 | 31.9728 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 20 | P8 | 132 | 34.5876 | 26 | SAL: analyzed a month later than other samples.
2012-13 | 20 | P8 | 133 | 34.5133 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 20 | P8 | 134 | 34.4652 | 2 | SAL: analyzed a month later than other samples.
2012-13 | 20 | P8 | 135 | 34.368 | 6
2012-13 | 20 | P8 | 136 | 34.3004 | 4 | SAL: . Sample not sealed. Insert crushed into neck
2012-13 | 20 | P8 | 150 | 31.954
2012-13 | 24 | P10 | 158 | 31.9885
2012-13 | 29 | P12 | 200 | 34.6485 | 6
2012-13 | 29 | P12 | 201 | 34.6264
2012-13 | 29 | P12 | 202 | 34.5874
2012-13 | 29 | P12 | 203 | 34.5106
2012-13 | 29 | P12 | 204 | 34.4461
2012-13 | 29 | P12 | 205 | 34.3517 | 6
2012-13 | 32 | P12 | 223 | 34.5838
2012-13 | 32 | P12 | 226 | 34.3641
2012-13 | 32 | P12 | 229 | 34.0571
2012-13 | 32 | P12 | 232 | 32.488 | 4 | SAL: . Sample not sealed. Insert crushed into neck
2012-13 | 36 | P14 | 261 | 32.5924
2012-13 | 39 | P16 | 270 | 34.6056
2012-13 | 39 | P16 | 271 | 34.5844
2012-13 | 41 | P16 | 295 | 34.6681
2012-13 | 41 | P16 | 296 | 34.6571
2012-13 | 41 | P16 | 297 | 34.6593 | 36 | ALL: Niskin may have closed early. Poor corresondence to samples from other Niskin at same depth.
2012-13 | 41 | P16 | 298 | 34.6272
2012-13 | 41 | P16 | 299 | 34.5863
2012-13 | 41 | P16 | 300 | 34.5114 | 4 | SAL: Insert not fully pushed in. Sample not sealed
2012-13 | 41 | P16 | 301 | 34.4535
2012-13 | 41 | P16 | 302 | 34.3784 | 2 | SAL: Insert not fully pushed in. Sample not sealed
2012-13 | 48 | P18 | 355 | 32.5419
2012-13 | 51 | P20 | 358 | 34.5987
2012-13 | 51 | P20 | 359 | 34.5806 | 3 | SAL: Ratio bouncing around a lot over multiple readings. Cell looked fine.
2012-13 | 52 | P20 | 371 | 34.6716
2012-13 | 52 | P20 | 372 | 34.658
2012-13 | 52 | P20 | 373 | 34.6572 | 6
2012-13 | 52 | P20 | 374 | 34.6311
2012-13 | 52 | P20 | 375 | 34.5873
2012-13 | 52 | P20 | 376 | 34.5051
2012-13 | 52 | P20 | 377 | 34.4539
2012-13 | 52 | P20 | 378 | 34.3758 | 6
2012-13 | 61 | P22 | 450 | 32.5567
2012-13 | 64 | P24 | 468 | 32.5458
2012-13 | 65 | P25 | 469 | 34.5835
2012-13 | 65 | P25 | 470 | 34.5815 | 2 | SAL: Changed spent water bath bulb 45 minutes before this sample
2012-13 | 65 | P25 | 471 | 34.5847
2012-13 | 65 | P25 | 472 | 34.5814
2012-13 | 65 | P25 | 473 | 34.5812
2012-13 | 65 | P25 | 474 | 34.5824
2012-13 | 65 | P25 | 475 | 34.5887
2012-13 | 65 | P25 | 476 | 34.5837
2012-13 | 65 | P25 | 477 | 34.5817
2012-13 | 65 | P25 | 478 | 34.583
2012-13 | 65 | P25 | 479 | 34.5839
2012-13 | 65 | P25 | 480 | 34.5829
2012-13 | 65 | P25 | 481 | 34.5828
2012-13 | 65 | P25 | 482 | 34.5815
2012-13 | 65 | P25 | 483 | 34.5816
2012-13 | 65 | P25 | 484 | 34.5824
2012-13 | 65 | P25 | 485 | 34.5831
2012-13 | 65 | P25 | 486 | 34.5816
2012-13 | 65 | P25 | 487 | 34.5829
2012-13 | 65 | P25 | 488 | 34.5814
2012-13 | 65 | P25 | 489 | 34.5795
2012-13 | 65 | P25 | 490 | 34.5824
2012-13 | 68 | P26 | 495 | 34.5969
2012-13 | 68 | P26 | 496 | 34.5825
2012-13 | 66 | P26 | 508 | 34.6798
2012-13 | 66 | P26 | 509 | 34.6732
2012-13 | 66 | P26 | 510 | 34.6548 | 6
2012-13 | 66 | P26 | 511 | 34.6274
2012-13 | 66 | P26 | 512 | 34.5829
2012-13 | 66 | P26 | 513 | 34.5008
2012-13 | 66 | P26 | 514 | 34.4499
2012-13 | 66 | P26 | 515 | 34.3774 | 6
2012-13 | 66 | P26 | 516 | 34.3054
2012-13 | 66 | P26 | 517 | 34.1994
2012-13 | 66 | P26 | 518 | 34.0436
2012-13 | 66 | P26 | 519 | 33.9139
2012-13 | 66 | P26 | 520 | 33.8334
2012-13 | 66 | P26 | 521 | 33.76
2012-13 | 66 | P26 | 522 | 33.7027
2012-13 | 66 | P26 | 523 | 33.6324
2012-13 | 66 | P26 | 524 | 33.271
2012-13 | 66 | P26 | 525 | 32.765
2012-13 | 66 | P26 | 526 | 32.7326
2012-13 | 66 | P26 | 527 | 32.7032
2012-13 | 66 | P26 | 528 | 32.6096
2012-13 | 66 | P26 | 529 | 32.5561
2012-13 | 72 | PA006 | 583 | 33.8133 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 72 | PA006 | 584 | 33.7414 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 72 | PA006 | 585 | 33.6255 | 4 | SAL: No headspace. Emptied some sample to mix. Screw cap was also loose and insert pushed out.
2012-13 | 72 | PA006 | 586 | 33.0974 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 72 | PA006 | 587 | 32.7713 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 72 | PA006 | 588 | 32.7247 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 72 | PA006 | 589 | 32.6915 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 72 | PA006 | 590 | 32.6353
2012-13 | 72 | PA006 | 594 | 32.5611 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 73 | P26 | 597 | 34.6819
2012-13 | 73 | P26 | 598 | 34.6722
2012-13 | 73 | P26 | 599 | 34.6551
2012-13 | 73 | P26 | 600 | 34.6333 | 4 | SAL: Insert not fully pushed in. Sample not sealed
2012-13 | 74 | P26 | 601 | 34.6844 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 74 | P26 | 602 | 34.6847 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 74 | P26 | 603 | 34.6823 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 74 | P26 | 604 | 34.6783 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 74 | P26 | 605 | 34.6584 | 3 | SAL: No headspace. Emptied some sample to mix
2012-13 | 77 | P26 | 620 | 34.5878
2012-13 | 77 | P26 | 623 | 34.3863
2012-13 | 77 | P26 | 626 | 34.1317
2012-13 | 77 | P26 | 629 | 32.5693
2012-13 | 77 | P26 | 631 | 34.5873
2012-13 | 77 | P26 | 632 | 34.4889
2012-13 | 77 | P26 | 633 | 34.387
2012-13 | 77 | P26 | 634 | 34.3849
2012-13 | 77 | P26 | 635 | 34.1322 | 3 | SAL: Insert not fully pushed in
2012-13 | 79 | P26 | 641 | 34.6431
2012-13 | 79 | P26 | 642 | 34.6187
2012-13 | 79 | P26 | 643 | 34.6188
2012-13 | 79 | P26 | 644 | 34.5424
2012-13 | 87 | P4 | 673 | 34.4721
2012-13 | 87 | P4 | 674 | 34.3925 | 6
2012-13 | 87 | P4 | 675 | 34.3121 | 6
2012-13 | 87 | P4 | 676 | 34.1784

## QF 2012-13 LOOP FINAL
Header: | Salinity samples were collected in 200 mL type II glass bottles with disposable plastic inserts and screw cap
supplied by Ocean Scientific International Limited. They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer
standardized with IAPSO standard seawater 1.6 months after collection. For details including a duplicate analysis, see document QF2012-13SAL.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Salinity:Bottle | Flag:Salinity Bottle | Comments
2012-13 |  |  | LOOP 1 | 32.5203
2012-13 |  |  | LOOP 2 | 32.5267
2012-13 |  |  | LOOP 3 | 32.5141
2012-13 |  |  | LOOP 4 | 32.5306
2012-13 |  |  | LOOP 5 | 31.9535
2012-13 |  |  | LOOP | 29.7818
2012-13 |  |  | LOOP | 28.694
2012-13 |  |  | LOOP | 29.7071
2012-13 |  |  | LOOP | 29.2408

## QF 2012-13 TRM FINAL
Header: | Salinity samples were collected in 200 mL type II glass bottles with disposable plastic inserts and screw cap
supplied by Ocean Scientific International Limited. They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer
standardized with IAPSO standard seawater 1.6 months after collection. For details including a duplicate analysis, see document QF2012-13SAL.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Salinity:Bottle | Flag:Salinity Bottle | Comments
2012-13 | 14 | P4 | 102 | 34.4641 | SAL: analyzed a month later than other samples.
2012-13 | 14 | P4 | 103 | 34.3915 | SAL: analyzed a month later than other samples.
2012-13 | 28 | P12 | 186 | 34.5876
2012-13 | 28 | P12 | 187 | 34.5322
2012-13 | 33 | P12 | 234 | 34.5585
2012-13 | 33 | P12 | 235 | 34.495
2012-13 | 46 | P16 | 340 | 34.5616
2012-13 | 46 | P16 | 341 | 34.4856
2012-13 | 59 | P20 | 418 | 34.5673
2012-13 | 59 | P20 | 419 | 34.4917
2012-13 | 78 | P26 | 532 | 34.5655
2012-13 | 78 | P26 | 533 | 34.4896
2012-13 | 82 | P26 | 648 | 34.4924
2012-13 | 82 | P26 | 649 | 34.4521
2012-13 | 86 | P16 | 660 | 34.5328
2012-13 | 86 | P16 | 661 | 34.4681

## Duplicates
Precision Statement for Replicate Samples Drawn From a Single Niskin Bottle:
Mean Difference for salinity for the range 34.3109 to 34.6622 PSU was 0.0015, k = 26 with 0 outlier(s) removed
Precision Analysis and the Determination of Outliers
Precision was determined by analyzing replicate samples drawn from one Niskin
Outliers are discarded on the basis of Chauvenet's criteria. The statistic is calculated using the following formula:
The maximum deviation, dmax, is compared with the individual residuals from the original concentrations (see Table 1, columns H, I and J).
If a replicate's residual is greater than dmax, this reading can be rejected based upon Chauvenet's criterion
Table 1. Replicate Analysis for Samples Drawn From a Single Niskin Bottle | [14 empty] | Table 2. Critical values for Chauvenet statistic
Event | Station | Sample_ Number | SAL1 | SAL2 | SALAvg | Diff | Residual1 (Xi-Xm) | Residual2 (Xi-Xm) | Residual3 (Xi-Xm) | Σ(Xi-Xm)2 | n | Reject? | SAL1 Comment | SAL2 Comment | n | Chauvenet
10 | P4 | 69 | #VALUE! | 34.4691 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 3 | 1.37
10 | P4 | 71 | #VALUE! | 34.3851 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 4 | 1.53
20 | P8 | 132 | #VALUE! | 34.5867 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 5 | 1.64
20 | P8 | 135 | #VALUE! | 34.3676 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 6 | 1.73
29 | P12 | 200 | #VALUE! | 34.649 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 7 | 1.8
29 | P12 | 205 | #VALUE! | 34.3508 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 8 | 1.87
41 | P16 | 297 | #VALUE! | 34.6564 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 9 | 1.91
52 | P20 | 373 | #VALUE! | 34.6573 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 10 | 1.97
52 | P20 | 378 | #VALUE! | 34.3753 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 20 | 2.24
66 | P26 | 510 | #VALUE! | 34.655 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 30 | 2.39
66 | P26 | 515 | #VALUE! | 34.3769 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 50 | 2.57
87 | P4 | 674 | #VALUE! | 34.393 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 80 | 2.74
87 | P4 | 675 | #VALUE! | 34.3133 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | 0 |  |  |  | 100 | 2.81
[15 empty] | 150 | 2.93
[15 empty] | 200 | 3.02
[9 empty] | Sum | #VALUE!
[9 empty] | Count |  | 0
[9 empty] | σ | #VALUE!
[9 empty] | dmax/σ for n (Table 2) | 2.39
[9 empty] | Thus dmax = 2.39 x σ | #VALUE!
[9 empty] | Mean Diff | #VALUE!
[9 empty] | MIN | #VALUE!
[9 empty] | MAX | #VALUE!
0 Outlier(s) Removed
[9 empty] | Mean Diff = | #VALUE!

## Flags
Institute of Ocean Sciences - Quality Flags for Water Sampling | Last Updated: October 2011
Double click document to use.

## Salinity Data Processing
Sample ID | Event | Station | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | New Calculated Salinity | Flag | Comments
2012-59#66 | 2 | P1 | 13 | 2012-11-07 14:15 | 24C | 0.8628 | 0 | -9.5367e-07 | 0.8628 | 29.6665 | 1.0662e-06 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#88 | 3 | P2 | 20 | 2012-11-07 15:17 | 24C | 0.90711 | 3.1999e-06 | -1.0133e-06 | 0.90711 | 31.3746 | 0.00012329 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#80 | 9 | P4 | 60 | 2012-11-07 14:57 | 24C | 0.98685 | 2.4908e-06 | -1.0133e-06 | 0.98685 | 34.4829 | 9.2996e-05 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#89 | 9 | P4 | 61 | 2012-11-07 15:20 | 24C | 0.98575 | 3.4551e-06 | -1.0133e-06 | 0.98575 | 34.4399 | 0.00013929 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#77 | 10 | P4 | 69 | 2012-11-07 14:45 | 24C | 0.98651 | 2.4909e-06 | -9.5367e-07 | 0.98651 | 34.4695 | 9.869e-05 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#79 | 10 | P4 | 69 | 2012-11-07 14:54 | 24C | 0.9865 | 3.2037e-06 | -1.0133e-06 | 0.9865 | 34.4691 | 0.00012803 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#65 | 10 | P4 | 71 | 2012-11-07 14:12 | 24C | 0.98448 | 2.7103e-06 | -1.0133e-06 | 0.98448 | 34.39 | 0.00010447 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#81 | 10 | P4 | 71 | 2012-11-07 14:59 | 24C | 0.98436 | 2.0841e-06 | -9.5367e-07 | 0.98435 | 34.385 | 8.3896e-05 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#76 | 10 | P4 | 72 | 2012-11-07 14:43 | 24C | 0.98237 | 2.8897e-06 | -9.5367e-07 | 0.98237 | 34.3072 | 0.00011213 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#87 | 13 | P4 | 90 | 2012-11-07 15:15 | 24C | 0.98719 | 3.1415e-06 | -9.5367e-07 | 0.98719 | 34.4965 | 0.00012514 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#67 | 13 | P4 | 93 | 2012-11-07 14:18 | 24C | 0.98471 | 3.6104e-06 | -9.5367e-07 | 0.98471 | 34.3991 | 0.00014155 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#84 | 13 | P4 | 96 | 2012-11-07 15:07 | 24C | 0.97728 | 2.4915e-06 | -1.0133e-06 | 0.97728 | 34.1076 | 9.8725e-05 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#70 | 13 | P4 | 99 | 2012-11-07 14:27 | 24C | 0.92252 | 3.4551e-06 | -9.5367e-07 | 0.92252 | 31.9715 | 0.00013277 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#86 | 14 | P4 | 102 | 2012-11-07 15:12 | 24C | 0.98637 | 3.6104e-06 | -1.0133e-06 | 0.98637 | 34.4641 | 0.00014493 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#68 | 14 | P4 | 103 | 2012-11-07 14:21 | 24C | 0.98452 | 9.1433e-06 | -9.5367e-07 | 0.98452 | 34.3915 | 0.0003615 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#69 | 17 | P6 | 128 | 2012-11-07 14:24 | 24C | 0.92255 | 2.2391e-06 | -9.5367e-07 | 0.92255 | 31.9728 | 8.7006e-05 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#75 | 20 | P8 | 132 | 2012-11-07 14:40 | 24C | 0.98954 | 2.7426e-06 | -9.5367e-07 | 0.98954 | 34.5886 | 0.00010865 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#83 | 20 | P8 | 132 | 2012-11-07 15:04 | 24C | 0.98949 | 2.4908e-06 | -9.5367e-07 | 0.98949 | 34.5867 | 9.9745e-05 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#71 | 20 | P8 | 133 | 2012-11-07 14:30 | 24C | 0.98762 | 3.6109e-06 | -1.0729e-06 | 0.98762 | 34.5133 | 0.00014159 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-59#74 | 20 | P8 | 134 | 2012-11-07 14:38 | 24C | 0.9864 | 2.7426e-06 | -1.0133e-06 | 0.9864 | 34.4652 | 0.00010451 | #VALUE! | 2 | SAL: analyzed a month later than other samples.
2012-57#110 | 20 | P8 | 135 | 2012-10-03 11:20 | 24C | 0.98391 | 1.4046e-06 | 2.1458e-06 | 0.98391 | 34.3676 | 5.2235e-05 | #VALUE!
2012-57#98 | 20 | P8 | 135 | 2012-10-03 10:29 | 24C | 0.98393 | 2.89e-06 | 2.1458e-06 | 0.98393 | 34.3685 | 0.00011213 | #VALUE!
2012-57#135 | 20 | P8 | 136 | 2012-10-03 13:58 | 24C | 0.9822 | 2.7425e-06 | 2.0266e-06 | 0.9822 | 34.3004 | 0.00010969 | #VALUE! | 4 | SAL: . Sample not sealed. Insert crushed into neck
2012-57#158 | 20 | P8 | 150 | 2012-10-03 15:41 | 24C | 0.92206 | 2.2394e-06 | 2.0266e-06 | 0.92206 | 31.954 | 8.4447e-05 | #VALUE!
2012-57#156 | 24 | P10 | 158 | 2012-10-03 15:35 | 24C | 0.92295 | 2.7426e-06 | 2.0862e-06 | 0.92296 | 31.9885 | 0.0001071 | #VALUE!
2012-57#123 | 28 | P12 | 186 | 2012-10-03 11:58 | 24C | 0.98951 | 2.4911e-06 | 2.0862e-06 | 0.98951 | 34.5876 | 9.9543e-05 | #VALUE!
2012-57#141 | 28 | P12 | 187 | 2012-10-03 14:14 | 24C | 0.9881 | 2.7425e-06 | 2.1458e-06 | 0.9881 | 34.5322 | 0.00010657 | #VALUE!
2012-57#106 | 29 | P12 | 200 | 2012-10-03 11:10 | 24C | 0.99105 | 2.7426e-06 | 2.0266e-06 | 0.99105 | 34.648 | 0.00010447 | #VALUE!
2012-57#119 | 29 | P12 | 200 | 2012-10-03 11:44 | 24C | 0.99107 | 3.2037e-06 | 2.2054e-06 | 0.99107 | 34.649 | 0.00012709 | #VALUE!
2012-57#140 | 29 | P12 | 201 | 2012-10-03 14:11 | 24C | 0.9905 | 2.4751e-06 | 2.1458e-06 | 0.9905 | 34.6264 | 9.6976e-05 | #VALUE!
2012-57#111 | 29 | P12 | 202 | 2012-10-03 11:23 | 24C | 0.98951 | 3.2145e-06 | 2.1458e-06 | 0.98951 | 34.5874 | 0.00012777 | #VALUE!
2012-57#117 | 29 | P12 | 203 | 2012-10-03 11:40 | 24C | 0.98755 | 3.4657e-06 | 2.1458e-06 | 0.98755 | 34.5106 | 0.00013478 | #VALUE!
2012-57#165 | 29 | P12 | 204 | 2012-10-03 16:00 | 24C | 0.98591 | 2.7757e-06 | 2.1458e-06 | 0.98591 | 34.4461 | 0.00010944 | #VALUE!
2012-57#151 | 29 | P12 | 205 | 2012-10-03 15:21 | 24C | 0.98348 | 2.4754e-06 | 2.1458e-06 | 0.98348 | 34.3508 | 9.6652e-05 | #VALUE!
2012-57#86 | 29 | P12 | 205 | 2012-10-03 09:55 | 24C | 0.98353 | 1.1196e-06 | 2.1458e-06 | 0.98353 | 34.3527 | 4.4364e-05 | #VALUE!
2012-57#161 | 32 | P12 | 223 | 2012-10-03 15:49 | 24C | 0.98941 | 2.2391e-06 | 2.1458e-06 | 0.98942 | 34.5838 | 8.8738e-05 | #VALUE!
2012-57#131 | 32 | P12 | 226 | 2012-10-03 13:48 | 24C | 0.98382 | 4.1014e-06 | 2.1458e-06 | 0.98382 | 34.3641 | 0.00016222 | #VALUE!
2012-57#107 | 32 | P12 | 229 | 2012-10-03 11:13 | 24C | 0.97599 | 2.2391e-06 | 2.0862e-06 | 0.97599 | 34.0571 | 8.9601e-05 | #VALUE!
2012-57#75 | 32 | P12 | 232 | 2012-10-03 09:21 | 24C | 0.93581 | 1.1196e-06 | 2.0862e-06 | 0.93581 | 32.488 | 4.5217e-05 | #VALUE! | 4 | SAL: . Sample not sealed. Insert crushed into neck
2012-57#84 | 33 | P12 | 234 | 2012-10-03 09:50 | 24C | 0.98877 | 3.2141e-06 | 2.0862e-06 | 0.98877 | 34.5585 | 0.00012639 | #VALUE!
2012-57#149 | 33 | P12 | 235 | 2012-10-03 15:15 | 24C | 0.98715 | 2.7426e-06 | 2.1458e-06 | 0.98715 | 34.495 | 0.00010447 | #VALUE!
2012-57#138 | 36 | P14 | 261 | 2012-10-03 14:06 | 24C | 0.93849 | 2.8795e-06 | 2.0862e-06 | 0.93849 | 32.5924 | 0.00011438 | #VALUE!
2012-57#120 | 39 | P16 | 270 | 2012-10-03 11:50 | 24C | 0.98997 | 4.259e-06 | 2.0266e-06 | 0.98997 | 34.6056 | 0.00016677 | #VALUE!
2012-57#136 | 39 | P16 | 271 | 2012-10-03 14:01 | 24C | 0.98943 | 3.4657e-06 | 2.1458e-06 | 0.98943 | 34.5844 | 0.00013702 | #VALUE!
2012-13#56 | 41 | P16 | 295 | 2012-10-04 15:45 | 24C | 0.99156 | 2.4908e-06 | -2.0266e-06 | 0.99156 | 34.6681 | 9.9734e-05 | #VALUE!
2012-13#8 | 41 | P16 | 296 | 2012-10-04 11:10 | 24C | 0.99128 | 3.3456e-06 | -1.9073e-06 | 0.99128 | 34.6571 | 0.00012993 | #VALUE!
2012-13#20 | 41 | P16 | 297 | 2012-10-04 11:41 | 24C | 0.99141 | 2.2397e-06 | -1.9073e-06 | 0.99141 | 34.6622 | 9.0433e-05 | #VALUE!
2012-13#57 | 41 | P16 | 297 | 2012-10-04 15:47 | 24C | 0.99127 | 4.3338e-06 | -2.0266e-06 | 0.99126 | 34.6564 | 0.00016509 | #VALUE!
2012-13#48 | 41 | P16 | 298 | 2012-10-04 15:25 | 24C | 0.99052 | 3.6239e-06 | -2.0266e-06 | 0.99052 | 34.6272 | 0.00013993 | #VALUE!
2012-13#51 | 41 | P16 | 299 | 2012-10-04 15:33 | 24C | 0.98948 | 3.1414e-06 | -1.9073e-06 | 0.98948 | 34.5863 | 0.00012309 | #VALUE!
2012-13#53 | 41 | P16 | 300 | 2012-10-04 15:38 | 24C | 0.98757 | 2.7425e-06 | -1.9073e-06 | 0.98757 | 34.5114 | 0.00010658 | #VALUE! | 4 | SAL: Insert not pushed in. Sample not sealed
2012-13#7 | 41 | P16 | 301 | 2012-10-04 11:08 | 24C | 0.9861 | 1.3713e-06 | -1.9073e-06 | 0.9861 | 34.4535 | 5.2235e-05 | #VALUE!
2012-13#37 | 41 | P16 | 302 | 2012-10-04 14:55 | 24C | 0.98419 | 1.1532e-06 | -2.0266e-06 | 0.98418 | 34.3784 | 4.2703e-05 | #VALUE!
2012-13#46 | 41 | P16 | 302 | 2012-10-04 15:18 | 24C | 0.9844 | 1.1199e-06 | -1.967e-06 | 0.9844 | 34.3868 | 4.4374e-05 | #VALUE! | 4 | SAL: Insert crushed into neck. Sample not sealed. Not included in duplicate analysis. Not reported.
2012-13#38 | 46 | P16 | 340 | 2012-10-04 14:57 | 24C | 0.98885 | 2.2401e-06 | -1.967e-06 | 0.98885 | 34.5616 | 9.0443e-05 | #VALUE!
2012-13#52 | 46 | P16 | 341 | 2012-10-04 15:35 | 24C | 0.98692 | 2.4915e-06 | -2.0266e-06 | 0.98691 | 34.4856 | 9.4938e-05 | #VALUE!
2012-13#49 | 48 | P18 | 355 | 2012-10-04 15:27 | 24C | 0.93719 | 2.8897e-06 | -1.967e-06 | 0.93719 | 32.5419 | 0.00011404 | #VALUE!
2012-13#41 | 51 | P20 | 358 | 2012-10-04 15:05 | 24C | 0.9898 | 2.2127e-06 | -1.9073e-06 | 0.98979 | 34.5987 | 8.363e-05 | #VALUE!
2012-13#47 | 51 | P20 | 359 | 2012-10-04 15:21 | 24C | 0.98934 | 1.8095e-05 | -1.967e-06 | 0.98933 | 34.5806 | 0.00071038 | #VALUE! | 3 | SAL: Ratio bouncing around alot over multiple readings. Cell looked fine.
2012-13#5 | 52 | P20 | 371 | 2012-10-04 11:02 | 24C | 0.99165 | 2.8904e-06 | -2.0266e-06 | 0.99165 | 34.6716 | 0.00011471 | #VALUE!
2012-13#12 | 52 | P20 | 372 | 2012-10-04 11:20 | 24C | 0.99131 | 2.2397e-06 | -2.0266e-06 | 0.99131 | 34.6581 | 8.8728e-05 | #VALUE!
2012-13#30 | 52 | P20 | 373 | 2012-10-04 12:11 | 24C | 0.99128 | 1.3713e-06 | -1.9073e-06 | 0.99128 | 34.6571 | 5.4326e-05 | #VALUE!
2012-13#42 | 52 | P20 | 373 | 2012-10-04 15:08 | 24C | 0.99129 | 2.7103e-06 | -1.9073e-06 | 0.99129 | 34.6573 | 0.00010447 | #VALUE!
2012-13#39 | 52 | P20 | 374 | 2012-10-04 15:00 | 24C | 0.99062 | 2.4921e-06 | -1.9073e-06 | 0.99062 | 34.6311 | 9.7029e-05 | #VALUE!
2012-13#15 | 52 | P20 | 375 | 2012-10-04 11:28 | 24C | 0.98951 | 2.2391e-06 | -1.967e-06 | 0.9895 | 34.5873 | 8.959e-05 | #VALUE!
2012-13#10 | 52 | P20 | 376 | 2012-10-04 11:15 | 24C | 0.98741 | 3.1415e-06 | -2.0266e-06 | 0.98741 | 34.5051 | 0.00012104 | #VALUE!
2012-13#9 | 52 | P20 | 377 | 2012-10-04 11:13 | 24C | 0.98611 | 3.4389e-06 | -1.9073e-06 | 0.98611 | 34.4538 | 0.00013908 | #VALUE!
2012-13#16 | 52 | P20 | 378 | 2012-10-04 11:30 | 24C | 0.98413 | 2.4615e-06 | -1.9073e-06 | 0.98413 | 34.3763 | 9.4938e-05 | #VALUE!
2012-13#19 | 52 | P20 | 378 | 2012-10-04 11:37 | 24C | 0.98411 | 2.4918e-06 | -1.9073e-06 | 0.98411 | 34.3753 | 9.7029e-05 | #VALUE!
2012-13#27 | 59 | P20 | 418 | 2012-10-04 12:03 | 24C | 0.989 | 1.4046e-06 | -1.967e-06 | 0.989 | 34.5673 | 5.2235e-05 | #VALUE!
2012-13#6 | 59 | P20 | 419 | 2012-10-04 11:05 | 24C | 0.98707 | 2.2262e-06 | -1.9073e-06 | 0.98707 | 34.4917 | 8.7876e-05 | #VALUE!
2012-13#21 | 61 | P22 | 450 | 2012-10-04 11:44 | 24C | 0.93757 | 2.2394e-06 | -1.967e-06 | 0.93757 | 32.5567 | 8.7866e-05 | #VALUE!
2012-13#4 | 64 | P24 | 468 | 2012-10-04 10:59 | 24C | 0.9373 | 2.8792e-06 | -1.967e-06 | 0.93729 | 32.5458 | 0.00011385 | #VALUE!
2012-57#155 | 65 | P25 | 469 | 2012-10-03 15:33 | 24C | 0.98941 | 2.7262e-06 | 2.0266e-06 | 0.98941 | 34.5835 | 0.00010656 | #VALUE!
2012-57#145 | 65 | P25 | 470 | 2012-10-03 15:03 | 24C | 0.98936 | 1.1396e-06 | 2.0862e-06 | 0.98936 | 34.5815 | 4.5217e-05 | #VALUE! | 2 | SAL: Changed spent water bath bulb 45 minutes before this sample
2012-57#114 | 65 | P25 | 471 | 2012-10-03 11:31 | 24C | 0.98944 | 1.1529e-06 | 2.0266e-06 | 0.98944 | 34.5847 | 4.5217e-05 | #VALUE!
2012-57#102 | 65 | P25 | 472 | 2012-10-03 10:39 | 24C | 0.98935 | 1.1532e-06 | 2.1458e-06 | 0.98936 | 34.5814 | 4.265e-05 | #VALUE!
2012-57#92 | 65 | P25 | 473 | 2012-10-03 10:13 | 24C | 0.98935 | 2.2391e-06 | 2.0862e-06 | 0.98935 | 34.5812 | 8.8728e-05 | #VALUE!
2012-57#113 | 65 | P25 | 474 | 2012-10-03 11:28 | 24C | 0.98938 | 2.2401e-06 | 2.0862e-06 | 0.98938 | 34.5824 | 8.7025e-05 | #VALUE!
2012-57#166 | 65 | P25 | 475 | 2012-10-03 16:03 | 24C | 0.98954 | 3.3587e-06 | 2.0266e-06 | 0.98954 | 34.5887 | 0.00012795 | #VALUE!
2012-57#169 | 65 | P25 | 476 | 2012-10-03 16:11 | 24C | 0.98941 | 2.4912e-06 | 2.1458e-06 | 0.98941 | 34.5837 | 9.9543e-05 | #VALUE!
2012-57#154 | 65 | P25 | 477 | 2012-10-03 15:30 | 24C | 0.98936 | 2.7268e-06 | 2.1458e-06 | 0.98936 | 34.5817 | 0.00010658 | #VALUE!
2012-57#105 | 65 | P25 | 478 | 2012-10-03 11:07 | 24C | 0.98939 | 3.4553e-06 | 2.0862e-06 | 0.9894 | 34.583 | 0.00013516 | #VALUE!
2012-57#109 | 65 | P25 | 479 | 2012-10-03 11:18 | 24C | 0.98942 | 4.4783e-06 | 2.1458e-06 | 0.98942 | 34.5839 | 0.00017631 | #VALUE!
2012-57#103 | 65 | P25 | 480 | 2012-10-03 10:42 | 24C | 0.98939 | 1.3712e-06 | 2.0862e-06 | 0.98939 | 34.5829 | 5.2235e-05 | #VALUE!
2012-57#87 | 65 | P25 | 481 | 2012-10-03 09:58 | 24C | 0.98939 | 3.4658e-06 | 2.0862e-06 | 0.98939 | 34.5828 | 0.00013528 | #VALUE!
2012-57#171 | 65 | P25 | 482 | 2012-10-03 16:16 | 24C | 0.98936 | 2.2259e-06 | 2.0266e-06 | 0.98936 | 34.5815 | 8.7866e-05 | #VALUE!
2012-57#160 | 65 | P25 | 483 | 2012-10-03 15:46 | 24C | 0.98936 | 3.581e-06 | 2.2054e-06 | 0.98936 | 34.5816 | 0.00014324 | #VALUE!
2012-57#94 | 65 | P25 | 484 | 2012-10-03 10:18 | 24C | 0.98938 | 3.4666e-06 | 2.0266e-06 | 0.98938 | 34.5824 | 0.00013511 | #VALUE!
2012-57#100 | 65 | P25 | 485 | 2012-10-03 10:34 | 24C | 0.9894 | 2.4915e-06 | 2.0266e-06 | 0.9894 | 34.5831 | 9.4885e-05 | #VALUE!
2012-57#82 | 65 | P25 | 486 | 2012-10-03 09:44 | 24C | 0.98936 | 2.7105e-06 | 2.1458e-06 | 0.98936 | 34.5816 | 0.00011074 | #VALUE!
2012-57#80 | 65 | P25 | 487 | 2012-10-03 09:33 | 24C | 0.98939 | 2.0945e-06 | 2.1458e-06 | 0.98939 | 34.5829 | 8.3668e-05 | #VALUE!
2012-57#78 | 65 | P25 | 488 | 2012-10-03 09:28 | 24C | 0.98935 | 3.2142e-06 | 2.0862e-06 | 0.98935 | 34.5814 | 0.00012382 | #VALUE!
2012-57#125 | 65 | P25 | 489 | 2012-10-03 12:03 | 24C | 0.9893 | 1.4046e-06 | 2.1458e-06 | 0.98931 | 34.5795 | 5.7458e-05 | #VALUE!
2012-57#167 | 65 | P25 | 490 | 2012-10-03 16:06 | 24C | 0.98938 | 2.2405e-06 | 2.1458e-06 | 0.98938 | 34.5824 | 9.0443e-05 | #VALUE!
2012-13#26 | 68 | P26 | 495 | 2012-10-04 12:00 | 24C | 0.98975 | 3.8436e-06 | -1.9073e-06 | 0.98975 | 34.5969 | 0.00015239 | #VALUE!
2012-13#58 | 68 | P26 | 496 | 2012-10-04 15:50 | 24C | 0.98938 | 3.6104e-06 | -2.0266e-06 | 0.98938 | 34.5825 | 0.00014188 | #VALUE!
2012-57#93 | 66 | P26 | 508 | 2012-10-03 10:16 | 24C | 0.99186 | 2.2135e-06 | 2.1458e-06 | 0.99186 | 34.6798 | 8.5352e-05 | #VALUE!
2012-57#144 | 66 | P26 | 509 | 2012-10-03 14:24 | 24C | 0.99169 | 4.3186e-06 | 2.0862e-06 | 0.99169 | 34.6732 | 0.00016823 | #VALUE!
2012-57#130 | 66 | P26 | 510 | 2012-10-03 13:45 | 24C | 0.99121 | 1.6606e-06 | 2.0862e-06 | 0.99122 | 34.6545 | 6.649e-05 | #VALUE!
2012-57#132 | 66 | P26 | 510 | 2012-10-03 13:50 | 24C | 0.99123 | 3.4659e-06 | 2.0862e-06 | 0.99123 | 34.655 | 0.00013613 | #VALUE!
2012-57#116 | 66 | P26 | 511 | 2012-10-03 11:36 | 24C | 0.99052 | 3.0569e-06 | 2.0862e-06 | 0.99052 | 34.6274 | 0.00012172 | #VALUE!
2012-57#139 | 66 | P26 | 512 | 2012-10-03 14:08 | 24C | 0.98939 | 2.5749e-06 | 2.1458e-06 | 0.98939 | 34.5829 | 0.00010189 | #VALUE!
2012-57#121 | 66 | P26 | 513 | 2012-10-03 11:53 | 24C | 0.9873 | 2.0838e-06 | 2.0862e-06 | 0.9873 | 34.5008 | 8.1638e-05 | #VALUE!
2012-57#79 | 66 | P26 | 514 | 2012-10-03 09:31 | 24C | 0.986 | 2.2394e-06 | 2.1458e-06 | 0.98601 | 34.4499 | 8.7013e-05 | #VALUE!
2012-57#108 | 66 | P26 | 515 | 2012-10-03 11:15 | 24C | 0.98417 | 1.4046e-06 | 2.0266e-06 | 0.98417 | 34.378 | 5.4326e-05 | #VALUE!
2012-57#134 | 66 | P26 | 515 | 2012-10-03 13:55 | 24C | 0.98414 | 2.2391e-06 | 2.0862e-06 | 0.98415 | 34.3769 | 8.7077e-05 | #VALUE!
2012-57#112 | 66 | P26 | 516 | 2012-10-03 11:25 | 24C | 0.98232 | 2.7426e-06 | 2.0862e-06 | 0.98232 | 34.3054 | 0.00011082 | #VALUE!
2012-57#168 | 66 | P26 | 517 | 2012-10-03 16:08 | 24C | 0.97962 | 2.4909e-06 | 2.0862e-06 | 0.97962 | 34.1994 | 9.7462e-05 | #VALUE!
2012-57#115 | 66 | P26 | 518 | 2012-10-03 11:34 | 24C | 0.97565 | 2.7425e-06 | 2.0862e-06 | 0.97565 | 34.0436 | 0.00010656 | #VALUE!
2012-57#153 | 66 | P26 | 519 | 2012-10-03 15:27 | 24C | 0.97233 | 3.1414e-06 | 2.2054e-06 | 0.97234 | 33.9138 | 0.00012618 | #VALUE!
2012-57#122 | 66 | P26 | 520 | 2012-10-03 11:56 | 24C | 0.97028 | 2.2394e-06 | 2.0266e-06 | 0.97028 | 33.8334 | 8.5299e-05 | #VALUE!
2012-57#96 | 66 | P26 | 521 | 2012-10-03 10:23 | 24C | 0.9684 | 2.4909e-06 | 2.0862e-06 | 0.96841 | 33.76 | 9.87e-05 | #VALUE!
2012-57#124 | 66 | P26 | 522 | 2012-10-03 12:01 | 24C | 0.96694 | 2.2394e-06 | 2.0266e-06 | 0.96694 | 33.7027 | 8.9564e-05 | #VALUE!
2012-57#127 | 66 | P26 | 523 | 2012-10-03 12:09 | 24C | 0.96514 | 2.89e-06 | 2.1458e-06 | 0.96515 | 33.6324 | 0.00011606 | #VALUE!
2012-57#126 | 66 | P26 | 524 | 2012-10-03 12:06 | 24C | 0.9559 | 3.2144e-06 | 2.0862e-06 | 0.9559 | 33.271 | 0.00012554 | #VALUE!
2012-57#162 | 66 | P26 | 525 | 2012-10-03 15:52 | 24C | 0.94292 | 2.0804e-06 | 2.0862e-06 | 0.94292 | 32.765 | 7.9109e-05 | #VALUE!
2012-57#142 | 66 | P26 | 526 | 2012-10-03 14:17 | 24C | 0.94209 | 3.1999e-06 | 2.1458e-06 | 0.94209 | 32.7326 | 0.00012618 | #VALUE!
2012-57#81 | 66 | P26 | 527 | 2012-10-03 09:36 | 24C | 0.94134 | 2.1649e-06 | 2.0862e-06 | 0.94134 | 32.7032 | 8.1626e-05 | #VALUE!
2012-57#88 | 66 | P26 | 528 | 2012-10-03 10:01 | 24C | 0.93893 | 2.7426e-06 | 2.2054e-06 | 0.93893 | 32.6096 | 0.00010451 | #VALUE!
2012-57#76 | 66 | P26 | 529 | 2012-10-03 09:23 | 24C | 0.93756 | 2.7426e-06 | 2.0862e-06 | 0.93756 | 32.5561 | 0.00010865 | #VALUE!
2012-13#25 |  | P26 | 532 | 2012-10-04 11:58 | 24C | 0.98895 | 1.4046e-06 | -1.9073e-06 | 0.98895 | 34.5655 | 5.5371e-05 | #VALUE!
2012-13#17 |  | P26 | 533 | 2012-10-04 11:32 | 24C | 0.98702 | 2.4915e-06 | -1.9073e-06 | 0.98702 | 34.4896 | 9.6976e-05 | #VALUE!
2012-57#95 | 72 | PA006 | 583 | 2012-10-03 10:21 | 24C | 0.96977 | 2.2391e-06 | 2.1458e-06 | 0.96977 | 33.8133 | 9.0433e-05 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#147 | 72 | PA006 | 584 | 2012-10-03 15:09 | 24C | 0.96793 | 3.6104e-06 | 2.0266e-06 | 0.96793 | 33.7414 | 0.00014304 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#159 | 72 | PA006 | 585 | 2012-10-03 15:44 | 24C | 0.96497 | 3.2145e-06 | 2.1458e-06 | 0.96497 | 33.6255 | 0.00012387 | #VALUE! | 4 | SAL: No headspace. Emptied some sample to mix. Screw cap was also loose and insert pushed out.
2012-57#77 | 72 | PA006 | 586 | 2012-10-03 09:26 | 24C | 0.95145 | 2.4911e-06 | 2.1458e-06 | 0.95145 | 33.0974 | 9.869e-05 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#164 | 72 | PA006 | 587 | 2012-10-03 15:57 | 24C | 0.94308 | 2.407e-06 | 2.0862e-06 | 0.94309 | 32.7713 | 9.3191e-05 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#143 | 72 | PA006 | 588 | 2012-10-03 14:21 | 24C | 0.94189 | 3.4657e-06 | 2.1458e-06 | 0.94189 | 32.7247 | 0.00013596 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#83 | 72 | PA006 | 589 | 2012-10-03 09:47 | 24C | 0.94104 | 2.4776e-06 | 2.0862e-06 | 0.94104 | 32.6915 | 9.7029e-05 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#97 | 72 | PA006 | 590 | 2012-10-03 10:26 | 24C | 0.93959 | 3.046e-06 | 2.0862e-06 | 0.93959 | 32.6353 | 0.00011668 | #VALUE!
2012-57#137 | 72 | PA006 | 594 | 2012-10-03 14:03 | 24C | 0.93769 | 1.8038e-06 | 2.0266e-06 | 0.93769 | 32.5611 | 6.6764e-05 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-13#24 | 73 | P26 | 597 | 2012-10-04 11:55 | 24C | 0.99192 | 2.7425e-06 | -1.9073e-06 | 0.99191 | 34.6819 | 0.00010761 | #VALUE!
2012-13#54 | 73 | P26 | 598 | 2012-10-04 15:40 | 24C | 0.99167 | 2.7426e-06 | -2.0266e-06 | 0.99167 | 34.6722 | 0.00010447 | #VALUE!
2012-13#13 | 73 | P26 | 599 | 2012-10-04 11:23 | 24C | 0.99123 | 3.3587e-06 | -1.967e-06 | 0.99123 | 34.6551 | 0.00013098 | #VALUE!
2012-13#11 | 73 | P26 | 600 | 2012-10-04 11:18 | 24C | 0.99068 | 2.8769e-06 | -2.0266e-06 | 0.99067 | 34.6333 | 0.00011351 | #VALUE! | 4 | SAL: Insert not pushed in. Sample not sealed
2012-57#157 | 74 | P26 | 601 | 2012-10-03 15:38 | 24C | 0.99197 | 3.4657e-06 | 2.0862e-06 | 0.99198 | 34.6844 | 0.00013822 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#89 | 74 | P26 | 602 | 2012-10-03 10:04 | 24C | 0.99198 | 2.4777e-06 | 2.0266e-06 | 0.99198 | 34.6847 | 9.4938e-05 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#170 | 74 | P26 | 603 | 2012-10-03 16:13 | 24C | 0.99192 | 1.3712e-06 | 2.0862e-06 | 0.99192 | 34.6823 | 5.4326e-05 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#152 | 74 | P26 | 604 | 2012-10-03 15:24 | 24C | 0.99182 | 3.2149e-06 | 2.0862e-06 | 0.99182 | 34.6782 | 0.00012805 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-57#91 | 74 | P26 | 605 | 2012-10-03 10:10 | 24C | 0.99131 | 1.3713e-06 | 2.0266e-06 | 0.99131 | 34.6584 | 5.4326e-05 | #VALUE! | 3 | SAL: No headspace. Emptied some sample to mix
2012-13#55 | 77 | P26 | 620 | 2012-10-04 15:42 | 24C | 0.98952 | 3.3588e-06 | -1.967e-06 | 0.98952 | 34.5878 | 0.00012994 | #VALUE!
2012-13#50 | 77 | P26 | 623 | 2012-10-04 15:30 | 24C | 0.98439 | 2.4908e-06 | -2.0266e-06 | 0.98439 | 34.3863 | 9.9734e-05 | #VALUE!
2012-13#29 | 77 | P26 | 626 | 2012-10-04 12:08 | 24C | 0.9779 | 1.4046e-06 | -1.9073e-06 | 0.97789 | 34.1317 | 5.4326e-05 | #VALUE!
2012-13#31 | 77 | P26 | 629 | 2012-10-04 12:14 | 24C | 0.9379 | 2.4909e-06 | -2.0266e-06 | 0.9379 | 32.5693 | 9.4066e-05 | #VALUE!
2012-57#146 | 77 | P26 | 631 | 2012-10-03 15:06 | 24C | 0.9895 | 2.4909e-06 | 2.0862e-06 | 0.9895 | 34.5873 | 9.4885e-05 | #VALUE!
2012-57#99 | 77 | P26 | 632 | 2012-10-03 10:31 | 24C | 0.987 | 3.2148e-06 | 2.0266e-06 | 0.987 | 34.4889 | 0.000126 | #VALUE!
2012-57#150 | 77 | P26 | 633 | 2012-10-03 15:18 | 24C | 0.9844 | 3.8542e-06 | 2.0266e-06 | 0.98441 | 34.387 | 0.00014997 | #VALUE!
2012-57#129 | 77 | P26 | 634 | 2012-10-03 13:42 | 24C | 0.98435 | 2.4911e-06 | 2.0862e-06 | 0.98435 | 34.3849 | 9.869e-05 | #VALUE!
2012-57#90 | 77 | P26 | 635 | 2012-10-03 10:07 | 24C | 0.97791 | 1.1199e-06 | 2.1458e-06 | 0.97791 | 34.1322 | 4.4364e-05 | #VALUE! | 3 | SAL: Insert not fully pushed in
2012-57#118 | 79 | P26 | 641 | 2012-10-03 11:42 | 24C | 0.99092 | 2.4908e-06 | 2.0266e-06 | 0.99093 | 34.6431 | 9.87e-05 | #VALUE!
2012-57#133 | 79 | P26 | 642 | 2012-10-03 13:53 | 24C | 0.9903 | 2.4909e-06 | 2.0862e-06 | 0.9903 | 34.6187 | 9.869e-05 | #VALUE!
2012-57#148 | 79 | P26 | 643 | 2012-10-03 15:12 | 24C | 0.99031 | 1.373e-05 | 2.1458e-06 | 0.99031 | 34.6188 | 0.00054054 | #VALUE!
2012-57#101 | 79 | P26 | 644 | 2012-10-03 10:36 | 24C | 0.98836 | 2.666e-06 | 2.0862e-06 | 0.98836 | 34.5424 | 0.0001056 | #VALUE!
2012-13#34 | 82 | P26 | 648 | 2012-10-04 14:47 | 24C | 0.98709 | 3.4659e-06 | -2.0862e-06 | 0.98709 | 34.4924 | 0.00013824 | #VALUE!
2012-13#14 | 82 | P26 | 649 | 2012-10-04 11:25 | 24C | 0.98607 | 2.7261e-06 | -1.967e-06 | 0.98606 | 34.4521 | 0.00010656 | #VALUE!
2012-13#22 | 86 | P16 | 660 | 2012-10-04 11:46 | 24C | 0.98812 | 2.2391e-06 | -1.967e-06 | 0.98812 | 34.5328 | 8.7067e-05 | #VALUE!
2012-13#40 | 86 | P16 | 661 | 2012-10-04 15:03 | 24C | 0.98647 | 3.6235e-06 | -2.0266e-06 | 0.98647 | 34.4681 | 0.00013998 | #VALUE!
2012-13#45 | 87 | P4 | 673 | 2012-10-04 15:16 | 24C | 0.98658 | 2.2394e-06 | -2.0266e-06 | 0.98657 | 34.4722 | 8.958e-05 | #VALUE!
2012-13#28 | 87 | P4 | 674 | 2012-10-04 12:06 | 24C | 0.98453 | 4.6217e-06 | -2.0266e-06 | 0.98453 | 34.3919 | 0.00018033 | #VALUE!
2012-13#33 | 87 | P4 | 674 | 2012-10-04 12:19 | 24C | 0.98456 | 2.4908e-06 | -2.0266e-06 | 0.98456 | 34.393 | 9.869e-05 | #VALUE!
2012-13#35 | 87 | P4 | 675 | 2012-10-04 14:50 | 24C | 0.98247 | 1.3882e-06 | -1.9073e-06 | 0.98246 | 34.3109 | 5.2235e-05 | #VALUE!
2012-13#36 | 87 | P4 | 675 | 2012-10-04 14:52 | 24C | 0.98253 | 1.3713e-06 | -2.0266e-06 | 0.98253 | 34.3133 | 5.5428e-05 | #VALUE!
2012-13#43 | 87 | P4 | 676 | 2012-10-04 15:10 | 24C | 0.97909 | 2.7261e-06 | -1.9073e-06 | 0.97909 | 34.1784 | 0.00010658 | #VALUE!
2012-13#18 |  |  | LOOP 1 | 2012-10-04 11:34 | 24C | 0.93664 | 2.2259e-06 | -1.9073e-06 | 0.93664 | 32.5203 | 8.7876e-05 | #VALUE!
2012-13#59 |  |  | LOOP 2 | 2012-10-04 15:52 | 24C | 0.93681 | 2.8794e-06 | -1.967e-06 | 0.9368 | 32.5267 | 0.0001132 | #VALUE!
2012-13#23 |  |  | LOOP 3 | 2012-10-04 11:52 | 24C | 0.93648 | 2.1278e-06 | -2.0266e-06 | 0.93648 | 32.5141 | 8.0956e-05 | #VALUE!
2012-57#85 |  |  | LOOP 4 | 2012-10-03 09:53 | 24C | 0.9369 | 3.3595e-06 | 2.1458e-06 | 0.9369 | 32.5306 | 0.00012993 | #VALUE!
2012-13#32 |  |  | LOOP 5 | 2012-10-04 12:16 | 24C | 0.92205 | 1.1529e-06 | -2.0862e-06 | 0.92205 | 31.9535 | 4.5861e-05 | #VALUE!
2012-59#72 |  |  | LOOP | 2012-11-07 14:32 | 24C | 0.8658 | 3.2248e-06 | -1.0133e-06 | 0.8658 | 29.7818 | 0.00012219 | #VALUE!
2012-59#78 |  |  | LOOP | 2012-11-07 14:51 | 24C | 0.8374 | 1.1199e-06 | -1.0133e-06 | 0.8374 | 28.694 | 4.265e-05 | #VALUE!
2012-59#82 |  |  | LOOP | 2012-11-07 15:02 | 24C | 0.86385 | 3.2148e-06 | -1.0729e-06 | 0.86385 | 29.7071 | 0.00012584 | #VALUE!
2012-59#85 |  |  | LOOP | 2012-11-07 15:09 | 24C | 0.85169 | 2.4909e-06 | -1.0133e-06 | 0.85169 | 29.2408 | 9.6511e-05 | #VALUE!

## RAW Salinometer Output
Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
2012-57#75 | 232 | 2012-10-03 09:21 | 24C | 0.93581 | 1.1196e-06 | 2.0862e-06 | 0.93581 | 32.488 | 4.5217e-05 | P12
2012-57#76 | 529 | 2012-10-03 09:23 | 24C | 0.93756 | 2.7426e-06 | 2.0862e-06 | 0.93756 | 32.5561 | 0.00010865 | P26
2012-57#77 | 586 | 2012-10-03 09:26 | 24C | 0.95145 | 2.4911e-06 | 2.1458e-06 | 0.95145 | 33.0974 | 9.869e-05 | PA006
2012-57#78 | 488 | 2012-10-03 09:28 | 24C | 0.98935 | 3.2142e-06 | 2.0862e-06 | 0.98935 | 34.5814 | 0.00012382 | P25
2012-57#79 | 514 | 2012-10-03 09:31 | 24C | 0.986 | 2.2394e-06 | 2.1458e-06 | 0.98601 | 34.4499 | 8.7013e-05 | P26
2012-57#80 | 487 | 2012-10-03 09:33 | 24C | 0.98939 | 2.0945e-06 | 2.1458e-06 | 0.98939 | 34.5829 | 8.3668e-05 | P25
2012-57#81 | 527 | 2012-10-03 09:36 | 24C | 0.94134 | 2.1649e-06 | 2.0862e-06 | 0.94134 | 32.7032 | 8.1626e-05 | P26
2012-57#82 | 486 | 2012-10-03 09:44 | 24C | 0.98936 | 2.7105e-06 | 2.1458e-06 | 0.98936 | 34.5816 | 0.00011074 | P25
2012-57#83 | 589 | 2012-10-03 09:47 | 24C | 0.94104 | 2.4776e-06 | 2.0862e-06 | 0.94104 | 32.6915 | 9.7029e-05 | PA006
2012-57#84 | 234 | 2012-10-03 09:50 | 24C | 0.98877 | 3.2141e-06 | 2.0862e-06 | 0.98877 | 34.5585 | 0.00012639 | P12
2012-57#85 | LOOP 4 | 2012-10-03 09:53 | 24C | 0.9369 | 3.3595e-06 | 2.1458e-06 | 0.9369 | 32.5306 | 0.00012993
2012-57#86 | 205 | 2012-10-03 09:55 | 24C | 0.98353 | 1.1196e-06 | 2.1458e-06 | 0.98353 | 34.3527 | 4.4364e-05 | P12
2012-57#87 | 481 | 2012-10-03 09:58 | 24C | 0.98939 | 3.4658e-06 | 2.0862e-06 | 0.98939 | 34.5828 | 0.00013528 | P25
2012-57#88 | 528 | 2012-10-03 10:01 | 24C | 0.93893 | 2.7426e-06 | 2.2054e-06 | 0.93893 | 32.6096 | 0.00010451 | P26
2012-57#89 | 602 | 2012-10-03 10:04 | 24C | 0.99198 | 2.4777e-06 | 2.0266e-06 | 0.99198 | 34.6847 | 9.4938e-05 | P26
2012-57#90 | 635 | 2012-10-03 10:07 | 24C | 0.97791 | 1.1199e-06 | 2.1458e-06 | 0.97791 | 34.1322 | 4.4364e-05 | P26
2012-57#91 | 605 | 2012-10-03 10:10 | 24C | 0.99131 | 1.3713e-06 | 2.0266e-06 | 0.99131 | 34.6584 | 5.4326e-05 | P26
2012-57#92 | 473 | 2012-10-03 10:13 | 24C | 0.98935 | 2.2391e-06 | 2.0862e-06 | 0.98935 | 34.5812 | 8.8728e-05 | P25
2012-57#93 | 508 | 2012-10-03 10:16 | 24C | 0.99186 | 2.2135e-06 | 2.1458e-06 | 0.99186 | 34.6798 | 8.5352e-05 | P26
2012-57#94 | 484 | 2012-10-03 10:18 | 24C | 0.98938 | 3.4666e-06 | 2.0266e-06 | 0.98938 | 34.5824 | 0.00013511 | P25
2012-57#95 | 583 | 2012-10-03 10:21 | 24C | 0.96977 | 2.2391e-06 | 2.1458e-06 | 0.96977 | 33.8133 | 9.0433e-05 | PA006
2012-57#96 | 521 | 2012-10-03 10:23 | 24C | 0.9684 | 2.4909e-06 | 2.0862e-06 | 0.96841 | 33.76 | 9.87e-05 | P26
2012-57#97 | 590 | 2012-10-03 10:26 | 24C | 0.93959 | 3.046e-06 | 2.0862e-06 | 0.93959 | 32.6353 | 0.00011668 | PA006
2012-57#98 | 135 | 2012-10-03 10:29 | 24C | 0.98393 | 2.89e-06 | 2.1458e-06 | 0.98393 | 34.3685 | 0.00011213 | P8
2012-57#99 | 632 | 2012-10-03 10:31 | 24C | 0.987 | 3.2148e-06 | 2.0266e-06 | 0.987 | 34.4889 | 0.000126 | P26
2012-57#100 | 485 | 2012-10-03 10:34 | 24C | 0.9894 | 2.4915e-06 | 2.0266e-06 | 0.9894 | 34.5831 | 9.4885e-05 | P25
2012-57#101 | 644 | 2012-10-03 10:36 | 24C | 0.98836 | 2.666e-06 | 2.0862e-06 | 0.98836 | 34.5424 | 0.0001056 | P26
2012-57#102 | 472 | 2012-10-03 10:39 | 24C | 0.98935 | 1.1532e-06 | 2.1458e-06 | 0.98936 | 34.5814 | 4.265e-05 | P25
2012-57#103 | 480 | 2012-10-03 10:42 | 24C | 0.98939 | 1.3712e-06 | 2.0862e-06 | 0.98939 | 34.5829 | 5.2235e-05 | P25
2012-57#104 | 1 | 2012-10-03 10:44 | 24C | 0.98987 | 3.2141e-06 | 2.0862e-06 | 0.98987 | 34.6018 | 0.00012382 | DWR
2012-57#105 | 478 | 2012-10-03 11:07 | 24C | 0.98939 | 3.4553e-06 | 2.0862e-06 | 0.9894 | 34.583 | 0.00013516 | P25
2012-57#106 | 200 | 2012-10-03 11:10 | 24C | 0.99105 | 2.7426e-06 | 2.0266e-06 | 0.99105 | 34.648 | 0.00010447 | P12
2012-57#107 | 229 | 2012-10-03 11:13 | 24C | 0.97599 | 2.2391e-06 | 2.0862e-06 | 0.97599 | 34.0571 | 8.9601e-05 | P12
2012-57#108 | 515 | 2012-10-03 11:15 | 24C | 0.98417 | 1.4046e-06 | 2.0266e-06 | 0.98417 | 34.378 | 5.4326e-05 | P26
2012-57#109 | 479 | 2012-10-03 11:18 | 24C | 0.98942 | 4.4783e-06 | 2.1458e-06 | 0.98942 | 34.5839 | 0.00017631 | P25
2012-57#110 | 135 | 2012-10-03 11:20 | 24C | 0.98391 | 1.4046e-06 | 2.1458e-06 | 0.98391 | 34.3676 | 5.2235e-05 | P8
2012-57#111 | 202 | 2012-10-03 11:23 | 24C | 0.98951 | 3.2145e-06 | 2.1458e-06 | 0.98951 | 34.5874 | 0.00012777 | P12
2012-57#112 | 516 | 2012-10-03 11:25 | 24C | 0.98232 | 2.7426e-06 | 2.0862e-06 | 0.98232 | 34.3054 | 0.00011082 | P26
2012-57#113 | 474 | 2012-10-03 11:28 | 24C | 0.98938 | 2.2401e-06 | 2.0862e-06 | 0.98938 | 34.5824 | 8.7025e-05 | P25
2012-57#114 | 471 | 2012-10-03 11:31 | 24C | 0.98944 | 1.1529e-06 | 2.0266e-06 | 0.98944 | 34.5847 | 4.5217e-05 | P25
2012-57#115 | 518 | 2012-10-03 11:34 | 24C | 0.97565 | 2.7425e-06 | 2.0862e-06 | 0.97565 | 34.0436 | 0.00010656 | P26
2012-57#116 | 511 | 2012-10-03 11:36 | 24C | 0.99052 | 3.0569e-06 | 2.0862e-06 | 0.99052 | 34.6274 | 0.00012172 | P26
2012-57#117 | 203 | 2012-10-03 11:40 | 24C | 0.98755 | 3.4657e-06 | 2.1458e-06 | 0.98755 | 34.5106 | 0.00013478 | P12
2012-57#118 | 641 | 2012-10-03 11:42 | 24C | 0.99092 | 2.4908e-06 | 2.0266e-06 | 0.99093 | 34.6431 | 9.87e-05 | P26
2012-57#119 | 200 | 2012-10-03 11:44 | 24C | 0.99107 | 3.2037e-06 | 2.2054e-06 | 0.99107 | 34.649 | 0.00012709 | P12
2012-57#120 | 270 | 2012-10-03 11:50 | 24C | 0.98997 | 4.259e-06 | 2.0266e-06 | 0.98997 | 34.6056 | 0.00016677 | P16
2012-57#121 | 513 | 2012-10-03 11:53 | 24C | 0.9873 | 2.0838e-06 | 2.0862e-06 | 0.9873 | 34.5008 | 8.1638e-05 | P26
2012-57#122 | 520 | 2012-10-03 11:56 | 24C | 0.97028 | 2.2394e-06 | 2.0266e-06 | 0.97028 | 33.8334 | 8.5299e-05 | P26
2012-57#123 | 186 | 2012-10-03 11:58 | 24C | 0.98951 | 2.4911e-06 | 2.0862e-06 | 0.98951 | 34.5876 | 9.9543e-05 | P12
2012-57#124 | 522 | 2012-10-03 12:01 | 24C | 0.96694 | 2.2394e-06 | 2.0266e-06 | 0.96694 | 33.7027 | 8.9564e-05 | P26
2012-57#125 | 489 | 2012-10-03 12:03 | 24C | 0.9893 | 1.4046e-06 | 2.1458e-06 | 0.98931 | 34.5795 | 5.7458e-05 | P25
2012-57#126 | 524 | 2012-10-03 12:06 | 24C | 0.9559 | 3.2144e-06 | 2.0862e-06 | 0.9559 | 33.271 | 0.00012554 | P26
2012-57#127 | 523 | 2012-10-03 12:09 | 24C | 0.96514 | 2.89e-06 | 2.1458e-06 | 0.96515 | 33.6324 | 0.00011606 | P26
2012-57#128 | 8 | 2012-10-03 12:12 | 24C | 0.98979 | 2.4912e-06 | 2.0266e-06 | 0.98979 | 34.5986 | 9.9553e-05 | DWR
2012-57#129 | 634 | 2012-10-03 13:42 | 24C | 0.98435 | 2.4911e-06 | 2.0862e-06 | 0.98435 | 34.3849 | 9.869e-05 | P26
2012-57#130 | 510 | 2012-10-03 13:45 | 24C | 0.99121 | 1.6606e-06 | 2.0862e-06 | 0.99122 | 34.6545 | 6.649e-05 | P26
2012-57#131 | 226 | 2012-10-03 13:48 | 24C | 0.98382 | 4.1014e-06 | 2.1458e-06 | 0.98382 | 34.3641 | 0.00016222 | P12
2012-57#132 | 510 | 2012-10-03 13:50 | 24C | 0.99123 | 3.4659e-06 | 2.0862e-06 | 0.99123 | 34.655 | 0.00013613 | P26
2012-57#133 | 642 | 2012-10-03 13:53 | 24C | 0.9903 | 2.4909e-06 | 2.0862e-06 | 0.9903 | 34.6187 | 9.869e-05 | P26
2012-57#134 | 515 | 2012-10-03 13:55 | 24C | 0.98414 | 2.2391e-06 | 2.0862e-06 | 0.98415 | 34.3769 | 8.7077e-05 | P26
2012-57#135 | 136 | 2012-10-03 13:58 | 24C | 0.9822 | 2.7425e-06 | 2.0266e-06 | 0.9822 | 34.3004 | 0.00010969 | P8
2012-57#136 | 271 | 2012-10-03 14:01 | 24C | 0.98943 | 3.4657e-06 | 2.1458e-06 | 0.98943 | 34.5844 | 0.00013702 | P16
2012-57#137 | 594 | 2012-10-03 14:03 | 24C | 0.93769 | 1.8038e-06 | 2.0266e-06 | 0.93769 | 32.5611 | 6.6764e-05 | PA006
2012-57#138 | 261 | 2012-10-03 14:06 | 24C | 0.93849 | 2.8795e-06 | 2.0862e-06 | 0.93849 | 32.5924 | 0.00011438 | P14
2012-57#139 | 512 | 2012-10-03 14:08 | 24C | 0.98939 | 2.5749e-06 | 2.1458e-06 | 0.98939 | 34.5829 | 0.00010189 | P26
2012-57#140 | 201 | 2012-10-03 14:11 | 24C | 0.9905 | 2.4751e-06 | 2.1458e-06 | 0.9905 | 34.6264 | 9.6976e-05 | P12
2012-57#141 | 187 | 2012-10-03 14:14 | 24C | 0.9881 | 2.7425e-06 | 2.1458e-06 | 0.9881 | 34.5322 | 0.00010657 | P12
2012-57#142 | 526 | 2012-10-03 14:17 | 24C | 0.94209 | 3.1999e-06 | 2.1458e-06 | 0.94209 | 32.7326 | 0.00012618 | P26
2012-57#143 | 588 | 2012-10-03 14:21 | 24C | 0.94189 | 3.4657e-06 | 2.1458e-06 | 0.94189 | 32.7247 | 0.00013596 | PA006
2012-57#144 | 509 | 2012-10-03 14:24 | 24C | 0.99169 | 4.3186e-06 | 2.0862e-06 | 0.99169 | 34.6732 | 0.00016823 | P26
2012-57#145 | 470 | 2012-10-03 15:03 | 24C | 0.98936 | 1.1396e-06 | 2.0862e-06 | 0.98936 | 34.5815 | 4.5217e-05 | P25
2012-57#146 | 631 | 2012-10-03 15:06 | 24C | 0.9895 | 2.4909e-06 | 2.0862e-06 | 0.9895 | 34.5873 | 9.4885e-05 | P26
2012-57#147 | 584 | 2012-10-03 15:09 | 24C | 0.96793 | 3.6104e-06 | 2.0266e-06 | 0.96793 | 33.7414 | 0.00014304 | PA006
2012-57#148 | 643 | 2012-10-03 15:12 | 24C | 0.99031 | 1.373e-05 | 2.1458e-06 | 0.99031 | 34.6188 | 0.00054054 | P26
2012-57#149 | 235 | 2012-10-03 15:15 | 24C | 0.98715 | 2.7426e-06 | 2.1458e-06 | 0.98715 | 34.495 | 0.00010447 | P12
2012-57#150 | 633 | 2012-10-03 15:18 | 24C | 0.9844 | 3.8542e-06 | 2.0266e-06 | 0.98441 | 34.387 | 0.00014997 | P26
2012-57#151 | 205 | 2012-10-03 15:21 | 24C | 0.98348 | 2.4754e-06 | 2.1458e-06 | 0.98348 | 34.3508 | 9.6652e-05 | P12
2012-57#152 | 604 | 2012-10-03 15:24 | 24C | 0.99182 | 3.2149e-06 | 2.0862e-06 | 0.99182 | 34.6782 | 0.00012805 | P26
2012-57#153 | 519 | 2012-10-03 15:27 | 24C | 0.97233 | 3.1414e-06 | 2.2054e-06 | 0.97234 | 33.9138 | 0.00012618 | P26
2012-57#154 | 477 | 2012-10-03 15:30 | 24C | 0.98936 | 2.7268e-06 | 2.1458e-06 | 0.98936 | 34.5817 | 0.00010658 | P25
2012-57#155 | 469 | 2012-10-03 15:33 | 24C | 0.98941 | 2.7262e-06 | 2.0266e-06 | 0.98941 | 34.5835 | 0.00010656 | P25
2012-57#156 | 158 | 2012-10-03 15:35 | 24C | 0.92295 | 2.7426e-06 | 2.0862e-06 | 0.92296 | 31.9885 | 0.0001071 | P10
2012-57#157 | 601 | 2012-10-03 15:38 | 24C | 0.99197 | 3.4657e-06 | 2.0862e-06 | 0.99198 | 34.6844 | 0.00013822 | P26
2012-57#158 | 150 | 2012-10-03 15:41 | 24C | 0.92206 | 2.2394e-06 | 2.0266e-06 | 0.92206 | 31.954 | 8.4447e-05 | P8
2012-57#159 | 585 | 2012-10-03 15:44 | 24C | 0.96497 | 3.2145e-06 | 2.1458e-06 | 0.96497 | 33.6255 | 0.00012387 | PA006
2012-57#160 | 483 | 2012-10-03 15:46 | 24C | 0.98936 | 3.581e-06 | 2.2054e-06 | 0.98936 | 34.5816 | 0.00014324 | P25
2012-57#161 | 223 | 2012-10-03 15:49 | 24C | 0.98941 | 2.2391e-06 | 2.1458e-06 | 0.98942 | 34.5838 | 8.8738e-05 | P12
2012-57#162 | 525 | 2012-10-03 15:52 | 24C | 0.94292 | 2.0804e-06 | 2.0862e-06 | 0.94292 | 32.765 | 7.9109e-05 | P26
2012-57#163 | 2 | 2012-10-03 15:54 | 24C | 0.98966 | 2.7425e-06 | 2.0862e-06 | 0.98966 | 34.5934 | 0.00010456 | DWR
2012-57#164 | 587 | 2012-10-03 15:57 | 24C | 0.94308 | 2.407e-06 | 2.0862e-06 | 0.94309 | 32.7713 | 9.3191e-05 | PA006
2012-57#165 | 204 | 2012-10-03 16:00 | 24C | 0.98591 | 2.7757e-06 | 2.1458e-06 | 0.98591 | 34.4461 | 0.00010944 | P12
2012-57#166 | 475 | 2012-10-03 16:03 | 24C | 0.98954 | 3.3587e-06 | 2.0266e-06 | 0.98954 | 34.5887 | 0.00012795 | P25
2012-57#167 | 490 | 2012-10-03 16:06 | 24C | 0.98938 | 2.2405e-06 | 2.1458e-06 | 0.98938 | 34.5824 | 9.0443e-05 | P25
2012-57#168 | 517 | 2012-10-03 16:08 | 24C | 0.97962 | 2.4909e-06 | 2.0862e-06 | 0.97962 | 34.1994 | 9.7462e-05 | P26
2012-57#169 | 476 | 2012-10-03 16:11 | 24C | 0.98941 | 2.4912e-06 | 2.1458e-06 | 0.98941 | 34.5837 | 9.9543e-05 | P25
2012-57#170 | 603 | 2012-10-03 16:13 | 24C | 0.99192 | 1.3712e-06 | 2.0862e-06 | 0.99192 | 34.6823 | 5.4326e-05 | P26
2012-57#171 | 482 | 2012-10-03 16:16 | 24C | 0.98936 | 2.2259e-06 | 2.0266e-06 | 0.98936 | 34.5815 | 8.7866e-05 | P25
Calibration#1 | K15 | 2012-10-04 10:11 | 24C | 0.99979 | 9.3671e-06 | 4.9472e-06 | 0.99979
Calibration#2 | K15 | 2012-10-04 10:12 | 24C | 0.99979 | 2.7426e-06 | -1.967e-06 | 0.99979
2012-13#1 | std | 2012-10-04 10:13 | 24C | 0.9998 | 2.4784e-06 | -1.9073e-06 | 0.99979 | 34.9919 | 9.8098e-05
2012-13#2 | 10 | 2012-10-04 10:19 | 24C | 0.98971 | 3.595e-06 | -1.9073e-06 | 0.98971 | 34.5953 | 0.00014165 | DWR
2012-13#3 | 5 | 2012-10-04 10:56 | 24C | 0.98979 | 2.8795e-06 | -1.9073e-06 | 0.98979 | 34.5986 | 0.00011403 | DWR
2012-13#4 | 468 | 2012-10-04 10:59 | 24C | 0.9373 | 2.8792e-06 | -1.967e-06 | 0.93729 | 32.5458 | 0.00011385 | P24
2012-13#5 | 371 | 2012-10-04 11:02 | 24C | 0.99165 | 2.8904e-06 | -2.0266e-06 | 0.99165 | 34.6716 | 0.00011471 | P20
2012-13#6 | 419 | 2012-10-04 11:05 | 24C | 0.98707 | 2.2262e-06 | -1.9073e-06 | 0.98707 | 34.4917 | 8.7876e-05 | P20
2012-13#7 | 301 | 2012-10-04 11:08 | 24C | 0.9861 | 1.3713e-06 | -1.9073e-06 | 0.9861 | 34.4535 | 5.2235e-05 | P16
2012-13#8 | 296 | 2012-10-04 11:10 | 24C | 0.99128 | 3.3456e-06 | -1.9073e-06 | 0.99128 | 34.6571 | 0.00012993 | P16
2012-13#9 | 377 | 2012-10-04 11:13 | 24C | 0.98611 | 3.4389e-06 | -1.9073e-06 | 0.98611 | 34.4538 | 0.00013908 | P20
2012-13#10 | 376 | 2012-10-04 11:15 | 24C | 0.98741 | 3.1415e-06 | -2.0266e-06 | 0.98741 | 34.5051 | 0.00012104 | P20
2012-13#11 | 600 | 2012-10-04 11:18 | 24C | 0.99068 | 2.8769e-06 | -2.0266e-06 | 0.99067 | 34.6333 | 0.00011351 | P26
2012-13#12 | 372 | 2012-10-04 11:20 | 24C | 0.99131 | 2.2397e-06 | -2.0266e-06 | 0.99131 | 34.6581 | 8.8728e-05 | P20
2012-13#13 | 599 | 2012-10-04 11:23 | 24C | 0.99123 | 3.3587e-06 | -1.967e-06 | 0.99123 | 34.6551 | 0.00013098 | P26
2012-13#14 | 649 | 2012-10-04 11:25 | 24C | 0.98607 | 2.7261e-06 | -1.967e-06 | 0.98606 | 34.4521 | 0.00010656 | P26
2012-13#15 | 375 | 2012-10-04 11:28 | 24C | 0.98951 | 2.2391e-06 | -1.967e-06 | 0.9895 | 34.5873 | 8.959e-05 | P20
2012-13#16 | 378 | 2012-10-04 11:30 | 24C | 0.98413 | 2.4615e-06 | -1.9073e-06 | 0.98413 | 34.3763 | 9.4938e-05 | P20
2012-13#17 | 533 | 2012-10-04 11:32 | 24C | 0.98702 | 2.4915e-06 | -1.9073e-06 | 0.98702 | 34.4896 | 9.6976e-05 | P26
2012-13#18 | LOOP 1 | 2012-10-04 11:34 | 24C | 0.93664 | 2.2259e-06 | -1.9073e-06 | 0.93664 | 32.5203 | 8.7876e-05
2012-13#19 | 378 | 2012-10-04 11:37 | 24C | 0.98411 | 2.4918e-06 | -1.9073e-06 | 0.98411 | 34.3753 | 9.7029e-05 | P20
2012-13#20 | 297 | 2012-10-04 11:41 | 24C | 0.99141 | 2.2397e-06 | -1.9073e-06 | 0.99141 | 34.6622 | 9.0433e-05 | P16
2012-13#21 | 450 | 2012-10-04 11:44 | 24C | 0.93757 | 2.2394e-06 | -1.967e-06 | 0.93757 | 32.5567 | 8.7866e-05 | P22
2012-13#22 | 660 | 2012-10-04 11:46 | 24C | 0.98812 | 2.2391e-06 | -1.967e-06 | 0.98812 | 34.5328 | 8.7067e-05 | P16
2012-13#23 | LOOP 3 | 2012-10-04 11:52 | 24C | 0.93648 | 2.1278e-06 | -2.0266e-06 | 0.93648 | 32.5141 | 8.0956e-05
2012-13#24 | 597 | 2012-10-04 11:55 | 24C | 0.99192 | 2.7425e-06 | -1.9073e-06 | 0.99191 | 34.6819 | 0.00010761 | P26
2012-13#25 | 532 | 2012-10-04 11:58 | 24C | 0.98895 | 1.4046e-06 | -1.9073e-06 | 0.98895 | 34.5655 | 5.5371e-05 | P26
2012-13#26 | 495 | 2012-10-04 12:00 | 24C | 0.98975 | 3.8436e-06 | -1.9073e-06 | 0.98975 | 34.5969 | 0.00015239 | P26
2012-13#27 | 418 | 2012-10-04 12:03 | 24C | 0.989 | 1.4046e-06 | -1.967e-06 | 0.989 | 34.5673 | 5.2235e-05 | P20
2012-13#28 | 674 | 2012-10-04 12:06 | 24C | 0.98453 | 4.6217e-06 | -2.0266e-06 | 0.98453 | 34.3919 | 0.00018033 | P4
2012-13#29 | 626 | 2012-10-04 12:08 | 24C | 0.9779 | 1.4046e-06 | -1.9073e-06 | 0.97789 | 34.1317 | 5.4326e-05 | P26
2012-13#30 | 373 | 2012-10-04 12:11 | 24C | 0.99128 | 1.3713e-06 | -1.9073e-06 | 0.99128 | 34.6571 | 5.4326e-05 | P20
2012-13#31 | 629 | 2012-10-04 12:14 | 24C | 0.9379 | 2.4909e-06 | -2.0266e-06 | 0.9379 | 32.5693 | 9.4066e-05 | P26
2012-13#32 | LOOP 5 | 2012-10-04 12:16 | 24C | 0.92205 | 1.1529e-06 | -2.0862e-06 | 0.92205 | 31.9535 | 4.5861e-05
2012-13#33 | 674 | 2012-10-04 12:19 | 24C | 0.98456 | 2.4908e-06 | -2.0266e-06 | 0.98456 | 34.393 | 9.869e-05 | P4
2012-13#34 | 648 | 2012-10-04 14:47 | 24C | 0.98709 | 3.4659e-06 | -2.0862e-06 | 0.98709 | 34.4924 | 0.00013824 | P26
2012-13#35 | 675 | 2012-10-04 14:50 | 24C | 0.98247 | 1.3882e-06 | -1.9073e-06 | 0.98246 | 34.3109 | 5.2235e-05 | P4
2012-13#36 | 675 | 2012-10-04 14:52 | 24C | 0.98253 | 1.3713e-06 | -2.0266e-06 | 0.98253 | 34.3133 | 5.5428e-05 | P4
2012-13#37 | 302 | 2012-10-04 14:55 | 24C | 0.98419 | 1.1532e-06 | -2.0266e-06 | 0.98418 | 34.3784 | 4.2703e-05 | P16
2012-13#38 | 340 | 2012-10-04 14:57 | 24C | 0.98885 | 2.2401e-06 | -1.967e-06 | 0.98885 | 34.5616 | 9.0443e-05 | P16
2012-13#39 | 374 | 2012-10-04 15:00 | 24C | 0.99062 | 2.4921e-06 | -1.9073e-06 | 0.99062 | 34.6311 | 9.7029e-05 | P20
2012-13#40 | 661 | 2012-10-04 15:03 | 24C | 0.98647 | 3.6235e-06 | -2.0266e-06 | 0.98647 | 34.4681 | 0.00013998 | P16
2012-13#41 | 358 | 2012-10-04 15:05 | 24C | 0.9898 | 2.2127e-06 | -1.9073e-06 | 0.98979 | 34.5987 | 8.363e-05 | P20
2012-13#42 | 373 | 2012-10-04 15:08 | 24C | 0.99129 | 2.7103e-06 | -1.9073e-06 | 0.99129 | 34.6573 | 0.00010447 | P20
2012-13#43 | 676 | 2012-10-04 15:10 | 24C | 0.97909 | 2.7261e-06 | -1.9073e-06 | 0.97909 | 34.1784 | 0.00010658 | P4
2012-13#44 | 4 | 2012-10-04 15:13 | 24C | 0.98977 | 2.5747e-06 | -2.0266e-06 | 0.98977 | 34.5977 | 0.00010258 | DWR
2012-13#45 | 673 | 2012-10-04 15:16 | 24C | 0.98658 | 2.2394e-06 | -2.0266e-06 | 0.98657 | 34.4722 | 8.958e-05 | P4
2012-13#46 | 302 | 2012-10-04 15:18 | 24C | 0.9844 | 1.1199e-06 | -1.967e-06 | 0.9844 | 34.3868 | 4.4374e-05 | P16
2012-13#47 | 359 | 2012-10-04 15:21 | 24C | 0.98934 | 1.8095e-05 | -1.967e-06 | 0.98933 | 34.5806 | 0.00071038 | P20
2012-13#48 | 298 | 2012-10-04 15:25 | 24C | 0.99052 | 3.6239e-06 | -2.0266e-06 | 0.99052 | 34.6272 | 0.00013993 | P16
2012-13#49 | 355 | 2012-10-04 15:27 | 24C | 0.93719 | 2.8897e-06 | -1.967e-06 | 0.93719 | 32.5419 | 0.00011404 | P18
2012-13#50 | 623 | 2012-10-04 15:30 | 24C | 0.98439 | 2.4908e-06 | -2.0266e-06 | 0.98439 | 34.3863 | 9.9734e-05 | P26
2012-13#51 | 299 | 2012-10-04 15:33 | 24C | 0.98948 | 3.1414e-06 | -1.9073e-06 | 0.98948 | 34.5863 | 0.00012309 | P16
2012-13#52 | 341 | 2012-10-04 15:35 | 24C | 0.98692 | 2.4915e-06 | -2.0266e-06 | 0.98691 | 34.4856 | 9.4938e-05 | P16
2012-13#53 | 300 | 2012-10-04 15:38 | 24C | 0.98757 | 2.7425e-06 | -1.9073e-06 | 0.98757 | 34.5114 | 0.00010658 | P16
2012-13#54 | 598 | 2012-10-04 15:40 | 24C | 0.99167 | 2.7426e-06 | -2.0266e-06 | 0.99167 | 34.6722 | 0.00010447 | P26
2012-13#55 | 620 | 2012-10-04 15:42 | 24C | 0.98952 | 3.3588e-06 | -1.967e-06 | 0.98952 | 34.5878 | 0.00012994 | P26
2012-13#56 | 295 | 2012-10-04 15:45 | 24C | 0.99156 | 2.4908e-06 | -2.0266e-06 | 0.99156 | 34.6681 | 9.9734e-05 | P16
2012-13#57 | 297 | 2012-10-04 15:47 | 24C | 0.99127 | 4.3338e-06 | -2.0266e-06 | 0.99126 | 34.6564 | 0.00016509 | P16
2012-13#58 | 496 | 2012-10-04 15:50 | 24C | 0.98938 | 3.6104e-06 | -2.0266e-06 | 0.98938 | 34.5825 | 0.00014188 | P26
2012-13#59 | LOOP 2 | 2012-10-04 15:52 | 24C | 0.93681 | 2.8794e-06 | -1.967e-06 | 0.9368 | 32.5267 | 0.0001132
2012-13#60 | 10 | 2012-10-04 15:55 | 24C | 0.98973 | 3.466e-06 | -2.0266e-06 | 0.98973 | 34.5961 | 0.00013386 | DWR
2012-59#66 | 13 | 2012-11-07 14:15 | 24C | 0.8628 | 0 | -9.5367e-07 | 0.8628 | 29.6665 | 1.0662e-06 | P1
2012-59#88 | 20 | 2012-11-07 15:17 | 24C | 0.90711 | 3.1999e-06 | -1.0133e-06 | 0.90711 | 31.3746 | 0.00012329 | P2
2012-59#80 | 60 | 2012-11-07 14:57 | 24C | 0.98685 | 2.4908e-06 | -1.0133e-06 | 0.98685 | 34.4829 | 9.2996e-05 | P4
2012-59#89 | 61 | 2012-11-07 15:20 | 24C | 0.98575 | 3.4551e-06 | -1.0133e-06 | 0.98575 | 34.4399 | 0.00013929 | P4
2012-59#77 | 69 | 2012-11-07 14:45 | 24C | 0.98651 | 2.4909e-06 | -9.5367e-07 | 0.98651 | 34.4695 | 9.869e-05 | P4
2012-59#79 | 69 | 2012-11-07 14:54 | 24C | 0.9865 | 3.2037e-06 | -1.0133e-06 | 0.9865 | 34.4691 | 0.00012803 | P4
2012-59#65 | 71 | 2012-11-07 14:12 | 24C | 0.98448 | 2.7103e-06 | -1.0133e-06 | 0.98448 | 34.39 | 0.00010447 | P4 2012-13
2012-59#81 | 71 | 2012-11-07 14:59 | 24C | 0.98436 | 2.0841e-06 | -9.5367e-07 | 0.98435 | 34.385 | 8.3896e-05 | P4
2012-59#76 | 72 | 2012-11-07 14:43 | 24C | 0.98237 | 2.8897e-06 | -9.5367e-07 | 0.98237 | 34.3072 | 0.00011213 | P4
2012-59#87 | 90 | 2012-11-07 15:15 | 24C | 0.98719 | 3.1415e-06 | -9.5367e-07 | 0.98719 | 34.4965 | 0.00012514 | P4
2012-59#67 | 93 | 2012-11-07 14:18 | 24C | 0.98471 | 3.6104e-06 | -9.5367e-07 | 0.98471 | 34.3991 | 0.00014155 | P4
2012-59#84 | 96 | 2012-11-07 15:07 | 24C | 0.97728 | 2.4915e-06 | -1.0133e-06 | 0.97728 | 34.1076 | 9.8725e-05 | P4
2012-59#70 | 99 | 2012-11-07 14:27 | 24C | 0.92252 | 3.4551e-06 | -9.5367e-07 | 0.92252 | 31.9715 | 0.00013277 | P4
2012-59#86 | 102 | 2012-11-07 15:12 | 24C | 0.98637 | 3.6104e-06 | -1.0133e-06 | 0.98637 | 34.4641 | 0.00014493 | P4
2012-59#68 | 103 | 2012-11-07 14:21 | 24C | 0.98452 | 9.1433e-06 | -9.5367e-07 | 0.98452 | 34.3915 | 0.0003615 | P4
2012-59#69 | 128 | 2012-11-07 14:24 | 24C | 0.92255 | 2.2391e-06 | -9.5367e-07 | 0.92255 | 31.9728 | 8.7006e-05 | P6
2012-59#75 | 132 | 2012-11-07 14:40 | 24C | 0.98954 | 2.7426e-06 | -9.5367e-07 | 0.98954 | 34.5886 | 0.00010865 | P8
2012-59#83 | 132 | 2012-11-07 15:04 | 24C | 0.98949 | 2.4908e-06 | -9.5367e-07 | 0.98949 | 34.5867 | 9.9745e-05 | P8
2012-59#71 | 133 | 2012-11-07 14:30 | 24C | 0.98762 | 3.6109e-06 | -1.0729e-06 | 0.98762 | 34.5133 | 0.00014159 | P8
2012-59#74 | 134 | 2012-11-07 14:38 | 24C | 0.9864 | 2.7426e-06 | -1.0133e-06 | 0.9864 | 34.4652 | 0.00010451 | P8
2012-59#73 | DWR | 2012-11-07 14:36 | 24C | 0.99014 | 6.664e-08 | -1.0729e-06 | 0.99014 | 34.6122 | 0 | 25 2012-12
2012-59#72 | LOOP | 2012-11-07 14:32 | 24C | 0.8658 | 3.2248e-06 | -1.0133e-06 | 0.8658 | 29.7818 | 0.00012219 | JF3
2012-59#78 | LOOP | 2012-11-07 14:51 | 24C | 0.8374 | 1.1199e-06 | -1.0133e-06 | 0.8374 | 28.694 | 4.265e-05 | JF1
2012-59#82 | LOOP | 2012-11-07 15:02 | 24C | 0.86385 | 3.2148e-06 | -1.0729e-06 | 0.86385 | 29.7071 | 0.00012584 | JF4
2012-59#85 | LOOP | 2012-11-07 15:09 | 24C | 0.85169 | 2.4909e-06 | -1.0133e-06 | 0.85169 | 29.2408 | 9.6511e-05 | JF2