## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2025 | Updated channel names & formats in TOB files. GG |
| 1 Dec 2020 | Added HPLC data. S.H. |

## PROCESSING NOTES

Cruise: 2012-14

Agency: OSD

Location: Strait of Georgia

Project: Venus/Strait of Georgia

Party Chief: Macoun P.

Platform: John P. Tully

Date: February 22, 2012 – February 27, 2012

Processed by: Germaine Gatien

Date of Processing: 3 April 2012 – 25 April 2012

Number of original HEX files: 22

Number of CTD files:  22

Number of bottle casts:	22

Number of original TSG files: 3

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#0997), a Wet Labs Eco-AFL/FL Fluorometer (#2215), a pumped SeaPoint fluorometer with a 10X cable (#2228) and an altimeter (#1204). Data logging computer #3 was used.

A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #0603 and a flow meter.

The data logging computer was #3.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, plus details about the rosette. There was no information about the TSG. There was a note warning that the fluorometer configuration parameters used at sea was incorrect.

The rosette logs were in good order.

Both a SeaPoint and a WetLabs ECO fluorometer were used for this cruise. The comparison between the two fluorometers and extracted chlorophyll samples shows that both sensors read too high when chlorophyll is very low. Both sensors grow closer to the extracted CHL values as CHL approaches 1ug/L. The ECO fluorometer read about 50% higher than the SeaPoint. The range of CHL was very limited so further inter-comparisons are needed.

A change to the sampling rate of the WetLabs ECO fluorometer has led to better detail in the fluorometry than was seen during 2011 cruises.

There was no salinity calibration sampling, but during the previous cruise the primary sensors provided values very close to bottle samples.

There was no calibration sampling for Dissolved Oxygen so the calibration parameters found for 2012-01 were used. No estimate can be made of the accuracy, though near-surface saturation values are mostly ~95%, which is reasonable for the time of year and conditions.

The recalibration of the TSG salinity is based on comparisons with CTD data from this cruise and an earlier one. The other cruise had very noisy TSG salinity and the salinity gradients were fairly high for this cruise, so only a rough estimate could be made.

The thermosalinograph fluorescence data are raw, with volts as units. The TSG fluorometer was cleaned prior to this cruise and a pre-cruise calibration was done to obtain a blank voltage. However, a problem was found with the recalibration during a previous cruise, so offset = 0 and scale factor = 1 was used.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as sampling notes summarizing problems and points of interest with reference to processing.

Both a SeaPoint and a WetLabs ECO fluorometer were used for this cruise. It was noted at sea that the ECO readings were much lower than the SeaPoint; this turned out to be due to an error in the configuration file. The TSG technician thought the range had been changed, but later discovered that an extra step was needed to do this. So while a scale factor of 6 (range 0-30mg/L) was intended, 25 was actually used (range 0-125ug/L).  The corrected configuration file produces more reasonable readings. There was also a change to the sampling rate of the ECO fluorometer producing more detail than when this type of sensor was used in 2011.

Extracted chlorophyll and nutrient data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. The pressure and conductivity sensors were recalibrated in March and April 2011 and the dissolved oxygen sensor in December 2010. While they have been used on many cruises since then, the calibration information is poor for most of the cruises due to the CTD data being limited to 1Hz, delays in salinity analysis and scatter in comparisons. Fortunately 2012-01 which was run immediately before this cruise provided useful information.

The calibration constants were checked for all instruments. The only errors found were in the ECO fluorometer parameters (as mentioned above), and in the serial number of the SBE DO sensor. That was corrected and the file was saved as 2012-14-ctd.xmlcon.

##### Initial Rosette File Conversion and DO Calibration Study

There was no DO calibration sampling for this cruise, but there was good DO sampling during the cruise which immediately preceded it, 2012-01. The calibration study included an analysis of the hysteresis parameter, E. The parameters found for 2012-01 were entered into the DO section of the configuration file:

E=0.3

Soc = 0.4801

Offset=-0.4816

The file was saved as 2012-14-ctd-new.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new configuration parameters. They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files and none looked out of line.

A preliminary header check shows that the SeaPoint fluorescence may have gone off-scale during cast #1.

An addsamp file was prepared and then converted to CST files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2012-14-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2012-14chl.xls. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2012-14chl.csv which was then converted to individual CHL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2012-14nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers and saved as 2012-14-nuts.csv. The file was converted to individual NUT files.

The CHL and NUT files were merged with CST files in 2 steps.

After the 2nd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

The data from the SAMAVG files were exported to a spreadsheet 2012-14-fluorometer-study.xls.

- The ratio of each fluorometer to extracted CHL was plotted against extracted CHL. Both had a high ratio at very low CHL but the ratio fell as CHL rose. For the ECO the ratio started at ~3.5 and fell to ~1.2 when CHL → 1ug/L. For the SeaPoint the ratio started at 2.3 and fell to ~0.9 when CHL → 1ug/L. There was only one CHL value >1ug/L, at 12.1ug/L and for that case the ratio was 1.7 for the ECO and 1.0 for the SeaPoint. So the ECO appears to be giving values that are too high in all cases. The SeaPoint fluorescence is too high at low CHL but not by as much as the ECO and is a little low for CHL ~1ug/L.

- A fit of fluorescence against CHL shows a change in slope at about 0.5 to 0.6ug/L, with a steeper slope below that level. The two fits were:

ECO = 1.097*Extracted CHL +0.3727

SeaPoint = 0.768*Extracted CHL +0.2685

The offsets are larger than the dark values which are ~0.1ug/L for both fluorometers.

- A fit of one fluorometer against the other looks fairly tight

ECO FL = 1.45 * SeaPoint - 0.024

For the one high CHL values the ECO fluorescence is ~1.7 times the SeaPoint.

COMPARE was run comparing the two fluorometer channels with extracted CHL. This shows that the differences between both fluorometers and CHL show no pressure dependence, except that there was a much larger difference for the single surface sample with high CHL. The comparison file was saved as 2012-14-fl-chl-comp1.xls.

The range of extracted CHL is small with the exception of 1 value, so the conclusions may not apply to other cruises. More information is needed about how these 2 types of fluorometer compare. A trial conversion of ECO fluorescence using a scale factor of 20 instead of 25 did produce values more in line with the SeaPoint, but this study was not pursued because of the limited range of CHL.

##### Conversion of Full Files from Raw Data

All files were converted using 2012-14-ctd-new.con.

A few casts were examined and all expected channels are present.

The descent rate was fairly steady; for a few casts the descent rate was low in the top 50m. This was noted in the log book and was due to the winch operator believing there was not enough weight on the rosette.

The two temperature and conductivity channels are close during the downcasts but there are larger differences of variable sign during upcasts. This is likely something to do with how the CTD is mounted.

Altimetry looks useful when the CTD got near the bottom, dissolved oxygen, ECO and SeaPoint fluorescence and transmissivity look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 2 casts to determine the offset between the DO channel and the primary temperature. It is very hard to judge because the upcast temperature is noisy and there are a lot of stops for bottles, but a setting of 4s looks best.

ALIGNCTD was used to advance the DO Voltage by 4s relative to the pressure.

##### CELLTM

Tests were run using a variety of settings on 2 casts to see which brought upcast and downcast data closest in T-S space. However, there was no clear “best choice” with some features even better with no adjustment. During 2012-01 there were similar problems, but some settings clearly improved the data, so the same settings will be used for this cruise.

CELLTM was run using (α = 0.0245, β=9.5) for the primary and (α = 0.03, β=9) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

There are no deep casts, but one casts was examined to check for agreement between the pairs of T and C sensors. One cast from 2012-01 is show for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2012-01-0050 | 800 1400 2500 3500 4200 | +0.0005 +0.0007 +0.0014 +0.0020 +0.0021 | -0.00010 -0.00011 -0.00014 -0.00016 -0.00015 | -0.0016 -0.0022 -0.0034 -0.0041 -0.0044 | High, XX Noisy |
| 2012-14-0019 | 400 | -0.0001 | +0.00005 | +0.0006 | Mod, F. Steady |

It was noted during 2012-01 that there was some pressure-dependence in temperature, conductivity and salinity differences. While there are possible explanations for that because slight misalignment can affect the differences in high gradient zones, but the comparison with bottles confirmed some pressure-dependence in the secondary salinity.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are some small negative values in the ECO fluorescence, as usual. The maximum SeaPoint fluorescence value occurred during cast #1 is 14.872. A plot shows no evidence that the sensor was off-scale during cast #1.

Surface Check was run and shows an average surface pressure for the cruise was 3.0db which looks reasonable for the area and conditions.

The Cross-Reference Check was compared with the log book and no problems were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and MRG files were exported to spreadsheets. All but one of the CLN files got within 15m of the bottom and the header values look reasonable. Only 1 of the bottle files got within 15m of the bottom and that header is correct. However, there were 8 other casts which did not sample close to the bottom but have header readings; those header entries were removed. The spreadsheet was recreated to ensure all corrections had been made. CLEAN was rerun on the MRG files.

The bottom depths are close to those listed in the log book, though there were some differences for the last 5 casts, perhaps the log was updated after the header was entered; erasures suggest that was the case. The headers were adjusted to match the log.

##### 13. Shift

Fluorescence

Tests were run on two casts to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence traces look like that of the temperature trace. This task was complicated by:

- Noisy upcast temperature data so the vertical offset is confused.

- Many stops for bottles on upcasts so that fluorescence gets opportunities to “catch up”

- Very low fluorescence values so there is little variability to aid estimate of vertical offset.

For the ECO fluorometer, plots suggest that a shift by either +24 or +48 records improves the data. The optimal choice is likely +48, but at times that looks like it might be too high. The value used on other recent uses of the ECO sensor was +48 records. For 2012-01 it was also hard to make a judgment for similar reasons. In the absence of a clear result from the tests, +48 records will be applied. The fluorescence is so low that the difference is unlikely to be significant.

SHIFT was run on all casts to advance the ECO fluorescence channel by +48 records.

For the SeaPoint fluorometer the usual +24 record setting looks reasonable.

Conductivity

Tests were run on the primary conductivity channel only, since there are questions about the secondary salinity calibration (based on the 2012-01 bottle comparison.) The results were plotted for 4 casts using a variety of shifts and the best results were with an advance of -0.4 records.

SHIFT was run on all casts using that setting.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. For some casts it looks like the channel had been shifted too much, and for others not enough, but overall, the alignment looks ok. No further shift will be applied.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The conductivity and temperature sensors were recalibrated in late March 2011 and were used for 7 previous cruises, for 6 of which the data were averaged on acquisition. The secondary sensors were selected for 6 cruises due to more noise in the primary, but for the most recent the primary looked better. The comparisons for 1 cruise were based on only 5 bottles and another had bottles that had been stored for 4 months before analysis, so the comparison was noisy and not trusted. The best 2 comparisons showed the primary to be low by 0.0011 and 0.0005 and the secondary to be low by 0.0005 and 0.0003. For cruise 2012-01 the usual 24Hz acquisition was used and the primary salinity was selected for archiving. The primary was found to be low by 0.0004 and the secondary was low by 0.0028 with some pressure-dependence.

2. Dissolved Oxygen

The DO sensor was recalibrated in Dec.2010. It was used for 2011-60, 2011-25*, 2011-63, 2011-10, 2011-11 and 2011-76. As mentioned above the CTD data were averaged over 24 scans on acquisition. Most bottle comparisons were noisy and/or had few bottles. For 2012-01 there was a good comparison and hysteresis tests were run.

3. Pressure

The sensor was recalibrated in April 2011 and was used for was used for 2011-60, 2011-25*, 2011-63, 2011-10, 2011-11, 2011-76 and 2012-01. No further offset was applied to any of those cruises.

*Note: Cruise 2011-25 has not been processed yet.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity values below 250db were a little below the historic minimum except for the casts at the northern end of the cruise near Texada Island. Temperatures were a little low at mid-depths to the north and at the bottom towards the south. These do not look like evidence of instrument or calibration problems since the historic ranges are overly severe for near-shore casts. Similar results were found in the same area during 2011-11 in November 2011.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

The bottle comparison from 2012-01 suggested that the secondary conductivity sensor is showing some pressure-dependence, so the primary channels were chosen for archiving. Before editing all files were copied to *.EDT.

CTDEDIT was used to clean salinity where unstable features appear to be due to mismatch of conductivity and temperature and likely to affect the bin-averaged values, and to remove records corrupted by shed wakes mostly near the surface and bottom of the casts.

All edited files were copied to *.EDT.

##### Initial Recalibration

The pressure looks ok.

In the absence of calibration sampling, Salinity and SBE Dissolved Oxygen will not be recalibrated since they were found to require no correction during 2012-01.

The ECO fluorescence channel will be adjusted by adding 0.031ug/L to avoid negative numbers.

CALIBRATE was run using file 2012-14-recal1.ccf to apply the offset to the Fluorescence channel.

The SAM and MRGCLN2 files were also recalibrated.

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering both fluorescence channels before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the 2 fluorescence channels in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity and fluorescence data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

There was no Dissolved Oxygen sampling, so no estimate can be made of

the accuracy of the Oxygen:Dissolved:SBE data cannot be made.

Calibration was based on the results of cruise 2012-01.

For details on the processing see processing report: 2012-14-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

Another special file was prepared for Dr. Peña containing data from the top 50db of all CTD casts that included rosette sampling (all casts for this cruise.) The data were clipped to 50db, sigma-t was derived, and the data were exported to file 2012-14-SOG.csv.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged from 89% to 98% with most ~95%.

##### 24. Final Bottle Files

The MRGCOR` files were put through SORT to order on increasing pressure.

REMOVE was on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

A header check was run on the final files and no problems were found.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. No problems were found. The file was saved as 2012-014-bottles-final.csv.

Standards check was run on all files and no problems were found.

##### 25. Thermosalinograph Data

Data were provided in 3 hex files.

There were no loop samples.

a.) Checking calibrations

The calibrations were checked and the only problem concerns the fluorometry. The configuration file contains values from a 2012 calibration that was found to have been done wrong, producing many negative values. For 2012-01 it was decided that given there had been a major cleaning of the sensor and no recalibration after that, the best procedure would be to use a blank output of 0 and scale factor of 1 in order to obtain raw values in volts. This calibration file was saved as 2012-14-tsg.xmlcon.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

A check was made of fluorescence and the maximum value was ~2.5V.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Track plots and files were examined to determine if the problem of positions getting stuck occurred during this cruise. The judgment is a little difficult in that there were many stations run in quiet waters, so that there was no motion. A short section at the end of file #3 would appear to be at the dock or at SI. Similarly, a section at the end of file #2 occurred during a long stop in Saanich Inlet while repairs were done and diving operations were underway. So there is likely no problem with the positions; if they did become stuck it is of no significance. There is a gap in the tracks between the end of file #2 and the beginning of file #3, but there is a time gap of about an hour and a half, so this makes sense.

Time-series plots were produced. There is a sometimes noise in the salinity and in the intake temperature, but the ship was frequently in waters with high variability. The temperature and fluorescence look reasonable. The flow rate dropped a little around 0800 on the 23rd and then remained steady until the end of file #3 when it dropped to 0.

c.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2012-14-ctd-tsg-comp.xls.

All files overlapped at least 1 CTD cast so the files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files. The TSG data were added to 2012-14-ctd-tsg-comp.xls. There were 22 matches.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude were all <0.00023° and in longitude <0.00041°. The median differences were 0.00000° and 0.00004°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected throughout the cruise. The differences between the two temperatures had a median difference of 0.354, 0.332 and 0.307Cº  for the 3 files but the standard deviations were 0.09, 0.10 and 0.24 Cº;, respectively. During a quiet section of file #3 the median difference was 0.291Cº; with a standard deviation of 0.040Cº. During 2012-01 it was about 0.28Cº.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. During stops the intake and lab temperatures differed by a median of 0.032Cº. Excluding a few cases with high standard deviation in the temperature did not change this value significantly.

When all data were included the TSG intake temperature was higher than the CTD by a median value of 0.003 and when cases with high standard deviation in the TSG temperature are excluded it was high by 0.004Cº. This is much smaller than the difference found during 2012-01 when the data were very noisy.

The TSG salinity is lower than that of the CTD by a median value of 0.011 but the standard deviation is 0.41. When cases with a high standard deviation in the TSG salinity were excluded, the TSG was lower than the CTD by 0.007. Even with those exclusions the standard deviation is 0.04 and results range from the TSG being low by 0.133 to being high by 0.051. We can conclude that the TSG is likely reading low, but not by how much.

The comparison of the TSG raw fluorescence to CTD fluorescence shows a similar pattern for both fluorometers.

CTD Wetlabs ECO FL = 9.2486 * TSG(Volts) - 0.6217

CTD Seapoint FL         = 8.0291 * TSG(Volts) - 0.6192

While the two offsets are remarkably close, applying this fit produces values that are much too high. The range is so narrow that the fit is likely meaningless.

To do a few more checks against CHL, the near-surface rosette values were extracted and combined with the TSG fluorescence. The times were matched but we only have start times for the casts, so there will be significant errors due to the difference in time.

The TSG raw fluorescence has a median value of ~0.3 times the rosette bottle values when 1 bottle is excluded, with a range of 0.2 to 0.5. The ratio is ~0.5 when CHL is about 0.3ug/L and ~0.20 when CHL is about 1ug/L.  This is similar to the 2012-01 results.

(See 2012-14-ctd-tsg-comp.xls.)

- Calibration History

The TSG primary temperature and conductivity were recalibrated in March 2011 and the only other use since then was for 2012-01 in February 2012. For that cruise the data were very noisy, especially salinity, and a correction of +0.2 was applied to the TSG salinity, but it was considered a very rough estimate.

Conclusions

1. The TSG clock appears to have worked well.

2. The flow rate was fairly steady.

3. The temperature in the loop increases by a median value of ~0.32Cº during stops for CTD casts and ~0.30Cº in a section where the differences are very quiet. The values are reasonable for winter.

4. The TSG intake temperature is closer to the CTD than for 2012-14, and is as close as we can expect.

5. The TSG Salinity is noisy at times, but this looks like real variations, not instrumental noise. The comparison with CTD casts shows the salinity to be low by a median of 0.007, but the range of values is large. The TSG salinity is clearly a little low, but by how much? During 2012-01 a value of 0.02 was used for recalibration and in that case there were loop samples, but there was also worse noise in the salinity than found in this cruise. For 2012-01 the use of +0.02 as a correction looked like an underestimate based on the evidence, and for this cruise it looks like an overestimate. An addition of 0.02 will be applied to salinity, but if we obtain clearer evidence from future cruises we can fine-tune the recalibration later.

6. The fluorescence data have been converted with Vblank = 0 and Scale Factor=1, so the units are volts and the data are nominal. The pre-cruise measurement of Vblank as 0.168 appears to be wrong. The range of chlorophyll measured is too small to try to calibrate the sensor. It is recommended that this sensor be serviced at the factory.

f.) Editing

The ATC files were copied to *.EDT.

The ATC files were opened in CTDEDIT. The only editing required was the removal of records from the end of file #3 because the flow was very low or turned off during the last 82 minutes. The corresponding time and position data were left in the file so the track plot is preserved.

The edited files were copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

File 2012-14-tsg-recal1.ccf was prepared to adjust salinity by adding 0.02. A few values were checked to ensure it was applied correctly and it was.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Uploy0 and Flag. Position:New was removed since there were very few 0 values, so it is not very useful.

REORDER was used to place Temperature:Secondary ahead of Temperature:Primary and to rename them as Temperature:Primary and Temperature:Lab. The reorder is to ensure that programs pick the proxy intake temperature preferentially.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars (notes from log book)

General – Lanyard on Niskin 1 too short, impossible to arm.

1. Seapoint fluorometer saturating

7. Bottom shoaling to 114 by end of cast

9. Depth 418 from sounder, 411 from altimeter

10. Unstable water at 10m bottle, large sal, temp sensor offsets.

11. Cast terminated – wire jumped the shieve – no one on board to re-terminate

14 to the end – Transmissometer not cleaned because Kimwipes not available

19. Operator used descent rate 0.5m/s for first 50m because insufficient weight on the rosette.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2710 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2128 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1396DR | 26Jan2012 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 23Apr11 | Factory |  |  |
| Eco-AFL Fluorometer | 2215 | 7feb2012 |  |  |  |
| SeaPoint Fluorometer | 2228 |  |  |  |  |
| Pressure Sensor | 63507 | 12Apr2011 | Factory |  |  |
| Altimeter | 1204 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2012-14

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 26Mar11 | Factory |  |  |
| Conductivity | 2488 | 26Mar11 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18Sep01 | “ |  |  |
| Temperature:Secondary | 603 | 3Mar11 | “ |  |  |