## comparison-1119
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b
10 | 151.357 | 7.7349 | 7.736 | 29.6725 | 29.6734 | 6.8745 | 2.3448 | 4 | 4 | 67 | 6.9793 | 0.82361 | 5.7482 |  | 2.7716e-06 | 0.4394 | -0.2067 | Summary of Soc Voffset including the original values in the factory calibration
51 | 76.286 | 7.8537 | 7.8544 | 29.6144 | 29.616 | 6.8584 | 2.3384 | 3 | 3 | 231 | 6.897 | 0.82085 | 5.6615 |  | 3.9745e-05 |  |  | # of bottles |  | m | b | Soc | Voffset | R2
10 | 11.306 | 7.921 | 7.9153 | 29.4586 | 29.4481 | 6.8548 | 2.3131 | 11 | 11 | 74 | 6.8367 | 0.80973 | 5.5359 |  | 6.7672e-05 |  |  | used in fit | Original factory values | 0.4278 | -0.21856 | 0.4278 | -0.5109
40 | 76.106 | 7.8247 | 7.8258 | 29.5714 | 29.5727 | 6.8649 | 2.3075 | 2 | 2 | 183 | 6.9032 | 0.80728 | 5.5728 |  | 6.9313e-05 |  |  | 276 | all data | 0.3997 | -0.1154 | 0.3997 | -0.28872 | 0.8599
1 | 1.907 | 7.7938 | 7.7936 | 29.5563 | 29.561 | 6.8704 | 2.3601 | 14 | 14 | 14 | 6.8432 | 0.83022 | 5.6814 |  | 8.9801e-05 |  |  | 274 | excl. outliers diff>0.05 | 0.4018 | -0.1204 | 0.4018 | -0.29965 | 0.9229
23 | 51.165 | 7.9216 | 7.9231 | 29.5667 | 29.5676 | 6.8499 | 2.277 | 10 | 10 | 150 | 6.8668 | 0.79372 | 5.4503 |  | 0.00011025 |  |  | 267 | excl. outliers diff>0.02 | 0.4179 | -0.1574 | 0.4179 | -0.37665 | 0.9621
13 | 21.677 | 7.9709 | 7.9714 | 29.4879 | 29.4882 | 6.8457 | 2.2763 | 11 | 11 | 110 | 6.837 | 0.79337 | 5.4242 |  | 0.00011692 |  |  | 260 | excl. outliers diff>0.01 | 0.4313 | -0.1883 | 0.4313 | -0.43659 | 0.9746
4 | 51.209 | 7.8253 | 7.8255 | 29.523 | 29.5248 | 6.867 | 2.3006 | 8 | 8 | 47 | 6.8833 | 0.80407 | 5.5347 |  | 0.00012014 |  |  | 240 | excl. outliers diff>0.006 | 0.4386 | -0.2049 | 0.4386 | -0.46717 | 0.9871
10 | 31.262 | 7.951 | 7.9512 | 29.5253 | 29.5266 | 6.8472 | 2.2677 | 9 | 9 | 72 | 6.8467 | 0.78986 | 5.4079 |  | 0.00014659 |  |  | 234 | excl. outliers diff>0.005 | 0.4376 | -0.2026 | 0.4376 | -0.46298 | 0.9892
40 | 96.981 | 7.8146 | 7.8164 | 29.5823 | 29.5832 | 6.866 | 2.3058 | 1 | 1 | 182 | 6.9227 | 0.80664 | 5.5841 |  | 0.00015106 |  |  | 226 | excl. outliers diff>0.004 | 0.4394 | -0.2067 | 0.4394 | -0.47041 | 0.9909
11 | 151.386 | 7.7374 | 7.7374 | 29.6722 | 29.6746 | 6.8741 | 2.3436 | 4 | 4 | 79 | 6.979 | 0.82291 | 5.7431 |  | 0.00016295 |  | range reduced (DO<6) | 205 | excl. outliers diff>0.003 | 0.4464 | -0.2229 | 0.4464 | -0.49933 | 0.9917
66 | 2.047 | 8.2637 | 8.2559 | 29.0404 | 29.083 | 6.8201 | 2.4293 | 9 | 9 | 267 | 6.7962 | 0.86088 | 5.8507 |  | 0.00016859 |  |  | 190 | excl. outliers diff>0.0025 | 0.4507 | -0.2328 | 0.4507 | -0.51653 | 0.9937
50 | 31.345 | 7.8611 | 7.8617 | 29.5728 | 29.5746 | 6.8591 | 2.334 | 7 | 7 | 225 | 6.8582 | 0.81904 | 5.6171 |  | 0.00016859 |  |  | 167 | excl. outliers diff>0.002 | 0.4534 | -0.2391 | 0.4534 | -0.52735 | 0.9955
20 | 76.565 | 7.9113 | 7.9114 | 29.5557 | 29.5569 | 6.852 | 2.2703 | 10 | 10 | 135 | 6.8912 | 0.79106 | 5.4514 | 6 | 0.00018219
11 | 51.565 | 7.7822 | 7.7832 | 29.5724 | 29.5733 | 6.8715 | 2.3263 | 8 | 8 | 83 | 6.8879 | 0.81527 | 5.6156 |  | 0.00021152 |  |  | 2012-04 | excl. outliers diff>0.007 | 0.4442 | -0.2152 | 0.4442 | -0.4822 | 0.9998
35 | 21.394 | 7.8938 | 7.8952 | 29.5003 | 29.5012 | 6.8572 | 2.3086 | 3 | 3 | 168 | 6.8478 | 0.80747 | 5.5294 |  | 0.00021989
67 | 51.211 | 7.8947 | 7.8953 | 29.582 | 29.5833 | 6.8534 | 2.2885 | 2 | 2 | 269 | 6.8702 | 0.79911 | 5.49 |  | 0.00022147
66 | 101.315 | 7.8283 | 7.8296 | 29.6444 | 29.6455 | 6.861 | 2.3306 | 3 | 3 | 261 | 6.9217 | 0.81716 | 5.6562 |  | 0.00022438
51 | 21.21 | 7.8821 | 7.8806 | 29.5444 | 29.5518 | 6.8571 | 2.3444 | 6 | 6 | 234 | 6.8474 | 0.8232 | 5.6367 |  | 0.00022863
61 | 101.676 | 7.8549 | 7.8558 | 29.6128 | 29.614 | 6.8583 | 2.3241 | 2 | 2 | 252 | 6.9194 | 0.81427 | 5.6343 |  | 0.00025212
51 | 51.046 | 7.8664 | 7.8674 | 29.5849 | 29.5855 | 6.8577 | 2.3366 | 4 | 4 | 232 | 6.8742 | 0.82026 | 5.6386 |  | 0.00025903
36 | 56.786 | 7.8366 | 7.8384 | 29.5568 | 29.5576 | 6.8637 | 2.3002 | 1 | 1 | 171 | 6.885 | 0.80429 | 5.5375 |  | 0.0002608
2 | 251.694 | 7.6552 | 7.6561 | 29.7871 | 29.7886 | 6.8819 | 2.3341 | 3 | 3 | 17 | 7.0769 | 0.81865 | 5.7934 | 4 | 0.00026683
41 | 51.212 | 7.908 | 7.9089 | 29.5456 | 29.5471 | 6.853 | 2.2803 | 4 | 4 | 192 | 6.8698 | 0.795 | 5.4615 |  | 0.0002823
66 | 31.187 | 8.0202 | 8.0299 | 29.5506 | 29.5482 | 6.8353 | 2.2709 | 6 | 6 | 264 | 6.8351 | 0.79143 | 5.4095 |  | 0.00028961
23 | 200.947 | 7.7434 | 7.7443 | 29.6694 | 29.6706 | 6.8733 | 2.3208 | 4 | 4 | 144 | 7.0226 | 0.81336 | 5.7119 |  | 0.00029703
12 | 31.214 | 7.9089 | 7.9099 | 29.5201 | 29.5213 | 6.854 | 2.2562 | 9 | 9 | 96 | 6.8532 | 0.78498 | 5.3796 |  | 0.00032129
11 | 76.607 | 7.7836 | 7.7844 | 29.573 | 29.5744 | 6.8713 | 2.3215 | 7 | 7 | 82 | 6.9098 | 0.81371 | 5.6226 | 3 | 0.0003364
11 | 101.309 | 7.7822 | 7.7835 | 29.5755 | 29.5767 | 6.8714 | 2.3165 | 6 | 6 | 81 | 6.9318 | 0.81152 | 5.6253 |  | 0.00035653
57 | 51.033 | 7.9333 | 7.9333 | 29.5609 | 29.5623 | 6.8484 | 2.2805 | 4 | 4 | 246 | 6.8651 | 0.79497 | 5.4576 |  | 0.00037106
51 | 31.342 | 7.8664 | 7.8667 | 29.5801 | 29.5816 | 6.858 | 2.342 | 5 | 5 | 233 | 6.8571 | 0.82199 | 5.6364 |  | 0.00037625
4 | 126.094 | 7.7492 | 7.7499 | 29.6306 | 29.6324 | 6.8741 | 2.3468 | 5 | 5 | 44 | 6.9565 | 0.82489 | 5.7383 |  | 0.00039098
1 | 11.292 | 7.7902 | 7.7909 | 29.568 | 29.5694 | 6.8705 | 2.3624 | 13 | 13 | 13 | 6.8515 | 0.83172 | 5.6985 |  | 0.00039115
9 | 31.139 | 7.9493 | 7.9502 | 29.5147 | 29.516 | 6.8479 | 2.2703 | 9 | 9 | 60 | 6.8473 | 0.79125 | 5.4179 |  | 0.00039735
11 | 126.607 | 7.7647 | 7.7656 | 29.6238 | 29.6249 | 6.872 | 2.3399 | 5 | 5 | 80 | 6.9549 | 0.82105 | 5.7103 | 3 | 0.00040171
35 | 58.695 | 7.8275 | 7.8284 | 29.5588 | 29.5602 | 6.865 | 2.3003 | 1 | 1 | 166 | 6.888 | 0.80364 | 5.5354 |  | 0.00040279
2 | 21.027 | 7.777 | 7.7778 | 29.5459 | 29.5473 | 6.8735 | 2.3392 | 11 | 11 | 25 | 6.863 | 0.82153 | 5.6381 | 3 | 0.00040411
12 | 176.121 | 7.7564 | 7.757 | 29.6089 | 29.6103 | 6.8739 | 2.3133 | 3 | 3 | 90 | 7.0011 | 0.81019 | 5.6722 |  | 0.00042269
9 | 51.279 | 7.8844 | 7.8851 | 29.5221 | 29.5234 | 6.8577 | 2.2843 | 8 | 8 | 59 | 6.8745 | 0.7966 | 5.4762 |  | 0.00042468
2 | 101.264 | 7.7085 | 7.7094 | 29.6877 | 29.6892 | 6.8779 | 2.3627 | 7 | 7 | 21 | 6.938 | 0.83191 | 5.7718 |  | 0.00043815
57 | 76.152 | 7.8882 | 7.8894 | 29.5761 | 29.5773 | 6.8547 | 2.306 | 3 | 3 | 245 | 6.8934 | 0.80614 | 5.557 |  | 0.00043875
66 | 50.941 | 7.8953 | 7.8972 | 29.5848 | 29.5857 | 6.8532 | 2.2905 | 5 | 5 | 263 | 6.8697 | 0.80017 | 5.4969 |  | 0.00043995
35 | 31.368 | 7.8899 | 7.8909 | 29.5053 | 29.5066 | 6.8576 | 2.3031 | 2 | 2 | 167 | 6.8569 | 0.80485 | 5.5187 |  | 0.00044043
9 | 2.007 | 7.9308 | 7.9298 | 29.4093 | 29.4049 | 6.8555 | 2.3594 | 12 | 12 | 63 | 6.8293 | 0.8296 | 5.6656 |  | 0.00044223
2 | 76.667 | 7.7305 | 7.7314 | 29.6578 | 29.6593 | 6.8758 | 2.3655 | 8 | 8 | 22 | 6.9142 | 0.83316 | 5.7606 |  | 0.00045694
9 | 201.607 | 7.7182 | 7.7192 | 29.7212 | 29.7222 | 6.8749 | 2.3375 | 2 | 2 | 53 | 7.0248 | 0.82089 | 5.7666 |  | 0.00047873
48 | 31.251 | 7.8571 | 7.8582 | 29.568 | 29.5694 | 6.86 | 2.3158 | 8 | 8 | 215 | 6.8589 | 0.81038 | 5.5584 | 6 | 0.00048811
10 | 126.268 | 7.7558 | 7.7566 | 29.6371 | 29.6385 | 6.8727 | 2.3442 | 5 | 5 | 68 | 6.9553 | 0.82384 | 5.7301 |  | 0.0005173
10 | 201.276 | 7.7062 | 7.7077 | 29.7354 | 29.7357 | 6.8761 | 2.3401 | 2 | 2 | 65 | 7.0257 | 0.82207 | 5.7756 | 36 | 0.00054067
4 | 101.23 | 7.7659 | 7.7667 | 29.6113 | 29.6127 | 6.8723 | 2.3501 | 6 | 6 | 45 | 6.9327 | 0.8254 | 5.7222 |  | 0.00054145
1 | 151.589 | 7.7478 | 7.749 | 29.6463 | 29.6473 | 6.8736 | 2.3417 | 6 | 6 | 6 | 6.9787 | 0.8228 | 5.7421 |  | 0.000548
11 | 201.111 | 7.7032 | 7.704 | 29.7381 | 29.7397 | 6.8765 | 2.3399 | 2 | 2 | 77 | 7.0259 | 0.82089 | 5.7675 | 46 | 0.00055768
1 | 126.459 | 7.739 | 7.7402 | 29.6546 | 29.6557 | 6.8746 | 2.3495 | 7 | 7 | 7 | 6.9573 | 0.82624 | 5.7484 |  | 0.0005778
40 | 51.258 | 7.8613 | 7.8663 | 29.5522 | 29.5515 | 6.86 | 2.3051 | 3 | 3 | 184 | 6.8766 | 0.80675 | 5.5477 |  | 0.00059448
10 | 51.544 | 7.8498 | 7.8501 | 29.5385 | 29.5398 | 6.8624 | 2.2931 | 8 | 8 | 71 | 6.8792 | 0.80027 | 5.5052 |  | 0.00059883
4 | 10.956 | 7.935 | 7.936 | 29.4435 | 29.4446 | 6.8533 | 2.369 | 11 | 11 | 50 | 6.835 | 0.83363 | 5.6978 |  | 0.00060869
20 | 176.25 | 7.7479 | 7.7488 | 29.6456 | 29.6469 | 6.8736 | 2.3186 | 6 | 6 | 131 | 7.0008 | 0.8127 | 5.6896 |  | 0.00061116
4 | 75.951 | 7.7676 | 7.7685 | 29.5448 | 29.546 | 6.875 | 2.3145 | 7 | 7 | 46 | 6.913 | 0.81092 | 5.6059 | 6 | 0.00061393
19 | 76.176 | 7.8905 | 7.8919 | 29.5672 | 29.5682 | 6.8548 | 2.2746 | 8 | 8 | 120 | 6.8935 | 0.79214 | 5.4606 | 6 | 0.00061554
42 | 1.817 | 8.3351 | 8.3493 | 29.4126 | 29.4133 | 6.7927 | 2.4143 | 11 | 11 | 207 | 6.7691 | 0.85476 | 5.786 |  | 0.00063588
66 | 126.077 | 7.8211 | 7.822 | 29.6574 | 29.6587 | 6.8616 | 2.3264 | 2 | 2 | 260 | 6.9442 | 0.81617 | 5.6677 |  | 0.00064337
9 | 220.384 | 7.6977 | 7.6987 | 29.7535 | 29.7547 | 6.8767 | 2.3369 | 1 | 1 | 52 | 7.0434 | 0.81946 | 5.7718 |  | 0.00065097
1 | 76.287 | 7.7375 | 7.7386 | 29.6509 | 29.6521 | 6.875 | 2.3621 | 9 | 9 | 9 | 6.9131 | 0.83188 | 5.7508 | 6 | 0.00065622
13 | 126.032 | 7.7627 | 7.7636 | 29.5954 | 29.5969 | 6.8735 | 2.3146 | 6 | 6 | 105 | 6.9559 | 0.81101 | 5.6413 |  | 0.00066213
10 | 101.296 | 7.7693 | 7.7704 | 29.6119 | 29.6129 | 6.8718 | 2.3421 | 6 | 6 | 69 | 6.9322 | 0.82174 | 5.6964 |  | 0.00068303
68 | 1.892 | 8.1654 | 8.1545 | 29.3065 | 29.3252 | 6.8236 | 2.362 | 6 | 6 | 279 | 6.7988 | 0.83044 | 5.646 |  | 0.00070657
41 | 101.286 | 7.8081 | 7.8091 | 29.5901 | 29.5916 | 6.8666 | 2.3131 | 2 | 2 | 190 | 6.9272 | 0.81037 | 5.6136 |  | 0.00070945
36 | 21.249 | 7.9389 | 7.9396 | 29.4953 | 29.4968 | 6.8504 | 2.3117 | 4 | 4 | 174 | 6.8411 | 0.80834 | 5.5299 |  | 0.00071896
12 | 51.647 | 7.8955 | 7.8963 | 29.5318 | 29.533 | 6.8556 | 2.2767 | 8 | 8 | 95 | 6.8727 | 0.79294 | 5.4497 |  | 0.00072519
19 | 31.471 | 7.9439 | 7.944 | 29.5318 | 29.5332 | 6.848 | 2.2749 | 10 | 10 | 122 | 6.8477 | 0.79217 | 5.4245 | 4 | 0.00073744
50 | 151.519 | 7.7535 | 7.7543 | 29.825 | 29.8263 | 6.8647 | 2.3689 | 2 | 2 | 220 | 6.9696 | 0.83492 | 5.8191 |  | 0.00074688
12 | 126.368 | 7.7666 | 7.7671 | 29.5869 | 29.5887 | 6.8733 | 2.3168 | 5 | 5 | 92 | 6.956 | 0.81207 | 5.6488 |  | 0.0007469
36 | 51.247 | 7.8605 | 7.8617 | 29.5434 | 29.5447 | 6.8605 | 2.2962 | 2 | 2 | 172 | 6.8771 | 0.80302 | 5.5224 |  | 0.0007618
42 | 151.163 | 7.8179 | 7.8192 | 29.6303 | 29.6314 | 6.8633 | 2.3227 | 3 | 3 | 199 | 6.9683 | 0.81313 | 5.6661 | 6 | 0.00077358
61 | 30.914 | 7.9386 | 7.9404 | 29.5493 | 29.5499 | 6.848 | 2.2808 | 5 | 5 | 255 | 6.8472 | 0.79469 | 5.4414 |  | 0.00077671
50 | 154.555 | 7.7512 | 7.7518 | 29.8281 | 29.8298 | 6.8649 | 2.3694 | 1 | 1 | 219 | 6.9726 | 0.8352 | 5.8235 |  | 0.00080354
2 | 51.175 | 7.7436 | 7.745 | 29.6329 | 29.6334 | 6.8749 | 2.3662 | 9 | 9 | 23 | 6.8907 | 0.83221 | 5.7345 |  | 0.00081117
3 | 176.253 | 7.7206 | 7.7213 | 29.6817 | 29.6833 | 6.8763 | 2.3428 | 3 | 3 | 30 | 7.0034 | 0.8236 | 5.768 |  | 0.00085542
9 | 76.317 | 7.8564 | 7.8587 | 29.5305 | 29.531 | 6.8617 | 2.2866 | 7 | 7 | 58 | 6.9004 | 0.79892 | 5.5129 |  | 0.00087653
61 | 1.852 | 8.205 | 8.2009 | 29.3177 | 29.3251 | 6.8169 | 2.3978 | 8 | 8 | 258 | 6.7924 | 0.84778 | 5.7585 |  | 0.00089587
13 | 51.646 | 7.914 | 7.9148 | 29.5377 | 29.5389 | 6.8524 | 2.2712 | 9 | 9 | 108 | 6.8696 | 0.79034 | 5.4293 | 3 | 0.00090626
12 | 76.253 | 7.8528 | 7.8538 | 29.5589 | 29.5601 | 6.861 | 2.2814 | 7 | 7 | 94 | 6.8996 | 0.79483 | 5.484 |  | 0.00091375
23 | 1.56 | 7.9511 | 7.9471 | 29.4023 | 29.4067 | 6.8527 | 2.3501 | 14 | 14 | 154 | 6.8262 | 0.82502 | 5.6318 |  | 0.00091379
68 | 31.375 | 8.0015 | 8.0025 | 29.5304 | 29.5316 | 6.8391 | 2.2793 | 3 | 3 | 276 | 6.839 | 0.7939 | 5.4295 |  | 0.0009154
3 | 31.488 | 7.9162 | 7.9181 | 29.4917 | 29.493 | 6.8541 | 2.2877 | 9 | 9 | 36 | 6.8536 | 0.79943 | 5.479 |  | 0.00092403
39 | 65.125 | 7.8136 | 7.8146 | 29.5776 | 29.579 | 6.8663 | 2.3101 | 1 | 1 | 177 | 6.8949 | 0.8074 | 5.5669 |  | 0.0009513
61 | 133.866 | 7.8326 | 7.8336 | 29.6881 | 29.6892 | 6.8584 | 2.3415 | 1 | 1 | 251 | 6.948 | 0.8231 | 5.7189 |  | 0.00095242
4 | 151.121 | 7.7222 | 7.7232 | 29.6625 | 29.6638 | 6.8769 | 2.3454 | 4 | 4 | 43 | 6.9815 | 0.82483 | 5.7586 |  | 0.00096563
24 | 178.806 | 7.7563 | 7.7574 | 29.6537 | 29.6549 | 6.8719 | 2.3116 | 2 | 2 | 156 | 7.0015 | 0.80999 | 5.6711 |  | 0.00097265
50 | 21.473 | 7.8634 | 7.864 | 29.5636 | 29.5651 | 6.8592 | 2.3285 | 8 | 8 | 226 | 6.8496 | 0.81545 | 5.5855 |  | 0.00097415
13 | 151.274 | 7.7619 | 7.763 | 29.6059 | 29.6066 | 6.8732 | 2.3155 | 5 | 5 | 104 | 6.9781 | 0.81172 | 5.6642 |  | 0.0010032
52 | 46.25 | 7.8764 | 7.8767 | 29.5579 | 29.5596 | 6.8574 | 2.3346 | 2 | 2 | 238 | 6.8697 | 0.82017 | 5.6343 |  | 0.0010249
23 | 76.163 | 7.9119 | 7.9127 | 29.5831 | 29.5843 | 6.8507 | 2.2739 | 9 | 9 | 149 | 6.8895 | 0.79142 | 5.4525 |  | 0.0010267
4 | 31.01 | 7.8877 | 7.889 | 29.5072 | 29.5084 | 6.8579 | 2.2883 | 9 | 9 | 48 | 6.8568 | 0.79774 | 5.47 |  | 0.0010365
57 | 31.158 | 7.9824 | 7.9823 | 29.5423 | 29.5438 | 6.8415 | 2.2747 | 5 | 5 | 247 | 6.8412 | 0.79178 | 5.4167 |  | 0.0010407
24 | 51.715 | 8.0075 | 8.0085 | 29.5481 | 29.5495 | 6.8373 | 2.252 | 7 | 7 | 161 | 6.8552 | 0.78177 | 5.3592 |  | 0.001046
12 | 151.158 | 7.7611 | 7.7617 | 29.6042 | 29.6058 | 6.8734 | 2.32 | 4 | 4 | 91 | 6.9782 | 0.81163 | 5.6638 |  | 0.0010606
1 | 51.823 | 7.751 | 7.7513 | 29.6326 | 29.6346 | 6.8737 | 2.3669 | 10 | 10 | 10 | 6.8902 | 0.83438 | 5.749 |  | 0.001069
57 | 127.487 | 7.8463 | 7.8477 | 29.6223 | 29.6235 | 6.8592 | 2.3179 | 1 | 1 | 243 | 6.9432 | 0.81068 | 5.6287 |  | 0.0010924
19 | 176.201 | 7.7585 | 7.7593 | 29.6234 | 29.6252 | 6.873 | 2.3102 | 4 | 4 | 116 | 7.0002 | 0.80951 | 5.6667 | 4 | 0.0011066
61 | 51.179 | 7.8912 | 7.8928 | 29.5751 | 29.5759 | 6.8543 | 2.3 | 4 | 4 | 254 | 6.871 | 0.80279 | 5.516 |  | 0.0011084
4 | 201.389 | 7.7065 | 7.7074 | 29.687 | 29.6885 | 6.8783 | 2.3363 | 2 | 2 | 41 | 7.028 | 0.821 | 5.77 |  | 0.0011094
48 | 51.421 | 7.8248 | 7.8255 | 29.5796 | 29.5809 | 6.8645 | 2.3195 | 7 | 7 | 214 | 6.881 | 0.81138 | 5.5831 |  | 0.0011248
68 | 51.386 | 7.9649 | 7.9611 | 29.5558 | 29.5593 | 6.8437 | 2.2771 | 2 | 2 | 275 | 6.8609 | 0.79272 | 5.4388 |  | 0.0011292
2 | 151.258 | 7.6964 | 7.6985 | 29.7133 | 29.713 | 6.8787 | 2.353 | 6 | 6 | 20 | 6.9833 | 0.82835 | 5.7846 |  | 0.0011581
39 | 31.148 | 7.849 | 7.8497 | 29.5307 | 29.5323 | 6.8629 | 2.2993 | 2 | 2 | 178 | 6.8617 | 0.80246 | 5.5063 |  | 0.0011622
13 | 76.195 | 7.8366 | 7.8379 | 29.5626 | 29.5638 | 6.8634 | 2.2891 | 8 | 8 | 107 | 6.9019 | 0.80029 | 5.5235 |  | 0.0011681
1 | 201.94 | 7.7512 | 7.7523 | 29.6474 | 29.6487 | 6.873 | 2.3293 | 4 | 4 | 4 | 7.0233 | 0.81798 | 5.7449 |  | 0.0011784
19 | 101.296 | 7.85 | 7.848 | 29.5764 | 29.5777 | 6.8607 | 2.2824 | 7 | 7 | 119 | 6.9214 | 0.79741 | 5.5192 |  | 0.0012128
36 | 1.791 | 8.0717 | 8.0708 | 29.3288 | 29.3334 | 6.8371 | 2.3932 | 6 | 6 | 176 | 6.8117 | 0.84608 | 5.7632 |  | 0.0012222
20 | 51.536 | 8.0006 | 8.0016 | 29.5312 | 29.5324 | 6.8392 | 2.2588 | 11 | 11 | 136 | 6.8568 | 0.78703 | 5.3965 |  | 0.0012226
48 | 126.21 | 7.776 | 7.7768 | 29.7876 | 29.7887 | 6.8628 | 2.3708 | 4 | 4 | 211 | 6.9453 | 0.83626 | 5.8081 |  | 0.0012312
9 | 126.426 | 7.7702 | 7.7709 | 29.6175 | 29.6187 | 6.8714 | 2.3349 | 5 | 5 | 56 | 6.9541 | 0.82052 | 5.706 |  | 0.0012447
40 | 31.434 | 7.934 | 7.935 | 29.51 | 29.5113 | 6.8505 | 2.3031 | 4 | 4 | 185 | 6.8501 | 0.80404 | 5.5077 |  | 0.0012581
48 | 151.283 | 7.7737 | 7.7752 | 29.7967 | 29.797 | 6.8628 | 2.3658 | 3 | 3 | 210 | 6.9676 | 0.83408 | 5.8116 |  | 0.0012628
9 | 176.456 | 7.7317 | 7.7328 | 29.695 | 29.696 | 6.8739 | 2.342 | 3 | 3 | 54 | 7.0013 | 0.82364 | 5.7665 |  | 0.0012737
3 | 126.346 | 7.7202 | 7.721 | 29.6772 | 29.6786 | 6.8765 | 2.3554 | 5 | 5 | 32 | 6.959 | 0.82955 | 5.7729 |  | 0.0012883
24 | 31.685 | 8.016 | 8.0178 | 29.5467 | 29.5479 | 6.8361 | 2.2475 | 8 | 8 | 162 | 6.8364 | 0.78212 | 5.3469 |  | 0.0012898
11 | 176.377 | 7.7254 | 7.7263 | 29.693 | 29.6943 | 6.875 | 2.3408 | 3 | 3 | 78 | 7.0023 | 0.82313 | 5.7638 |  | 0.0012905
1 | 252.226 | 7.7231 | 7.7302 | 29.6878 | 29.682 | 6.8756 | 2.3188 | 3 | 3 | 3 | 7.0713 | 0.81351 | 5.7525 |  | 0.0013046
42 | 50.859 | 7.9214 | 7.924 | 29.5361 | 29.5369 | 6.8513 | 2.2807 | 7 | 7 | 203 | 6.8679 | 0.79411 | 5.4539 |  | 0.0013186
10 | 176.272 | 7.707 | 7.7079 | 29.7294 | 29.7308 | 6.8763 | 2.3452 | 3 | 3 | 66 | 7.0034 | 0.82511 | 5.7785 |  | 0.0013268
23 | 151.427 | 7.7698 | 7.7709 | 29.6218 | 29.6229 | 6.8712 | 2.3102 | 6 | 6 | 146 | 6.9763 | 0.80706 | 5.6303 |  | 0.0013283
36 | 31.069 | 7.8973 | 7.8987 | 29.5225 | 29.5237 | 6.8557 | 2.2999 | 3 | 3 | 173 | 6.8547 | 0.80253 | 5.5011 |  | 0.0013313
23 | 176.157 | 7.7495 | 7.7503 | 29.6435 | 29.645 | 6.8735 | 2.3185 | 5 | 5 | 145 | 7.0006 | 0.81341 | 5.6944 |  | 0.0013429
1 | 101.953 | 7.74 | 7.7411 | 29.6509 | 29.652 | 6.8746 | 2.3563 | 8 | 8 | 8 | 6.9355 | 0.83003 | 5.7566 |  | 0.0013656
12 | 256.712 | 7.7018 | 7.7017 | 29.7397 | 29.7433 | 6.8766 | 2.3235 | 1 | 1 | 88 | 7.0763 | 0.81565 | 5.7718 |  | 0.0013866
3 | 101.208 | 7.7282 | 7.7292 | 29.6632 | 29.6644 | 6.8759 | 2.3596 | 6 | 6 | 33 | 6.9361 | 0.83154 | 5.7676 |  | 0.0014115
24 | 101.145 | 7.9595 | 7.9637 | 29.5891 | 29.5899 | 6.843 | 2.2464 | 5 | 5 | 159 | 6.9041 | 0.78177 | 5.3974 |  | 0.0014118
23 | 319.369 | 7.7234 | 7.7243 | 29.7728 | 29.7741 | 6.8717 | 2.3088 | 1 | 1 | 141 | 7.1284 | 0.80923 | 5.7685 |  | 0.0014339
23 | 21.48 | 7.9403 | 7.9412 | 29.5281 | 29.5292 | 6.8487 | 2.2877 | 12 | 12 | 152 | 6.8396 | 0.79707 | 5.4516 |  | 0.0014565
48 | 21.14 | 7.8753 | 7.8758 | 29.5538 | 29.5553 | 6.8577 | 2.3129 | 9 | 9 | 216 | 6.8479 | 0.80813 | 5.534 |  | 0.0014612
10 | 228.055 | 7.6995 | 7.7004 | 29.7514 | 29.7529 | 6.8765 | 2.3345 | 1 | 1 | 64 | 7.0502 | 0.82056 | 5.7851 |  | 0.001478
20 | 31.55 | 7.9957 | 7.9971 | 29.4813 | 29.4825 | 6.8422 | 2.2773 | 12 | 12 | 137 | 6.8422 | 0.79248 | 5.4223 |  | 0.0014798
48 | 76.346 | 7.8607 | 7.8614 | 29.5981 | 29.5997 | 6.858 | 2.3339 | 6 | 6 | 213 | 6.8967 | 0.81732 | 5.6368 |  | 0.0014861
19 | 126.274 | 7.7986 | 7.7994 | 29.5862 | 29.5875 | 6.8683 | 2.2987 | 6 | 6 | 118 | 6.9511 | 0.80484 | 5.5945 |  | 0.0014864
57 | 101.339 | 7.8573 | 7.8582 | 29.6047 | 29.6056 | 6.8583 | 2.318 | 2 | 2 | 244 | 6.9191 | 0.81032 | 5.6067 |  | 0.0015118
61 | 76.335 | 7.8686 | 7.8709 | 29.5833 | 29.5841 | 6.8575 | 2.3028 | 3 | 3 | 253 | 6.8962 | 0.80363 | 5.542 |  | 0.0015308
23 | 31.238 | 7.9386 | 7.9396 | 29.5579 | 29.5586 | 6.8476 | 2.2792 | 11 | 11 | 151 | 6.8471 | 0.79321 | 5.4312 |  | 0.0015664
57 | 20.962 | 8.0283 | 8.0291 | 29.5106 | 29.5119 | 6.8358 | 2.2681 | 6 | 6 | 248 | 6.8268 | 0.79149 | 5.4034 |  | 0.0015693
4 | 252.91 | 7.6593 | 7.6598 | 29.7818 | 29.7837 | 6.8814 | 2.3338 | 1 | 1 | 40 | 7.0776 | 0.82037 | 5.8062 |  | 0.0015714
19 | 251.78 | 7.753 | 7.7539 | 29.6644 | 29.6653 | 6.872 | 2.3084 | 2 | 2 | 114 | 7.0673 | 0.80919 | 5.7187 |  | 0.0015934
67 | 30.944 | 7.9851 | 7.9856 | 29.543 | 29.5444 | 6.8411 | 2.2803 | 3 | 3 | 270 | 6.8405 | 0.79366 | 5.429 |  | 0.0016234
13 | 176.21 | 7.7623 | 7.7631 | 29.6141 | 29.6154 | 6.8728 | 2.3124 | 4 | 4 | 103 | 7 | 0.81101 | 5.6771 |  | 0.0016306
50 | 2.088 | 8.0872 | 8.0833 | 29.3986 | 29.4058 | 6.8316 | 2.4248 | 10 | 10 | 228 | 6.8065 | 0.85708 | 5.8337 |  | 0.0016625
24 | 21.282 | 8.0208 | 8.0224 | 29.5461 | 29.5471 | 6.8354 | 2.2511 | 9 | 9 | 163 | 6.8266 | 0.78076 | 5.33 |  | 0.001678
12 | 11.301 | 7.9431 | 7.9441 | 29.4484 | 29.4488 | 6.8518 | 2.3019 | 11 | 11 | 98 | 6.8338 | 0.80308 | 5.4881 |  | 0.0016913
41 | 31.474 | 7.9277 | 7.928 | 29.5284 | 29.5305 | 6.8507 | 2.2865 | 5 | 5 | 193 | 6.8503 | 0.79629 | 5.4548 |  | 0.001695
50 | 11.317 | 7.8872 | 7.888 | 29.5497 | 29.5511 | 6.856 | 2.3442 | 9 | 9 | 227 | 6.8377 | 0.82164 | 5.6181 |  | 0.0017036
42 | 11.26 | 8.2093 | 8.1925 | 29.4254 | 29.4291 | 6.8115 | 2.3831 | 10 | 10 | 206 | 6.7952 | 0.84216 | 5.7226 |  | 0.0017301
23 | 126.593 | 7.7869 | 7.7879 | 29.6143 | 29.6155 | 6.8689 | 2.3084 | 7 | 7 | 147 | 6.9519 | 0.80936 | 5.6266 |  | 0.0017521
42 | 125.817 | 7.8016 | 7.8027 | 29.6085 | 29.6099 | 6.8668 | 2.3126 | 4 | 4 | 200 | 6.9492 | 0.80772 | 5.613 |  | 0.001758
50 | 101.431 | 7.8388 | 7.8396 | 29.6558 | 29.6568 | 6.8589 | 2.3451 | 4 | 4 | 222 | 6.9197 | 0.82551 | 5.7123 | 6 | 0.0017606
42 | 30.928 | 7.9641 | 7.9652 | 29.5303 | 29.5317 | 6.8449 | 2.2746 | 8 | 8 | 204 | 6.8442 | 0.79099 | 5.4137 |  | 0.001766
9 | 101.486 | 7.7932 | 7.794 | 29.5858 | 29.5871 | 6.8692 | 2.3343 | 6 | 6 | 57 | 6.9299 | 0.81718 | 5.663 |  | 0.0018
19 | 21.209 | 7.9518 | 7.9536 | 29.5258 | 29.5274 | 6.847 | 2.2772 | 11 | 11 | 123 | 6.8378 | 0.79207 | 5.416 |  | 0.0018187
1 | 176.324 | 7.7479 | 7.7483 | 29.6488 | 29.6509 | 6.8735 | 2.3368 | 5 | 5 | 5 | 7.0007 | 0.8219 | 5.7539 |  | 0.001819
68 | 21.014 | 8.0424 | 8.04 | 29.4788 | 29.4838 | 6.835 | 2.2931 | 4 | 4 | 277 | 6.8261 | 0.79905 | 5.4544 |  | 0.0018195
4 | 176.222 | 7.7073 | 7.7081 | 29.6864 | 29.688 | 6.8782 | 2.3432 | 3 | 3 | 42 | 7.0052 | 0.82474 | 5.7775 |  | 0.0018304
19 | 151.288 | 7.7625 | 7.7634 | 29.6171 | 29.6183 | 6.8726 | 2.3109 | 5 | 5 | 117 | 6.9775 | 0.81056 | 5.6557 |  | 0.001843
67 | 10.903 | 8.1539 | 8.1554 | 29.3394 | 29.34 | 6.8239 | 2.369 | 5 | 5 | 272 | 6.8069 | 0.83238 | 5.6659 |  | 0.0018467
13 | 201.24 | 7.7533 | 7.754 | 29.6366 | 29.6382 | 6.8732 | 2.3164 | 3 | 3 | 102 | 7.0229 | 0.81299 | 5.7095 | 6 | 0.0018515
66 | 76.102 | 7.8268 | 7.8274 | 29.6243 | 29.6263 | 6.8622 | 2.325 | 4 | 4 | 262 | 6.9005 | 0.81682 | 5.6364 |  | 0.0019063
13 | 101.479 | 7.8056 | 7.8055 | 29.5678 | 29.5694 | 6.868 | 2.2953 | 7 | 7 | 106 | 6.9288 | 0.80381 | 5.5694 |  | 0.0019581
66 | 175.214 | 7.8177 | 7.8185 | 29.6726 | 29.674 | 6.8614 | 2.3211 | 1 | 1 | 259 | 6.9879 | 0.81519 | 5.6964 |  | 0.0019778
3 | 201.646 | 7.7135 | 7.7137 | 29.6894 | 29.6914 | 6.877 | 2.3369 | 2 | 2 | 29 | 7.027 | 0.82213 | 5.7771 |  | 0.0019999
20 | 126.541 | 7.7831 | 7.7826 | 29.6009 | 29.6028 | 6.8701 | 2.3042 | 8 | 8 | 133 | 6.953 | 0.80778 | 5.6165 |  | 0.0020137
41 | 76.15 | 7.8455 | 7.8464 | 29.5679 | 29.5694 | 6.8618 | 2.2968 | 3 | 3 | 191 | 6.9002 | 0.80049 | 5.5236 |  | 0.0020252
41 | 167.107 | 7.7583 | 7.7578 | 29.7483 | 29.755 | 6.8674 | 2.3398 | 1 | 1 | 189 | 6.9863 | 0.81938 | 5.7245 |  | 0.0020287
24 | 11.282 | 8.0134 | 8.0164 | 29.5451 | 29.5454 | 6.8366 | 2.2554 | 10 | 10 | 164 | 6.819 | 0.78231 | 5.3346 |  | 0.0020349
40 | 21.468 | 7.9524 | 7.9534 | 29.4966 | 29.498 | 6.8482 | 2.3112 | 5 | 5 | 186 | 6.8392 | 0.80682 | 5.518 |  | 0.0020423
40 | 11.119 | 8.0513 | 8.0526 | 29.4467 | 29.4479 | 6.8351 | 2.3502 | 6 | 6 | 187 | 6.8176 | 0.8239 | 5.617 |  | 0.0020741
2 | 176.245 | 7.6995 | 7.7005 | 29.712 | 29.7132 | 6.8783 | 2.3459 | 5 | 5 | 19 | 7.0053 | 0.82616 | 5.7875 |  | 0.0020759
48 | 10.953 | 7.9824 | 7.9851 | 29.5007 | 29.5012 | 6.8434 | 2.2939 | 10 | 10 | 217 | 6.8253 | 0.80338 | 5.4833 |  | 0.0021304
19 | 1.851 | 7.9863 | 7.9909 | 29.4627 | 29.4475 | 6.8445 | 2.2909 | 13 | 13 | 125 | 6.8185 | 0.79775 | 5.4395 |  | 0.0021651
20 | 101.127 | 7.8698 | 7.8703 | 29.5682 | 29.5694 | 6.8579 | 2.278 | 9 | 9 | 134 | 6.9186 | 0.7964 | 5.51 |  | 0.0021722
41 | 20.906 | 7.9787 | 7.9801 | 29.5009 | 29.5022 | 6.844 | 2.2999 | 6 | 6 | 194 | 6.8346 | 0.80172 | 5.4794 |  | 0.0021736
11 | 227.642 | 7.7016 | 7.7019 | 29.7468 | 29.749 | 6.8764 | 2.334 | 1 | 1 | 76 | 7.0497 | 0.8211 | 5.7885 |  | 0.0022377
48 | 176.361 | 7.768 | 7.7687 | 29.8101 | 29.8117 | 6.8631 | 2.3629 | 2 | 2 | 209 | 6.9903 | 0.83385 | 5.8289 |  | 0.0022804
2 | 31.256 | 7.7881 | 7.7889 | 29.5605 | 29.5618 | 6.8711 | 2.3455 | 10 | 10 | 24 | 6.8697 | 0.82621 | 5.6758 | 6 | 0.0022822
20 | 20.928 | 8.0244 | 8.0245 | 29.4231 | 29.4225 | 6.8403 | 2.2995 | 13 | 13 | 138 | 6.8312 | 0.8014 | 5.4746 |  | 0.002291
19 | 11.392 | 7.9662 | 7.9672 | 29.5014 | 29.5021 | 6.8459 | 2.2859 | 12 | 12 | 124 | 6.8281 | 0.79541 | 5.4311 |  | 0.0023198
11 | 21.31 | 7.9243 | 7.9255 | 29.5047 | 29.5059 | 6.8523 | 2.2897 | 10 | 10 | 85 | 6.8429 | 0.79706 | 5.4542 |  | 0.0023553
9 | 21.225 | 7.9431 | 7.9435 | 29.5077 | 29.5087 | 6.8492 | 2.2778 | 10 | 10 | 61 | 6.8399 | 0.79178 | 5.4157 | 6 | 0.0024027
11 | 31.634 | 7.9433 | 7.9439 | 29.5261 | 29.527 | 6.8483 | 2.2695 | 9 | 9 | 84 | 6.8482 | 0.78808 | 5.3969 |  | 0.0024407
1 | 21.367 | 7.754 | 7.7612 | 29.6221 | 29.6142 | 6.8737 | 2.3711 | 12 | 12 | 12 | 6.8634 | 0.8327 | 5.7151 |  | 0.0024455
11 | 1.754 | 7.947 | 7.9471 | 29.437 | 29.4407 | 6.8517 | 2.3245 | 12 | 12 | 87 | 6.8254 | 0.81223 | 5.5438 |  | 0.0024664
13 | 11.425 | 7.9495 | 7.9503 | 29.4531 | 29.4543 | 6.8506 | 2.3005 | 12 | 12 | 111 | 6.8328 | 0.80163 | 5.4774 |  | 0.002488
20 | 201.279 | 7.746 | 7.7464 | 29.6634 | 29.666 | 6.8731 | 2.318 | 5 | 5 | 130 | 7.0228 | 0.81432 | 5.7188 |  | 0.0024913
13 | 279.624 | 7.7086 | 7.7095 | 29.7404 | 29.7418 | 6.8755 | 2.3176 | 1 | 1 | 100 | 7.096 | 0.81418 | 5.7774 |  | 0.0025252
19 | 201.412 | 7.7604 | 7.7613 | 29.6312 | 29.6325 | 6.8723 | 2.3063 | 3 | 3 | 115 | 7.0222 | 0.80924 | 5.6826 |  | 0.0025343
9 | 151.441 | 7.7586 | 7.7593 | 29.6452 | 29.6468 | 6.872 | 2.3399 | 4 | 4 | 55 | 6.977 | 0.8241 | 5.7497 |  | 0.0026329
23 | 251.29 | 7.7435 | 7.7445 | 29.6848 | 29.6859 | 6.8726 | 2.3113 | 3 | 3 | 143 | 7.0674 | 0.81155 | 5.7355 |  | 0.0026748
12 | 101.323 | 7.7778 | 7.7801 | 29.5754 | 29.5762 | 6.8721 | 2.3112 | 6 | 6 | 93 | 6.9325 | 0.81154 | 5.6261 |  | 0.0027244
52 | 31.086 | 7.8995 | 7.9009 | 29.5367 | 29.5377 | 6.8547 | 2.3482 | 3 | 3 | 239 | 6.8538 | 0.82238 | 5.6364 |  | 0.0027304
3 | 76.331 | 7.7473 | 7.7483 | 29.6247 | 29.6261 | 6.8746 | 2.3577 | 7 | 7 | 34 | 6.9128 | 0.83206 | 5.7519 |  | 0.0027685
50 | 51.1 | 7.8685 | 7.8689 | 29.5926 | 29.5935 | 6.8571 | 2.3398 | 6 | 6 | 224 | 6.8735 | 0.81861 | 5.6268 | 3 | 0.002779
24 | 150.955 | 7.768 | 7.7673 | 29.6419 | 29.6447 | 6.8706 | 2.3125 | 3 | 3 | 157 | 6.9753 | 0.81227 | 5.6658 |  | 0.0028613
3 | 1.628 | 7.9623 | 7.9623 | 29.3399 | 29.3411 | 6.8537 | 2.4619 | 12 | 12 | 39 | 6.8273 | 0.87217 | 5.9546 |  | 0.0028676
3 | 11.8 | 7.9466 | 7.9474 | 29.4581 | 29.4596 | 6.8509 | 2.3327 | 11 | 11 | 38 | 6.8333 | 0.81539 | 5.5718 |  | 0.0028885
13 | 31.915 | 7.9181 | 7.9215 | 29.5239 | 29.5243 | 6.8524 | 2.2784 | 10 | 10 | 109 | 6.8523 | 0.79155 | 5.424 |  | 0.0028916
24 | 76.357 | 8.0157 | 8.017 | 29.5524 | 29.5535 | 6.8359 | 2.2437 | 6 | 6 | 160 | 6.8754 | 0.77626 | 5.3371 |  | 0.0029076
20 | 151.796 | 7.7683 | 7.7692 | 29.6104 | 29.6116 | 6.872 | 2.3079 | 7 | 7 | 132 | 6.9774 | 0.81029 | 5.6537 |  | 0.0029106
11 | 11.653 | 7.9465 | 7.9474 | 29.472 | 29.4733 | 6.8503 | 2.3051 | 11 | 11 | 86 | 6.8326 | 0.80321 | 5.488 |  | 0.0029623 |  | 5.9546
20 | 324.098 | 7.7398 | 7.7406 | 29.7168 | 29.7179 | 6.8717 | 2.298 | 1 | 1 | 126 | 7.1327 | 0.80605 | 5.7493 |  | 0.0029849
24 | 126.047 | 7.8767 | 7.879 | 29.5985 | 29.5995 | 6.8555 | 2.2728 | 4 | 4 | 158 | 6.9383 | 0.795 | 5.516 | 6 | 0.0030171
12 | 1.631 | 7.9696 | 7.9612 | 29.2943 | 29.3271 | 6.8546 | 2.3585 | 12 | 12 | 99 | 6.8283 | 0.8265 | 5.6436 |  | 0.0031024
19 | 51.12 | 7.9149 | 7.9159 | 29.5461 | 29.5471 | 6.8519 | 2.2747 | 9 | 9 | 121 | 6.8687 | 0.7959 | 5.4668 |  | 0.0031052
61 | 21.405 | 7.9799 | 7.9814 | 29.522 | 29.5229 | 6.8428 | 2.2855 | 6 | 6 | 256 | 6.8339 | 0.79439 | 5.4288 |  | 0.0031515
13 | 251.968 | 7.7208 | 7.7217 | 29.7091 | 29.711 | 6.875 | 2.3218 | 2 | 2 | 101 | 7.0704 | 0.81671 | 5.7745 |  | 0.003211
39 | 21.183 | 7.9085 | 7.9095 | 29.4929 | 29.4943 | 6.8553 | 2.3146 | 3 | 3 | 179 | 6.8457 | 0.80706 | 5.5249 |  | 0.0032932
10 | 21.202 | 7.9685 | 7.9685 | 29.4956 | 29.498 | 6.8458 | 2.2818 | 10 | 10 | 73 | 6.8366 | 0.79261 | 5.4187 |  | 0.0033014
23 | 301.302 | 7.7387 | 7.739 | 29.7213 | 29.724 | 6.8716 | 2.3054 | 2 | 2 | 142 | 7.1119 | 0.80966 | 5.7582 |  | 0.0033509
51 | 108.899 | 7.8142 | 7.8151 | 29.7048 | 29.7065 | 6.8606 | 2.3529 | 1 | 1 | 229 | 6.9279 | 0.82376 | 5.7069 |  | 0.0033995
3 | 21.344 | 7.9288 | 7.9298 | 29.4741 | 29.4758 | 6.8529 | 2.3051 | 10 | 10 | 37 | 6.8436 | 0.80276 | 5.4938 |  | 0.0034098
48 | 101.504 | 7.8441 | 7.8452 | 29.6472 | 29.6481 | 6.8585 | 2.3443 | 5 | 5 | 212 | 6.9193 | 0.82685 | 5.7213 |  | 0.0034656
42 | 102.328 | 7.812 | 7.8128 | 29.5822 | 29.5837 | 6.8664 | 2.2991 | 5 | 5 | 201 | 6.9279 | 0.80003 | 5.5425 | 4 | 0.0035023
35 | 11.256 | 8.016 | 8.0135 | 29.4496 | 29.4512 | 6.8404 | 2.3412 | 4 | 4 | 169 | 6.8228 | 0.81847 | 5.5843 |  | 0.0035525
2 | 301.09 | 7.6567 | 7.6575 | 29.7911 | 29.7927 | 6.8814 | 2.3225 | 2 | 2 | 16 | 7.1214 | 0.81735 | 5.8207 |  | 0.0035586
2 | 202.441 | 7.7053 | 7.706 | 29.7072 | 29.7089 | 6.8776 | 2.3388 | 4 | 4 | 18 | 7.0282 | 0.8246 | 5.7954 |  | 0.0036043
20 | 324.18 | 7.74 | 7.7407 | 29.7157 | 29.7177 | 6.8717 | 2.2978 | 2 | 2 | 127 | 7.1328 | 0.80676 | 5.7545 |  | 0.0038245
20 | 301.376 | 7.7379 | 7.7387 | 29.7132 | 29.7146 | 6.8722 | 2.3026 | 3 | 3 | 128 | 7.1124 | 0.8089 | 5.7533 |  | 0.0038586
3 | 51.089 | 7.7733 | 7.7742 | 29.5625 | 29.5632 | 6.8733 | 2.3356 | 8 | 8 | 35 | 6.8893 | 0.82343 | 5.6729 |  | 0.0038666
19 | 307.307 | 7.7379 | 7.7386 | 29.7176 | 29.719 | 6.872 | 2.3022 | 1 | 1 | 113 | 7.1177 | 0.80878 | 5.7566 |  | 0.0038853
50 | 126.234 | 7.8151 | 7.8156 | 29.7107 | 29.7129 | 6.8601 | 2.3535 | 3 | 3 | 221 | 6.9429 | 0.83134 | 5.7719 |  | 0.0038938
51 | 100.855 | 7.8491 | 7.85 | 29.6308 | 29.6322 | 6.8584 | 2.3362 | 2 | 2 | 230 | 6.9187 | 0.82374 | 5.6993 |  | 0.0039179
23 | 101.294 | 7.8394 | 7.8402 | 29.5952 | 29.5965 | 6.8615 | 2.2921 | 8 | 8 | 148 | 6.9222 | 0.79645 | 5.5132 |  | 0.0040172
13 | 1.748 | 7.9557 | 7.9562 | 29.2997 | 29.3035 | 6.8565 | 2.3519 | 13 | 13 | 112 | 6.8302 | 0.82238 | 5.617 |  | 0.0043278
2 | 11.112 | 7.8552 | 7.8559 | 29.4806 | 29.482 | 6.8642 | 2.3411 | 12 | 12 | 26 | 6.8454 | 0.81751 | 5.5962 |  | 0.0044647
41 | 11.139 | 8.0114 | 8.0129 | 29.4746 | 29.4757 | 6.84 | 2.323 | 7 | 7 | 195 | 6.8223 | 0.80952 | 5.5228 |  | 0.0045044
52 | 2.07 | 8.4492 | 8.4405 | 29.2686 | 29.2767 | 6.7815 | 2.6773 | 6 | 6 | 242 | 6.759 | 0.96518 | 6.5236 |  | 0.0045216
24 | 301.802 | 7.7303 | 7.7312 | 29.7421 | 29.7434 | 6.872 | 2.3072 | 1 | 1 | 155 | 7.1127 | 0.81166 | 5.7731 | 3 | 0.0045523
39 | 2.089 | 8.1257 | 8.1286 | 29.3643 | 29.3639 | 6.8271 | 2.4196 | 5 | 5 | 181 | 6.8023 | 0.8518 | 5.7942 |  | 0.0046612
48 | 1.881 | 8.0891 | 8.0988 | 29.3548 | 29.3437 | 6.8333 | 2.3505 | 11 | 11 | 218 | 6.808 | 0.82114 | 5.5903 |  | 0.0049693
2 | 315.703 | 7.659 | 7.6599 | 29.7896 | 29.7912 | 6.8811 | 2.3181 | 1 | 1 | 15 | 7.1344 | 0.81706 | 5.8292 |  | 0.0052008
50 | 76.442 | 7.86 | 7.8619 | 29.5927 | 29.5957 | 6.8584 | 2.3286 | 5 | 5 | 223 | 6.8972 | 0.82191 | 5.6689 |  | 0.0054346
66 | 21.199 | 8.0589 | 8.0598 | 29.4753 | 29.4766 | 6.8326 | 2.3042 | 7 | 7 | 265 | 6.824 | 0.80019 | 5.4605 |  | 0.0055831
4 | 21.101 | 7.9271 | 7.9281 | 29.4677 | 29.469 | 6.8535 | 2.3296 | 10 | 10 | 49 | 6.844 | 0.8113 | 5.5525 |  | 0.005615
12 | 21.296 | 7.9234 | 7.9244 | 29.4981 | 29.4986 | 6.8527 | 2.2665 | 10 | 10 | 97 | 6.8433 | 0.78346 | 5.3615 |  | 0.0057236
61 | 11.296 | 8.0401 | 8.0411 | 29.4689 | 29.4702 | 6.8358 | 2.3165 | 7 | 7 | 257 | 6.8184 | 0.8053 | 5.4909 |  | 0.0058522
12 | 201.282 | 7.7371 | 7.7399 | 29.635 | 29.6322 | 6.8758 | 2.3123 | 2 | 2 | 89 | 7.0255 | 0.81517 | 5.727 | 6 | 0.0058629
10 | 76.077 | 7.7989 | 7.8009 | 29.5549 | 29.5557 | 6.8697 | 2.3079 | 7 | 7 | 70 | 6.9079 | 0.81353 | 5.6197 |  | 0.0061513
20 | 11.472 | 7.9903 | 7.9918 | 29.3321 | 29.3333 | 6.8497 | 2.3349 | 14 | 14 | 139 | 6.8321 | 0.81272 | 5.5526 |  | 0.0065142
68 | 97.14 | 7.8187 | 7.8195 | 29.6492 | 29.6509 | 6.8623 | 2.3281 | 1 | 1 | 274 | 6.9192 | 0.8097 | 5.6025 |  | 0.0065458
42 | 20.892 | 8.0042 | 8.0054 | 29.4976 | 29.4993 | 6.8401 | 2.3028 | 9 | 9 | 205 | 6.8309 | 0.79847 | 5.4543 |  | 0.0067033
35 | 1.942 | 8.0373 | 8.0359 | 29.3945 | 29.4017 | 6.8396 | 2.3666 | 5 | 5 | 170 | 6.814 | 0.83991 | 5.7231 |  | 0.006736
48 | 202.95 | 7.7485 | 7.7492 | 29.837 | 29.8385 | 6.8649 | 2.3569 | 1 | 1 | 208 | 7.016 | 0.83601 | 5.8654 |  | 0.0070696
52 | 46.263 | 7.8765 | 7.8767 | 29.5577 | 29.5595 | 6.8574 | 2.3345 | 1 | 1 | 237 | 6.8697 | 0.82621 | 5.6758 |  | 0.0071098
52 | 21.388 | 7.9566 | 7.9574 | 29.4927 | 29.4941 | 6.8478 | 2.3852 | 4 | 4 | 240 | 6.8387 | 0.83416 | 5.7046 |  | 0.0071887
67 | 1.672 | 8.2716 | 8.2691 | 29.1533 | 29.1595 | 6.8139 | 2.4236 | 6 | 6 | 273 | 6.7897 | 0.85098 | 5.7779 |  | 0.0072538
66 | 11.125 | 8.1306 | 8.1315 | 29.393 | 29.394 | 6.8251 | 2.3489 | 8 | 8 | 266 | 6.8082 | 0.81802 | 5.5693 |  | 0.0073732
67 | 21.028 | 8.0785 | 8.0789 | 29.4574 | 29.4593 | 6.8303 | 2.3119 | 4 | 4 | 271 | 6.8217 | 0.80128 | 5.4661 |  | 0.007872
1 | 31.022 | 7.7434 | 7.7434 | 29.6393 | 29.6417 | 6.8746 | 2.3733 | 11 | 11 | 11 | 6.8727 | 0.84414 | 5.8015 | 4 | 0.0080029
23 | 11.371 | 7.9311 | 7.9357 | 29.4236 | 29.4202 | 6.8548 | 2.3413 | 13 | 13 | 153 | 6.8368 | 0.81364 | 5.5627 |  | 0.008428
10 | 1.525 | 7.9259 | 7.9264 | 29.3822 | 29.3845 | 6.8575 | 2.3646 | 12 | 12 | 75 | 6.8308 | 0.82385 | 5.6275 |  | 0.0084458
1 | 300.784 | 7.7435 | 7.7412 | 29.6678 | 29.673 | 6.8733 | 2.3075 | 2 | 2 | 2 | 7.1131 | 0.79863 | 5.6808 |  | 0.0085985
41 | 1.902 | 8.2203 | 8.2665 | 29.3949 | 29.3784 | 6.8112 | 2.3742 | 8 | 8 | 196 | 6.7868 | 0.82792 | 5.619 |  | 0.0086006
20 | 1.83 | 7.9967 | 7.9952 | 29.0148 | 29.0283 | 6.8628 | 2.4021 | 15 | 15 | 140 | 6.8368 | 0.83937 | 5.7386 |  | 0.0094158
36 | 11.125 | 8.0295 | 8.0315 | 29.4503 | 29.4513 | 6.8383 | 2.3665 | 5 | 5 | 175 | 6.8207 | 0.82302 | 5.6136 |  | 0.010124
67 | 69.614 | 7.8771 | 7.8781 | 29.5917 | 29.593 | 6.8557 | 2.2928 | 1 | 1 | 268 | 6.8886 | 0.79038 | 5.4446 | 3 | 0.010371
2 | 1.747 | 7.9728 | 7.9653 | 29.4301 | 29.4345 | 6.848 | 2.411 | 13 | 13 | 27 | 6.8219 | 0.84175 | 5.7423 |  | 0.010967
42 | 76.479 | 7.8758 | 7.8761 | 29.5536 | 29.555 | 6.8577 | 2.2877 | 6 | 6 | 202 | 6.8966 | 0.78641 | 5.4235 | 4 | 0.012097
9 | 11.269 | 7.9256 | 7.9273 | 29.4628 | 29.46 | 6.8539 | 2.3351 | 11 | 11 | 62 | 6.8358 | 0.80517 | 5.504 |  | 0.014184
39 | 11.347 | 8.0328 | 8.036 | 29.4417 | 29.4423 | 6.8381 | 2.3787 | 4 | 4 | 180 | 6.8208 | 0.82216 | 5.6077 |  | 0.016331
24 | 1.658 | 8.4904 | 8.4914 | 28.5888 | 28.5874 | 6.8051 | 2.4667 | 11 | 11 | 165 | 6.7824 | 0.86077 | 5.8381 |  | 0.016397
57 | 11.103 | 8.2333 | 8.2342 | 29.3172 | 29.3184 | 6.8126 | 2.436 | 7 | 7 | 249 | 6.7963 | 0.84544 | 5.7459 |  | 0.018249
52 | 11.445 | 8.1431 | 8.1302 | 29.3838 | 29.3918 | 6.8236 | 2.4829 | 5 | 5 | 241 | 6.807 | 0.86534 | 5.8904 |  | 0.018956
51 | 1.947 | 8.4061 | 8.3964 | 29.2145 | 29.222 | 6.7905 | 2.5755 | 8 | 8 | 236 | 6.7675 | 0.94496 | 6.395 |  | 0.020005
4 | 1.633 | 7.979 | 7.9823 | 29.3913 | 29.3896 | 6.8488 | 2.4359 | 12 | 12 | 51 | 6.8226 | 0.8897 | 6.07 |  | 0.026052
40 | 1.852 | 8.5835 | 8.5872 | 29.0969 | 29.0924 | 6.7685 | 2.4878 | 7 | 7 | 188 | 6.7468 | 0.85499 | 5.7684 |  | 0.03143
42 | 175.67 | 7.7933 | 7.7939 | 29.6869 | 29.689 | 6.8646 | 2.3326 | 2 | 2 | 198 | 6.9914 | 0.78576 | 5.4936 | 4 | 0.032476
51 | 11.287 | 8.1737 | 8.1693 | 29.3522 | 29.3564 | 6.8203 | 2.5142 | 7 | 7 | 235 | 6.8038 | 0.86298 | 5.8715 |  | 0.035077
68 | 11.562 | 8.1019 | 8.1016 | 29.3907 | 29.3926 | 6.8297 | 2.3314 | 5 | 5 | 278 | 6.8129 | 0.78011 | 5.3149 |  | 0.037608
57 | 1.794 | 8.4422 | 8.4458 | 29.1206 | 29.1214 | 6.7891 | 2.644 | 8 | 8 | 250 | 6.7662 | 0.91093 | 6.1636 |  | 0.044125
42 | 223.091 | 7.7554 | 7.7565 | 29.7793 | 29.7807 | 6.8664 | 2.339 | 1 | 1 | 197 | 7.0357 | 0.76874 | 5.4086 | 4 | 0.052303
3 | 281.047 | 7.699 | 7.6998 | 29.716 | 29.7176 | 6.8781 | 2.3159 | 1 | 1 | 28 | 7.0999 | 0.89355 | 6.3441 |  | 0.08264
1 | 336.707 | 7.7169 | 7.7168 | 29.7049 | 29.7077 | 6.8758 | 2.3023 | 1 | 1 | 1 |  |  | -99 | 5
3 | 151.089 | 7.7202 | 7.7209 | 29.6791 | 29.6806 | 6.8765 | 2.3493 | 4 | 4 | 31 |  |  | -99 | 5
20 | 251.531 | 7.7486 | 7.7498 | 29.6704 | 29.6711 | 6.8724 | 2.3077 | 4 | 4 | 129 |  |  | -99 | 5

## CTD rosette data
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number
1 | 336.707 | 7.7169 | 7.7168 | 29.7049 | 29.7077 | 6.8758 | 2.3023 | 1 | 1 | 1
1 | 300.784 | 7.7435 | 7.7412 | 29.6678 | 29.673 | 6.8733 | 2.3075 | 2 | 2 | 2
1 | 252.226 | 7.7231 | 7.7302 | 29.6878 | 29.682 | 6.8756 | 2.3188 | 3 | 3 | 3
1 | 201.94 | 7.7512 | 7.7523 | 29.6474 | 29.6487 | 6.873 | 2.3293 | 4 | 4 | 4
1 | 176.324 | 7.7479 | 7.7483 | 29.6488 | 29.6509 | 6.8735 | 2.3368 | 5 | 5 | 5
1 | 151.589 | 7.7478 | 7.749 | 29.6463 | 29.6473 | 6.8736 | 2.3417 | 6 | 6 | 6
1 | 126.459 | 7.739 | 7.7402 | 29.6546 | 29.6557 | 6.8746 | 2.3495 | 7 | 7 | 7
1 | 101.953 | 7.74 | 7.7411 | 29.6509 | 29.652 | 6.8746 | 2.3563 | 8 | 8 | 8
1 | 76.287 | 7.7375 | 7.7386 | 29.6509 | 29.6521 | 6.875 | 2.3621 | 9 | 9 | 9
1 | 51.823 | 7.751 | 7.7513 | 29.6326 | 29.6346 | 6.8737 | 2.3669 | 10 | 10 | 10
1 | 31.022 | 7.7434 | 7.7434 | 29.6393 | 29.6417 | 6.8746 | 2.3733 | 11 | 11 | 11
1 | 21.367 | 7.754 | 7.7612 | 29.6221 | 29.6142 | 6.8737 | 2.3711 | 12 | 12 | 12
1 | 11.292 | 7.7902 | 7.7909 | 29.568 | 29.5694 | 6.8705 | 2.3624 | 13 | 13 | 13
1 | 1.907 | 7.7938 | 7.7936 | 29.5563 | 29.561 | 6.8704 | 2.3601 | 14 | 14 | 14
2 | 315.703 | 7.659 | 7.6599 | 29.7896 | 29.7912 | 6.8811 | 2.3181 | 1 | 1 | 15
2 | 301.09 | 7.6567 | 7.6575 | 29.7911 | 29.7927 | 6.8814 | 2.3225 | 2 | 2 | 16
2 | 251.694 | 7.6552 | 7.6561 | 29.7871 | 29.7886 | 6.8819 | 2.3341 | 3 | 3 | 17
2 | 202.441 | 7.7053 | 7.706 | 29.7072 | 29.7089 | 6.8776 | 2.3388 | 4 | 4 | 18
2 | 176.245 | 7.6995 | 7.7005 | 29.712 | 29.7132 | 6.8783 | 2.3459 | 5 | 5 | 19
2 | 151.258 | 7.6964 | 7.6985 | 29.7133 | 29.713 | 6.8787 | 2.353 | 6 | 6 | 20
2 | 101.264 | 7.7085 | 7.7094 | 29.6877 | 29.6892 | 6.8779 | 2.3627 | 7 | 7 | 21
2 | 76.667 | 7.7305 | 7.7314 | 29.6578 | 29.6593 | 6.8758 | 2.3655 | 8 | 8 | 22
2 | 51.175 | 7.7436 | 7.745 | 29.6329 | 29.6334 | 6.8749 | 2.3662 | 9 | 9 | 23
2 | 31.256 | 7.7881 | 7.7889 | 29.5605 | 29.5618 | 6.8711 | 2.3455 | 10 | 10 | 24
2 | 21.027 | 7.777 | 7.7778 | 29.5459 | 29.5473 | 6.8735 | 2.3392 | 11 | 11 | 25
2 | 11.112 | 7.8552 | 7.8559 | 29.4806 | 29.482 | 6.8642 | 2.3411 | 12 | 12 | 26
2 | 1.747 | 7.9728 | 7.9653 | 29.4301 | 29.4345 | 6.848 | 2.411 | 13 | 13 | 27
3 | 281.047 | 7.699 | 7.6998 | 29.716 | 29.7176 | 6.8781 | 2.3159 | 1 | 1 | 28
3 | 201.646 | 7.7135 | 7.7137 | 29.6894 | 29.6914 | 6.877 | 2.3369 | 2 | 2 | 29
3 | 176.253 | 7.7206 | 7.7213 | 29.6817 | 29.6833 | 6.8763 | 2.3428 | 3 | 3 | 30
3 | 151.089 | 7.7202 | 7.7209 | 29.6791 | 29.6806 | 6.8765 | 2.3493 | 4 | 4 | 31
3 | 126.346 | 7.7202 | 7.721 | 29.6772 | 29.6786 | 6.8765 | 2.3554 | 5 | 5 | 32
3 | 101.208 | 7.7282 | 7.7292 | 29.6632 | 29.6644 | 6.8759 | 2.3596 | 6 | 6 | 33
3 | 76.331 | 7.7473 | 7.7483 | 29.6247 | 29.6261 | 6.8746 | 2.3577 | 7 | 7 | 34
3 | 51.089 | 7.7733 | 7.7742 | 29.5625 | 29.5632 | 6.8733 | 2.3356 | 8 | 8 | 35
3 | 31.488 | 7.9162 | 7.9181 | 29.4917 | 29.493 | 6.8541 | 2.2877 | 9 | 9 | 36
3 | 21.344 | 7.9288 | 7.9298 | 29.4741 | 29.4758 | 6.8529 | 2.3051 | 10 | 10 | 37
3 | 11.8 | 7.9466 | 7.9474 | 29.4581 | 29.4596 | 6.8509 | 2.3327 | 11 | 11 | 38
3 | 1.628 | 7.9623 | 7.9623 | 29.3399 | 29.3411 | 6.8537 | 2.4619 | 12 | 12 | 39
4 | 252.91 | 7.6593 | 7.6598 | 29.7818 | 29.7837 | 6.8814 | 2.3338 | 1 | 1 | 40
4 | 201.389 | 7.7065 | 7.7074 | 29.687 | 29.6885 | 6.8783 | 2.3363 | 2 | 2 | 41
4 | 176.222 | 7.7073 | 7.7081 | 29.6864 | 29.688 | 6.8782 | 2.3432 | 3 | 3 | 42
4 | 151.121 | 7.7222 | 7.7232 | 29.6625 | 29.6638 | 6.8769 | 2.3454 | 4 | 4 | 43
4 | 126.094 | 7.7492 | 7.7499 | 29.6306 | 29.6324 | 6.8741 | 2.3468 | 5 | 5 | 44
4 | 101.23 | 7.7659 | 7.7667 | 29.6113 | 29.6127 | 6.8723 | 2.3501 | 6 | 6 | 45
4 | 75.951 | 7.7676 | 7.7685 | 29.5448 | 29.546 | 6.875 | 2.3145 | 7 | 7 | 46
4 | 51.209 | 7.8253 | 7.8255 | 29.523 | 29.5248 | 6.867 | 2.3006 | 8 | 8 | 47
4 | 31.01 | 7.8877 | 7.889 | 29.5072 | 29.5084 | 6.8579 | 2.2883 | 9 | 9 | 48
4 | 21.101 | 7.9271 | 7.9281 | 29.4677 | 29.469 | 6.8535 | 2.3296 | 10 | 10 | 49
4 | 10.956 | 7.935 | 7.936 | 29.4435 | 29.4446 | 6.8533 | 2.369 | 11 | 11 | 50
4 | 1.633 | 7.979 | 7.9823 | 29.3913 | 29.3896 | 6.8488 | 2.4359 | 12 | 12 | 51
9 | 220.384 | 7.6977 | 7.6987 | 29.7535 | 29.7547 | 6.8767 | 2.3369 | 1 | 1 | 52
9 | 201.607 | 7.7182 | 7.7192 | 29.7212 | 29.7222 | 6.8749 | 2.3375 | 2 | 2 | 53
9 | 176.456 | 7.7317 | 7.7328 | 29.695 | 29.696 | 6.8739 | 2.342 | 3 | 3 | 54
9 | 151.441 | 7.7586 | 7.7593 | 29.6452 | 29.6468 | 6.872 | 2.3399 | 4 | 4 | 55
9 | 126.426 | 7.7702 | 7.7709 | 29.6175 | 29.6187 | 6.8714 | 2.3349 | 5 | 5 | 56
9 | 101.486 | 7.7932 | 7.794 | 29.5858 | 29.5871 | 6.8692 | 2.3343 | 6 | 6 | 57
9 | 76.317 | 7.8564 | 7.8587 | 29.5305 | 29.531 | 6.8617 | 2.2866 | 7 | 7 | 58
9 | 51.279 | 7.8844 | 7.8851 | 29.5221 | 29.5234 | 6.8577 | 2.2843 | 8 | 8 | 59
9 | 31.139 | 7.9493 | 7.9502 | 29.5147 | 29.516 | 6.8479 | 2.2703 | 9 | 9 | 60
9 | 21.225 | 7.9431 | 7.9435 | 29.5077 | 29.5087 | 6.8492 | 2.2778 | 10 | 10 | 61
9 | 11.269 | 7.9256 | 7.9273 | 29.4628 | 29.46 | 6.8539 | 2.3351 | 11 | 11 | 62
9 | 2.007 | 7.9308 | 7.9298 | 29.4093 | 29.4049 | 6.8555 | 2.3594 | 12 | 12 | 63
10 | 228.055 | 7.6995 | 7.7004 | 29.7514 | 29.7529 | 6.8765 | 2.3345 | 1 | 1 | 64
10 | 201.276 | 7.7062 | 7.7077 | 29.7354 | 29.7357 | 6.8761 | 2.3401 | 2 | 2 | 65
10 | 176.272 | 7.707 | 7.7079 | 29.7294 | 29.7308 | 6.8763 | 2.3452 | 3 | 3 | 66
10 | 151.357 | 7.7349 | 7.736 | 29.6725 | 29.6734 | 6.8745 | 2.3448 | 4 | 4 | 67
10 | 126.268 | 7.7558 | 7.7566 | 29.6371 | 29.6385 | 6.8727 | 2.3442 | 5 | 5 | 68
10 | 101.296 | 7.7693 | 7.7704 | 29.6119 | 29.6129 | 6.8718 | 2.3421 | 6 | 6 | 69
10 | 76.077 | 7.7989 | 7.8009 | 29.5549 | 29.5557 | 6.8697 | 2.3079 | 7 | 7 | 70
10 | 51.544 | 7.8498 | 7.8501 | 29.5385 | 29.5398 | 6.8624 | 2.2931 | 8 | 8 | 71
10 | 31.262 | 7.951 | 7.9512 | 29.5253 | 29.5266 | 6.8472 | 2.2677 | 9 | 9 | 72
10 | 21.202 | 7.9685 | 7.9685 | 29.4956 | 29.498 | 6.8458 | 2.2818 | 10 | 10 | 73
10 | 11.306 | 7.921 | 7.9153 | 29.4586 | 29.4481 | 6.8548 | 2.3131 | 11 | 11 | 74
10 | 1.525 | 7.9259 | 7.9264 | 29.3822 | 29.3845 | 6.8575 | 2.3646 | 12 | 12 | 75
11 | 227.642 | 7.7016 | 7.7019 | 29.7468 | 29.749 | 6.8764 | 2.334 | 1 | 1 | 76
11 | 201.111 | 7.7032 | 7.704 | 29.7381 | 29.7397 | 6.8765 | 2.3399 | 2 | 2 | 77
11 | 176.377 | 7.7254 | 7.7263 | 29.693 | 29.6943 | 6.875 | 2.3408 | 3 | 3 | 78
11 | 151.386 | 7.7374 | 7.7374 | 29.6722 | 29.6746 | 6.8741 | 2.3436 | 4 | 4 | 79
11 | 126.607 | 7.7647 | 7.7656 | 29.6238 | 29.6249 | 6.872 | 2.3399 | 5 | 5 | 80
11 | 101.309 | 7.7822 | 7.7835 | 29.5755 | 29.5767 | 6.8714 | 2.3165 | 6 | 6 | 81
11 | 76.607 | 7.7836 | 7.7844 | 29.573 | 29.5744 | 6.8713 | 2.3215 | 7 | 7 | 82
11 | 51.565 | 7.7822 | 7.7832 | 29.5724 | 29.5733 | 6.8715 | 2.3263 | 8 | 8 | 83
11 | 31.634 | 7.9433 | 7.9439 | 29.5261 | 29.527 | 6.8483 | 2.2695 | 9 | 9 | 84
11 | 21.31 | 7.9243 | 7.9255 | 29.5047 | 29.5059 | 6.8523 | 2.2897 | 10 | 10 | 85
11 | 11.653 | 7.9465 | 7.9474 | 29.472 | 29.4733 | 6.8503 | 2.3051 | 11 | 11 | 86
11 | 1.754 | 7.947 | 7.9471 | 29.437 | 29.4407 | 6.8517 | 2.3245 | 12 | 12 | 87
12 | 256.712 | 7.7018 | 7.7017 | 29.7397 | 29.7433 | 6.8766 | 2.3235 | 1 | 1 | 88
12 | 201.282 | 7.7371 | 7.7399 | 29.635 | 29.6322 | 6.8758 | 2.3123 | 2 | 2 | 89
12 | 176.121 | 7.7564 | 7.757 | 29.6089 | 29.6103 | 6.8739 | 2.3133 | 3 | 3 | 90
12 | 151.158 | 7.7611 | 7.7617 | 29.6042 | 29.6058 | 6.8734 | 2.32 | 4 | 4 | 91
12 | 126.368 | 7.7666 | 7.7671 | 29.5869 | 29.5887 | 6.8733 | 2.3168 | 5 | 5 | 92
12 | 101.323 | 7.7778 | 7.7801 | 29.5754 | 29.5762 | 6.8721 | 2.3112 | 6 | 6 | 93
12 | 76.253 | 7.8528 | 7.8538 | 29.5589 | 29.5601 | 6.861 | 2.2814 | 7 | 7 | 94
12 | 51.647 | 7.8955 | 7.8963 | 29.5318 | 29.533 | 6.8556 | 2.2767 | 8 | 8 | 95
12 | 31.214 | 7.9089 | 7.9099 | 29.5201 | 29.5213 | 6.854 | 2.2562 | 9 | 9 | 96
12 | 21.296 | 7.9234 | 7.9244 | 29.4981 | 29.4986 | 6.8527 | 2.2665 | 10 | 10 | 97
12 | 11.301 | 7.9431 | 7.9441 | 29.4484 | 29.4488 | 6.8518 | 2.3019 | 11 | 11 | 98
12 | 1.631 | 7.9696 | 7.9612 | 29.2943 | 29.3271 | 6.8546 | 2.3585 | 12 | 12 | 99
13 | 279.624 | 7.7086 | 7.7095 | 29.7404 | 29.7418 | 6.8755 | 2.3176 | 1 | 1 | 100
13 | 251.968 | 7.7208 | 7.7217 | 29.7091 | 29.711 | 6.875 | 2.3218 | 2 | 2 | 101
13 | 201.24 | 7.7533 | 7.754 | 29.6366 | 29.6382 | 6.8732 | 2.3164 | 3 | 3 | 102
13 | 176.21 | 7.7623 | 7.7631 | 29.6141 | 29.6154 | 6.8728 | 2.3124 | 4 | 4 | 103
13 | 151.274 | 7.7619 | 7.763 | 29.6059 | 29.6066 | 6.8732 | 2.3155 | 5 | 5 | 104
13 | 126.032 | 7.7627 | 7.7636 | 29.5954 | 29.5969 | 6.8735 | 2.3146 | 6 | 6 | 105
13 | 101.479 | 7.8056 | 7.8055 | 29.5678 | 29.5694 | 6.868 | 2.2953 | 7 | 7 | 106
13 | 76.195 | 7.8366 | 7.8379 | 29.5626 | 29.5638 | 6.8634 | 2.2891 | 8 | 8 | 107
13 | 51.646 | 7.914 | 7.9148 | 29.5377 | 29.5389 | 6.8524 | 2.2712 | 9 | 9 | 108
13 | 31.915 | 7.9181 | 7.9215 | 29.5239 | 29.5243 | 6.8524 | 2.2784 | 10 | 10 | 109
13 | 21.677 | 7.9709 | 7.9714 | 29.4879 | 29.4882 | 6.8457 | 2.2763 | 11 | 11 | 110
13 | 11.425 | 7.9495 | 7.9503 | 29.4531 | 29.4543 | 6.8506 | 2.3005 | 12 | 12 | 111
13 | 1.748 | 7.9557 | 7.9562 | 29.2997 | 29.3035 | 6.8565 | 2.3519 | 13 | 13 | 112
19 | 307.307 | 7.7379 | 7.7386 | 29.7176 | 29.719 | 6.872 | 2.3022 | 1 | 1 | 113
19 | 251.78 | 7.753 | 7.7539 | 29.6644 | 29.6653 | 6.872 | 2.3084 | 2 | 2 | 114
19 | 201.412 | 7.7604 | 7.7613 | 29.6312 | 29.6325 | 6.8723 | 2.3063 | 3 | 3 | 115
19 | 176.201 | 7.7585 | 7.7593 | 29.6234 | 29.6252 | 6.873 | 2.3102 | 4 | 4 | 116
19 | 151.288 | 7.7625 | 7.7634 | 29.6171 | 29.6183 | 6.8726 | 2.3109 | 5 | 5 | 117
19 | 126.274 | 7.7986 | 7.7994 | 29.5862 | 29.5875 | 6.8683 | 2.2987 | 6 | 6 | 118
19 | 101.296 | 7.85 | 7.848 | 29.5764 | 29.5777 | 6.8607 | 2.2824 | 7 | 7 | 119
19 | 76.176 | 7.8905 | 7.8919 | 29.5672 | 29.5682 | 6.8548 | 2.2746 | 8 | 8 | 120
19 | 51.12 | 7.9149 | 7.9159 | 29.5461 | 29.5471 | 6.8519 | 2.2747 | 9 | 9 | 121
19 | 31.471 | 7.9439 | 7.944 | 29.5318 | 29.5332 | 6.848 | 2.2749 | 10 | 10 | 122
19 | 21.209 | 7.9518 | 7.9536 | 29.5258 | 29.5274 | 6.847 | 2.2772 | 11 | 11 | 123
19 | 11.392 | 7.9662 | 7.9672 | 29.5014 | 29.5021 | 6.8459 | 2.2859 | 12 | 12 | 124
19 | 1.851 | 7.9863 | 7.9909 | 29.4627 | 29.4475 | 6.8445 | 2.2909 | 13 | 13 | 125
20 | 324.098 | 7.7398 | 7.7406 | 29.7168 | 29.7179 | 6.8717 | 2.298 | 1 | 1 | 126
20 | 324.18 | 7.74 | 7.7407 | 29.7157 | 29.7177 | 6.8717 | 2.2978 | 2 | 2 | 127
20 | 301.376 | 7.7379 | 7.7387 | 29.7132 | 29.7146 | 6.8722 | 2.3026 | 3 | 3 | 128
20 | 251.531 | 7.7486 | 7.7498 | 29.6704 | 29.6711 | 6.8724 | 2.3077 | 4 | 4 | 129
20 | 201.279 | 7.746 | 7.7464 | 29.6634 | 29.666 | 6.8731 | 2.318 | 5 | 5 | 130
20 | 176.25 | 7.7479 | 7.7488 | 29.6456 | 29.6469 | 6.8736 | 2.3186 | 6 | 6 | 131
20 | 151.796 | 7.7683 | 7.7692 | 29.6104 | 29.6116 | 6.872 | 2.3079 | 7 | 7 | 132
20 | 126.541 | 7.7831 | 7.7826 | 29.6009 | 29.6028 | 6.8701 | 2.3042 | 8 | 8 | 133
20 | 101.127 | 7.8698 | 7.8703 | 29.5682 | 29.5694 | 6.8579 | 2.278 | 9 | 9 | 134
20 | 76.565 | 7.9113 | 7.9114 | 29.5557 | 29.5569 | 6.852 | 2.2703 | 10 | 10 | 135
20 | 51.536 | 8.0006 | 8.0016 | 29.5312 | 29.5324 | 6.8392 | 2.2588 | 11 | 11 | 136
20 | 31.55 | 7.9957 | 7.9971 | 29.4813 | 29.4825 | 6.8422 | 2.2773 | 12 | 12 | 137
20 | 20.928 | 8.0244 | 8.0245 | 29.4231 | 29.4225 | 6.8403 | 2.2995 | 13 | 13 | 138
20 | 11.472 | 7.9903 | 7.9918 | 29.3321 | 29.3333 | 6.8497 | 2.3349 | 14 | 14 | 139
20 | 1.83 | 7.9967 | 7.9952 | 29.0148 | 29.0283 | 6.8628 | 2.4021 | 15 | 15 | 140
23 | 319.369 | 7.7234 | 7.7243 | 29.7728 | 29.7741 | 6.8717 | 2.3088 | 1 | 1 | 141
23 | 301.302 | 7.7387 | 7.739 | 29.7213 | 29.724 | 6.8716 | 2.3054 | 2 | 2 | 142
23 | 251.29 | 7.7435 | 7.7445 | 29.6848 | 29.6859 | 6.8726 | 2.3113 | 3 | 3 | 143
23 | 200.947 | 7.7434 | 7.7443 | 29.6694 | 29.6706 | 6.8733 | 2.3208 | 4 | 4 | 144
23 | 176.157 | 7.7495 | 7.7503 | 29.6435 | 29.645 | 6.8735 | 2.3185 | 5 | 5 | 145
23 | 151.427 | 7.7698 | 7.7709 | 29.6218 | 29.6229 | 6.8712 | 2.3102 | 6 | 6 | 146
23 | 126.593 | 7.7869 | 7.7879 | 29.6143 | 29.6155 | 6.8689 | 2.3084 | 7 | 7 | 147
23 | 101.294 | 7.8394 | 7.8402 | 29.5952 | 29.5965 | 6.8615 | 2.2921 | 8 | 8 | 148
23 | 76.163 | 7.9119 | 7.9127 | 29.5831 | 29.5843 | 6.8507 | 2.2739 | 9 | 9 | 149
23 | 51.165 | 7.9216 | 7.9231 | 29.5667 | 29.5676 | 6.8499 | 2.277 | 10 | 10 | 150
23 | 31.238 | 7.9386 | 7.9396 | 29.5579 | 29.5586 | 6.8476 | 2.2792 | 11 | 11 | 151
23 | 21.48 | 7.9403 | 7.9412 | 29.5281 | 29.5292 | 6.8487 | 2.2877 | 12 | 12 | 152
23 | 11.371 | 7.9311 | 7.9357 | 29.4236 | 29.4202 | 6.8548 | 2.3413 | 13 | 13 | 153
23 | 1.56 | 7.9511 | 7.9471 | 29.4023 | 29.4067 | 6.8527 | 2.3501 | 14 | 14 | 154
24 | 301.802 | 7.7303 | 7.7312 | 29.7421 | 29.7434 | 6.872 | 2.3072 | 1 | 1 | 155
24 | 178.806 | 7.7563 | 7.7574 | 29.6537 | 29.6549 | 6.8719 | 2.3116 | 2 | 2 | 156
24 | 150.955 | 7.768 | 7.7673 | 29.6419 | 29.6447 | 6.8706 | 2.3125 | 3 | 3 | 157
24 | 126.047 | 7.8767 | 7.879 | 29.5985 | 29.5995 | 6.8555 | 2.2728 | 4 | 4 | 158
24 | 101.145 | 7.9595 | 7.9637 | 29.5891 | 29.5899 | 6.843 | 2.2464 | 5 | 5 | 159
24 | 76.357 | 8.0157 | 8.017 | 29.5524 | 29.5535 | 6.8359 | 2.2437 | 6 | 6 | 160
24 | 51.715 | 8.0075 | 8.0085 | 29.5481 | 29.5495 | 6.8373 | 2.252 | 7 | 7 | 161
24 | 31.685 | 8.016 | 8.0178 | 29.5467 | 29.5479 | 6.8361 | 2.2475 | 8 | 8 | 162
24 | 21.282 | 8.0208 | 8.0224 | 29.5461 | 29.5471 | 6.8354 | 2.2511 | 9 | 9 | 163
24 | 11.282 | 8.0134 | 8.0164 | 29.5451 | 29.5454 | 6.8366 | 2.2554 | 10 | 10 | 164
24 | 1.658 | 8.4904 | 8.4914 | 28.5888 | 28.5874 | 6.8051 | 2.4667 | 11 | 11 | 165
35 | 58.695 | 7.8275 | 7.8284 | 29.5588 | 29.5602 | 6.865 | 2.3003 | 1 | 1 | 166
35 | 31.368 | 7.8899 | 7.8909 | 29.5053 | 29.5066 | 6.8576 | 2.3031 | 2 | 2 | 167
35 | 21.394 | 7.8938 | 7.8952 | 29.5003 | 29.5012 | 6.8572 | 2.3086 | 3 | 3 | 168
35 | 11.256 | 8.016 | 8.0135 | 29.4496 | 29.4512 | 6.8404 | 2.3412 | 4 | 4 | 169
35 | 1.942 | 8.0373 | 8.0359 | 29.3945 | 29.4017 | 6.8396 | 2.3666 | 5 | 5 | 170
36 | 56.786 | 7.8366 | 7.8384 | 29.5568 | 29.5576 | 6.8637 | 2.3002 | 1 | 1 | 171
36 | 51.247 | 7.8605 | 7.8617 | 29.5434 | 29.5447 | 6.8605 | 2.2962 | 2 | 2 | 172
36 | 31.069 | 7.8973 | 7.8987 | 29.5225 | 29.5237 | 6.8557 | 2.2999 | 3 | 3 | 173
36 | 21.249 | 7.9389 | 7.9396 | 29.4953 | 29.4968 | 6.8504 | 2.3117 | 4 | 4 | 174
36 | 11.125 | 8.0295 | 8.0315 | 29.4503 | 29.4513 | 6.8383 | 2.3665 | 5 | 5 | 175
36 | 1.791 | 8.0717 | 8.0708 | 29.3288 | 29.3334 | 6.8371 | 2.3932 | 6 | 6 | 176
39 | 65.125 | 7.8136 | 7.8146 | 29.5776 | 29.579 | 6.8663 | 2.3101 | 1 | 1 | 177
39 | 31.148 | 7.849 | 7.8497 | 29.5307 | 29.5323 | 6.8629 | 2.2993 | 2 | 2 | 178
39 | 21.183 | 7.9085 | 7.9095 | 29.4929 | 29.4943 | 6.8553 | 2.3146 | 3 | 3 | 179
39 | 11.347 | 8.0328 | 8.036 | 29.4417 | 29.4423 | 6.8381 | 2.3787 | 4 | 4 | 180
39 | 2.089 | 8.1257 | 8.1286 | 29.3643 | 29.3639 | 6.8271 | 2.4196 | 5 | 5 | 181
40 | 96.981 | 7.8146 | 7.8164 | 29.5823 | 29.5832 | 6.866 | 2.3058 | 1 | 1 | 182
40 | 76.106 | 7.8247 | 7.8258 | 29.5714 | 29.5727 | 6.8649 | 2.3075 | 2 | 2 | 183
40 | 51.258 | 7.8613 | 7.8663 | 29.5522 | 29.5515 | 6.86 | 2.3051 | 3 | 3 | 184
40 | 31.434 | 7.934 | 7.935 | 29.51 | 29.5113 | 6.8505 | 2.3031 | 4 | 4 | 185
40 | 21.468 | 7.9524 | 7.9534 | 29.4966 | 29.498 | 6.8482 | 2.3112 | 5 | 5 | 186
40 | 11.119 | 8.0513 | 8.0526 | 29.4467 | 29.4479 | 6.8351 | 2.3502 | 6 | 6 | 187
40 | 1.852 | 8.5835 | 8.5872 | 29.0969 | 29.0924 | 6.7685 | 2.4878 | 7 | 7 | 188
41 | 167.107 | 7.7583 | 7.7578 | 29.7483 | 29.755 | 6.8674 | 2.3398 | 1 | 1 | 189
41 | 101.286 | 7.8081 | 7.8091 | 29.5901 | 29.5916 | 6.8666 | 2.3131 | 2 | 2 | 190
41 | 76.15 | 7.8455 | 7.8464 | 29.5679 | 29.5694 | 6.8618 | 2.2968 | 3 | 3 | 191
41 | 51.212 | 7.908 | 7.9089 | 29.5456 | 29.5471 | 6.853 | 2.2803 | 4 | 4 | 192
41 | 31.474 | 7.9277 | 7.928 | 29.5284 | 29.5305 | 6.8507 | 2.2865 | 5 | 5 | 193
41 | 20.906 | 7.9787 | 7.9801 | 29.5009 | 29.5022 | 6.844 | 2.2999 | 6 | 6 | 194
41 | 11.139 | 8.0114 | 8.0129 | 29.4746 | 29.4757 | 6.84 | 2.323 | 7 | 7 | 195
41 | 1.902 | 8.2203 | 8.2665 | 29.3949 | 29.3784 | 6.8112 | 2.3742 | 8 | 8 | 196
42 | 223.091 | 7.7554 | 7.7565 | 29.7793 | 29.7807 | 6.8664 | 2.339 | 1 | 1 | 197
42 | 175.67 | 7.7933 | 7.7939 | 29.6869 | 29.689 | 6.8646 | 2.3326 | 2 | 2 | 198
42 | 151.163 | 7.8179 | 7.8192 | 29.6303 | 29.6314 | 6.8633 | 2.3227 | 3 | 3 | 199
42 | 125.817 | 7.8016 | 7.8027 | 29.6085 | 29.6099 | 6.8668 | 2.3126 | 4 | 4 | 200
42 | 102.328 | 7.812 | 7.8128 | 29.5822 | 29.5837 | 6.8664 | 2.2991 | 5 | 5 | 201
42 | 76.479 | 7.8758 | 7.8761 | 29.5536 | 29.555 | 6.8577 | 2.2877 | 6 | 6 | 202
42 | 50.859 | 7.9214 | 7.924 | 29.5361 | 29.5369 | 6.8513 | 2.2807 | 7 | 7 | 203
42 | 30.928 | 7.9641 | 7.9652 | 29.5303 | 29.5317 | 6.8449 | 2.2746 | 8 | 8 | 204
42 | 20.892 | 8.0042 | 8.0054 | 29.4976 | 29.4993 | 6.8401 | 2.3028 | 9 | 9 | 205
42 | 11.26 | 8.2093 | 8.1925 | 29.4254 | 29.4291 | 6.8115 | 2.3831 | 10 | 10 | 206
42 | 1.817 | 8.3351 | 8.3493 | 29.4126 | 29.4133 | 6.7927 | 2.4143 | 11 | 11 | 207
48 | 202.95 | 7.7485 | 7.7492 | 29.837 | 29.8385 | 6.8649 | 2.3569 | 1 | 1 | 208
48 | 176.361 | 7.768 | 7.7687 | 29.8101 | 29.8117 | 6.8631 | 2.3629 | 2 | 2 | 209
48 | 151.283 | 7.7737 | 7.7752 | 29.7967 | 29.797 | 6.8628 | 2.3658 | 3 | 3 | 210
48 | 126.21 | 7.776 | 7.7768 | 29.7876 | 29.7887 | 6.8628 | 2.3708 | 4 | 4 | 211
48 | 101.504 | 7.8441 | 7.8452 | 29.6472 | 29.6481 | 6.8585 | 2.3443 | 5 | 5 | 212
48 | 76.346 | 7.8607 | 7.8614 | 29.5981 | 29.5997 | 6.858 | 2.3339 | 6 | 6 | 213
48 | 51.421 | 7.8248 | 7.8255 | 29.5796 | 29.5809 | 6.8645 | 2.3195 | 7 | 7 | 214
48 | 31.251 | 7.8571 | 7.8582 | 29.568 | 29.5694 | 6.86 | 2.3158 | 8 | 8 | 215
48 | 21.14 | 7.8753 | 7.8758 | 29.5538 | 29.5553 | 6.8577 | 2.3129 | 9 | 9 | 216
48 | 10.953 | 7.9824 | 7.9851 | 29.5007 | 29.5012 | 6.8434 | 2.2939 | 10 | 10 | 217
48 | 1.881 | 8.0891 | 8.0988 | 29.3548 | 29.3437 | 6.8333 | 2.3505 | 11 | 11 | 218
50 | 154.555 | 7.7512 | 7.7518 | 29.8281 | 29.8298 | 6.8649 | 2.3694 | 1 | 1 | 219
50 | 151.519 | 7.7535 | 7.7543 | 29.825 | 29.8263 | 6.8647 | 2.3689 | 2 | 2 | 220
50 | 126.234 | 7.8151 | 7.8156 | 29.7107 | 29.7129 | 6.8601 | 2.3535 | 3 | 3 | 221
50 | 101.431 | 7.8388 | 7.8396 | 29.6558 | 29.6568 | 6.8589 | 2.3451 | 4 | 4 | 222
50 | 76.442 | 7.86 | 7.8619 | 29.5927 | 29.5957 | 6.8584 | 2.3286 | 5 | 5 | 223
50 | 51.1 | 7.8685 | 7.8689 | 29.5926 | 29.5935 | 6.8571 | 2.3398 | 6 | 6 | 224
50 | 31.345 | 7.8611 | 7.8617 | 29.5728 | 29.5746 | 6.8591 | 2.334 | 7 | 7 | 225
50 | 21.473 | 7.8634 | 7.864 | 29.5636 | 29.5651 | 6.8592 | 2.3285 | 8 | 8 | 226
50 | 11.317 | 7.8872 | 7.888 | 29.5497 | 29.5511 | 6.856 | 2.3442 | 9 | 9 | 227
50 | 2.088 | 8.0872 | 8.0833 | 29.3986 | 29.4058 | 6.8316 | 2.4248 | 10 | 10 | 228
51 | 108.899 | 7.8142 | 7.8151 | 29.7048 | 29.7065 | 6.8606 | 2.3529 | 1 | 1 | 229
51 | 100.855 | 7.8491 | 7.85 | 29.6308 | 29.6322 | 6.8584 | 2.3362 | 2 | 2 | 230
51 | 76.286 | 7.8537 | 7.8544 | 29.6144 | 29.616 | 6.8584 | 2.3384 | 3 | 3 | 231
51 | 51.046 | 7.8664 | 7.8674 | 29.5849 | 29.5855 | 6.8577 | 2.3366 | 4 | 4 | 232
51 | 31.342 | 7.8664 | 7.8667 | 29.5801 | 29.5816 | 6.858 | 2.342 | 5 | 5 | 233
51 | 21.21 | 7.8821 | 7.8806 | 29.5444 | 29.5518 | 6.8571 | 2.3444 | 6 | 6 | 234
51 | 11.287 | 8.1737 | 8.1693 | 29.3522 | 29.3564 | 6.8203 | 2.5142 | 7 | 7 | 235
51 | 1.947 | 8.4061 | 8.3964 | 29.2145 | 29.222 | 6.7905 | 2.5755 | 8 | 8 | 236
52 | 46.263 | 7.8765 | 7.8767 | 29.5577 | 29.5595 | 6.8574 | 2.3345 | 1 | 1 | 237
52 | 46.25 | 7.8764 | 7.8767 | 29.5579 | 29.5596 | 6.8574 | 2.3346 | 2 | 2 | 238
52 | 31.086 | 7.8995 | 7.9009 | 29.5367 | 29.5377 | 6.8547 | 2.3482 | 3 | 3 | 239
52 | 21.388 | 7.9566 | 7.9574 | 29.4927 | 29.4941 | 6.8478 | 2.3852 | 4 | 4 | 240
52 | 11.445 | 8.1431 | 8.1302 | 29.3838 | 29.3918 | 6.8236 | 2.4829 | 5 | 5 | 241
52 | 2.07 | 8.4492 | 8.4405 | 29.2686 | 29.2767 | 6.7815 | 2.6773 | 6 | 6 | 242
57 | 127.487 | 7.8463 | 7.8477 | 29.6223 | 29.6235 | 6.8592 | 2.3179 | 1 | 1 | 243
57 | 101.339 | 7.8573 | 7.8582 | 29.6047 | 29.6056 | 6.8583 | 2.318 | 2 | 2 | 244
57 | 76.152 | 7.8882 | 7.8894 | 29.5761 | 29.5773 | 6.8547 | 2.306 | 3 | 3 | 245
57 | 51.033 | 7.9333 | 7.9333 | 29.5609 | 29.5623 | 6.8484 | 2.2805 | 4 | 4 | 246
57 | 31.158 | 7.9824 | 7.9823 | 29.5423 | 29.5438 | 6.8415 | 2.2747 | 5 | 5 | 247
57 | 20.962 | 8.0283 | 8.0291 | 29.5106 | 29.5119 | 6.8358 | 2.2681 | 6 | 6 | 248
57 | 11.103 | 8.2333 | 8.2342 | 29.3172 | 29.3184 | 6.8126 | 2.436 | 7 | 7 | 249
57 | 1.794 | 8.4422 | 8.4458 | 29.1206 | 29.1214 | 6.7891 | 2.644 | 8 | 8 | 250
61 | 133.866 | 7.8326 | 7.8336 | 29.6881 | 29.6892 | 6.8584 | 2.3415 | 1 | 1 | 251
61 | 101.676 | 7.8549 | 7.8558 | 29.6128 | 29.614 | 6.8583 | 2.3241 | 2 | 2 | 252
61 | 76.335 | 7.8686 | 7.8709 | 29.5833 | 29.5841 | 6.8575 | 2.3028 | 3 | 3 | 253
61 | 51.179 | 7.8912 | 7.8928 | 29.5751 | 29.5759 | 6.8543 | 2.3 | 4 | 4 | 254
61 | 30.914 | 7.9386 | 7.9404 | 29.5493 | 29.5499 | 6.848 | 2.2808 | 5 | 5 | 255
61 | 21.405 | 7.9799 | 7.9814 | 29.522 | 29.5229 | 6.8428 | 2.2855 | 6 | 6 | 256
61 | 11.296 | 8.0401 | 8.0411 | 29.4689 | 29.4702 | 6.8358 | 2.3165 | 7 | 7 | 257
61 | 1.852 | 8.205 | 8.2009 | 29.3177 | 29.3251 | 6.8169 | 2.3978 | 8 | 8 | 258
66 | 175.214 | 7.8177 | 7.8185 | 29.6726 | 29.674 | 6.8614 | 2.3211 | 1 | 1 | 259
66 | 126.077 | 7.8211 | 7.822 | 29.6574 | 29.6587 | 6.8616 | 2.3264 | 2 | 2 | 260
66 | 101.315 | 7.8283 | 7.8296 | 29.6444 | 29.6455 | 6.861 | 2.3306 | 3 | 3 | 261
66 | 76.102 | 7.8268 | 7.8274 | 29.6243 | 29.6263 | 6.8622 | 2.325 | 4 | 4 | 262
66 | 50.941 | 7.8953 | 7.8972 | 29.5848 | 29.5857 | 6.8532 | 2.2905 | 5 | 5 | 263
66 | 31.187 | 8.0202 | 8.0299 | 29.5506 | 29.5482 | 6.8353 | 2.2709 | 6 | 6 | 264
66 | 21.199 | 8.0589 | 8.0598 | 29.4753 | 29.4766 | 6.8326 | 2.3042 | 7 | 7 | 265
66 | 11.125 | 8.1306 | 8.1315 | 29.393 | 29.394 | 6.8251 | 2.3489 | 8 | 8 | 266
66 | 2.047 | 8.2637 | 8.2559 | 29.0404 | 29.083 | 6.8201 | 2.4293 | 9 | 9 | 267
67 | 69.614 | 7.8771 | 7.8781 | 29.5917 | 29.593 | 6.8557 | 2.2928 | 1 | 1 | 268
67 | 51.211 | 7.8947 | 7.8953 | 29.582 | 29.5833 | 6.8534 | 2.2885 | 2 | 2 | 269
67 | 30.944 | 7.9851 | 7.9856 | 29.543 | 29.5444 | 6.8411 | 2.2803 | 3 | 3 | 270
67 | 21.028 | 8.0785 | 8.0789 | 29.4574 | 29.4593 | 6.8303 | 2.3119 | 4 | 4 | 271
67 | 10.903 | 8.1539 | 8.1554 | 29.3394 | 29.34 | 6.8239 | 2.369 | 5 | 5 | 272
67 | 1.672 | 8.2716 | 8.2691 | 29.1533 | 29.1595 | 6.8139 | 2.4236 | 6 | 6 | 273
68 | 97.14 | 7.8187 | 7.8195 | 29.6492 | 29.6509 | 6.8623 | 2.3281 | 1 | 1 | 274
68 | 51.386 | 7.9649 | 7.9611 | 29.5558 | 29.5593 | 6.8437 | 2.2771 | 2 | 2 | 275
68 | 31.375 | 8.0015 | 8.0025 | 29.5304 | 29.5316 | 6.8391 | 2.2793 | 3 | 3 | 276
68 | 21.014 | 8.0424 | 8.04 | 29.4788 | 29.4838 | 6.835 | 2.2931 | 4 | 4 | 277
68 | 11.562 | 8.1019 | 8.1016 | 29.3907 | 29.3926 | 6.8297 | 2.3314 | 5 | 5 | 278
68 | 1.892 | 8.1654 | 8.1545 | 29.3065 | 29.3252 | 6.8236 | 2.362 | 6 | 6 | 279

## bottle data
Event_Number | Station | Sample_Number | Oxygen:Dissolved | Flag:Oxygen:Dissolved | Temperature:Draw | Comment
1 | NC 01 | 1 | -99 | 5 | 8 | OXY: Dosimat ran on. Endpoint overshot. Sample lost.
1 | NC 01 | 2 | 5.6808 |  | 7.9
1 | NC 01 | 3 | 5.7525 |  | 7.9
1 | NC 01 | 4 | 5.7449 |  | 7.9
1 | NC 01 | 5 | 5.7539 |  | 7.9
1 | NC 01 | 6 | 5.7421 |  | 7.8
1 | NC 01 | 7 | 5.7484 |  | 7.8
1 | NC 01 | 8 | 5.7566 |  | 7.8
1 | NC 01 | 9 | 5.7508 | 6 | #N/A
1 | NC 01 | 10 | 5.749 |  | 7.8
1 | NC 01 | 11 | 5.8015 | 4 | 7.9 | OXY: Air bubble on lens. Transmissivity only to 93%. No clear end point.
1 | NC 01 | 12 | 5.7151 |  | 7.9
1 | NC 01 | 13 | 5.6985 |  | 7.9
1 | NC 01 | 14 | 5.6814 |  | 7.9
2 | NC 02 | 15 | 5.8292 |  | 7.9
2 | NC 02 | 16 | 5.8207 |  | 7.9
2 | NC 02 | 17 | 5.7934 | 4 | 7.9 | OXY: A poor titration with no discernable endpoint. That it is not terribly out of line is a fluke.
2 | NC 02 | 18 | 5.7954 |  | 7.9
2 | NC 02 | 19 | 5.7875 |  | 7.9
2 | NC 02 | 20 | 5.7846 |  | 7.9
2 | NC 02 | 21 | 5.7718 |  | 7.9
2 | NC 02 | 22 | 5.7606 |  | 7.9
2 | NC 02 | 23 | 5.7345 |  | 8
2 | NC 02 | 24 | 5.6758 | 6 | #N/A
2 | NC 02 | 25 | 5.6381 | 3 | 8 | OXY: Air bubble on lens. Transmissivity only to 96%.
2 | NC 02 | 26 | 5.5962 |  | 8
2 | NC 02 | 27 | 5.7423 |  | 8
3 | NC 03 | 28 | 6.3441 |  | 8.1
3 | NC 03 | 29 | 5.7771 |  | 8
3 | NC 03 | 30 | 5.768 |  | 8 | OXY: Duplicate value not used. See notes.
3 | NC 03 | 31 | -99 | 5 | 8 | OXY: Air bubble on lens. Transmissivity only to 81%. Titration incomplete. Out of line in depth profile.
3 | NC 03 | 32 | 5.7729 |  | 8
3 | NC 03 | 33 | 5.7676 |  | 8
3 | NC 03 | 34 | 5.7519 |  | 8
3 | NC 03 | 35 | 5.6729 |  | 8
3 | NC 03 | 36 | 5.479 |  | 8.1
3 | NC 03 | 37 | 5.4938 |  | 8.2
3 | NC 03 | 38 | 5.5718 |  | 8.2
3 | NC 03 | 39 | 5.9546 |  | 8.2
4 | NC 4 | 40 | 5.8062 |  | 8.2
4 | NC 4 | 41 | 5.77 |  | 8.1
4 | NC 4 | 42 | 5.7775 |  | 8.1
4 | NC 4 | 43 | 5.7586 |  | 8.1
4 | NC 4 | 44 | 5.7383 |  | 8.1
4 | NC 4 | 45 | 5.7222 |  | 8.1
4 | NC 4 | 46 | 5.6059 | 6 | #N/A
4 | NC 4 | 47 | 5.5347 |  | 8.2
4 | NC 4 | 48 | 5.47 |  | 8.3
4 | NC 4 | 49 | 5.5525 |  | 8.3
4 | NC 4 | 50 | 5.6978 |  | 8.3
4 | NC 4 | 51 | 6.07 |  | 8.4
9 | NC 5 | 52 | 5.7718 |  | 8.2
9 | NC 5 | 53 | 5.7666 |  | 8.1
9 | NC 5 | 54 | 5.7665 |  | 8.1
9 | NC 5 | 55 | 5.7497 |  | 8.1
9 | NC 5 | 56 | 5.706 |  | 8.1
9 | NC 5 | 57 | 5.663 |  | 8.1
9 | NC 5 | 58 | 5.5129 |  | 8.2
9 | NC 5 | 59 | 5.4762 |  | 8.2
9 | NC 5 | 60 | 5.4179 |  | 8.2
9 | NC 5 | 61 | 5.4157 | 6 | #N/A
9 | NC 5 | 62 | 5.504 |  | 8.2
9 | NC 5 | 63 | 5.6656 |  | 8.2
10 | NC 6 | 64 | 5.7851 |  | 8.1
10 | NC 6 | 65 | 5.7756 | 36 | #N/A | OXY: Both duplicate endpoints ambiguous. Both looked to be late but adjusted
10 | NC 6 | 66 | 5.7785 |  | 8
10 | NC 6 | 67 | 5.7482 |  | 8.1
10 | NC 6 | 68 | 5.7301 |  | 8.1
10 | NC 6 | 69 | 5.6964 |  | 8.1
10 | NC 6 | 70 | 5.6197 |  | 8.1
10 | NC 6 | 71 | 5.5052 |  | 8.2
10 | NC 6 | 72 | 5.4079 |  | 8.3
10 | NC 6 | 73 | 5.4187 |  | 8.3
10 | NC 6 | 74 | 5.5359 |  | 8.3
10 | NC 6 | 75 | 5.6275 |  | 8.3
11 | NC 7 | 76 | 5.7885 |  | 8.1
11 | NC 7 | 77 | 5.7675 | 46 | #N/A | OXY: Replicate outliers according to the Chauvenet criteria. Duplicate titrations were both quite problematic.
11 | NC 7 | 78 | 5.7638 |  | 8.1
11 | NC 7 | 79 | 5.7431 |  | 8.1
11 | NC 7 | 80 | 5.7103 | 3 | 8.1 | OXY: Wavy, ugly curve. Endpoint ambiguous.
11 | NC 7 | 81 | 5.6253 |  | 8.1
11 | NC 7 | 82 | 5.6226 | 3 | 8.2 | OXY: Messy curve. What seems to be the best endpoint was chosen by AutoOxy. Seems out of line in the depth profile.
11 | NC 7 | 83 | 5.6156 |  | 8.1
11 | NC 7 | 84 | 5.3969 |  | 8.3
11 | NC 7 | 85 | 5.4542 |  | 8.3
11 | NC 7 | 86 | 5.488 |  | 8.3
11 | NC 7 | 87 | 5.5438 |  | 8.3
12 | NC 8 | 88 | 5.7718 |  | 8.2
12 | NC 8 | 89 | 5.727 | 6 | #N/A
12 | NC 8 | 90 | 5.6722 |  | 8.2
12 | NC 8 | 91 | 5.6638 |  | 8.2
12 | NC 8 | 92 | 5.6488 |  | 8.2
12 | NC 8 | 93 | 5.6261 |  | 8.2
12 | NC 8 | 94 | 5.484 |  | 8.3
12 | NC 8 | 95 | 5.4497 |  | 8.3
12 | NC 8 | 96 | 5.3796 |  | 8.3
12 | NC 8 | 97 | 5.3615 |  | 8.3
12 | NC 8 | 98 | 5.4881 |  | 8.3
12 | NC 8 | 99 | 5.6436 |  | 8.3
13 | NC 9 | 100 | 5.7774 |  | 8.3
13 | NC 9 | 101 | 5.7745 |  | 8.2
13 | NC 9 | 102 | 5.7095 | 6 | #N/A
13 | NC 9 | 103 | 5.6771 |  | 8.2
13 | NC 9 | 104 | 5.6642 |  | 8.2
13 | NC 9 | 105 | 5.6413 |  | 8.2
13 | NC 9 | 106 | 5.5694 |  | 8.2
13 | NC 9 | 107 | 5.5235 |  | 8.2
13 | NC 9 | 108 | 5.4293 | 3 | 8.3 | OXY: End point was unclear. No good inflection point but seemed very late. Adjusted.
13 | NC 9 | 109 | 5.424 |  | 8.3
13 | NC 9 | 110 | 5.4242 |  | 8.3
13 | NC 9 | 111 | 5.4774 |  | 8.3
13 | NC 9 | 112 | 5.617 |  | 8.3
19 | NC 10 | 113 | 5.7566 |  | 8.1
19 | NC 10 | 114 | 5.7187 |  | 8
19 | NC 10 | 115 | 5.6826 |  | 8.2
19 | NC 10 | 116 | 5.6667 | 4 | 8 | OXY: Titration incomplete
19 | NC 10 | 117 | 5.6557 |  | 8
19 | NC 10 | 118 | 5.5945 |  | 8
19 | NC 10 | 119 | 5.5192 |  | 8.1
19 | NC 10 | 120 | 5.4606 | 6 | #N/A
19 | NC 10 | 121 | 5.4668 |  | 8.1
19 | NC 10 | 122 | 5.4245 | 4 | 8.1 | OXY: A very messy titration with no clear endpoint
19 | NC 10 | 123 | 5.416 |  | 8.2
19 | NC 10 | 124 | 5.4311 |  | 8.2
19 | NC 10 | 125 | 5.4395 |  | 8.2
20 | NC 11 | 126 | 5.7493 |  | 8.1
20 | NC 11 | 127 | 5.7545 |  | 8
20 | NC 11 | 128 | 5.7533 |  | 8
20 | NC 11 | 129 | -99 | 5 | 8 | OXY: Bubble on lens. Only seven points plotted in the titration. No discernable end point. A clear outlier in the depth profile
20 | NC 11 | 130 | 5.7188 |  | 8
20 | NC 11 | 131 | 5.6896 |  | 8
20 | NC 11 | 132 | 5.6537 |  | 8
20 | NC 11 | 133 | 5.6165 |  | 8
20 | NC 11 | 134 | 5.51 |  | 8.1
20 | NC 11 | 135 | 5.4514 | 6 | #N/A
20 | NC 11 | 136 | 5.3965 |  | 8.2
20 | NC 11 | 137 | 5.4223 |  | 8.2
20 | NC 11 | 138 | 5.4746 |  | 8.2
20 | NC 11 | 139 | 5.5526 |  | 8.2
20 | NC 11 | 140 | 5.7386 |  | 8.2
23 | NC 12 | 141 | 5.7685 |  | 8.1
23 | NC 12 | 142 | 5.7582 |  | 8.1
23 | NC 12 | 143 | 5.7355 |  | 8.1
23 | NC 12 | 144 | 5.7119 |  | 8.1
23 | NC 12 | 145 | 5.6944 |  | 8.1
23 | NC 12 | 146 | 5.6303 |  | 8.1
23 | NC 12 | 147 | 5.6266 |  | 8.1 | OXY: Duplicate value not used. See notes.
23 | NC 12 | 148 | 5.5132 |  | 8.1
23 | NC 12 | 149 | 5.4525 |  | 8.2
23 | NC 12 | 150 | 5.4503 |  | 8.2
23 | NC 12 | 151 | 5.4312 |  | 8.2
23 | NC 12 | 152 | 5.4516 |  | 8.2
23 | NC 12 | 153 | 5.5627 |  | 8.2
23 | NC 12 | 154 | 5.6318 |  | 8.2
24 | NC 13 | 155 | 5.7731 | 3 | 8.2 | OXY: Curve does an odd jump up near the end point, making it difficult to clearly identify
24 | NC 13 | 156 | 5.6711 |  | 8.1
24 | NC 13 | 157 | 5.6658 |  | 8.1
24 | NC 13 | 158 | 5.516 | 6 | #N/A
24 | NC 13 | 159 | 5.3974 |  | 8.2
24 | NC 13 | 160 | 5.3371 |  | 8.3
24 | NC 13 | 161 | 5.3592 |  | 8.3
24 | NC 13 | 162 | 5.3469 |  | 8.3
24 | NC 13 | 163 | 5.33 |  | 8.3
24 | NC 13 | 164 | 5.3346 |  | 8.3
24 | NC 13 | 165 | 5.8381 |  | 8.6
35 | NC 23 | 166 | 5.5354 |  | 8.2
35 | NC 23 | 167 | 5.5187 |  | 8.2
35 | NC 23 | 168 | 5.5294 |  | 8.2
35 | NC 23 | 169 | 5.5843 |  | 8.3
35 | NC 23 | 170 | 5.7231 |  | 8.3
36 | NC 24 | 171 | 5.5375 |  | 8.3
36 | NC 24 | 172 | 5.5224 |  | 8.2
36 | NC 24 | 173 | 5.5011 |  | 8.2
36 | NC 24 | 174 | 5.5299 |  | 8.3
36 | NC 24 | 175 | 5.6136 |  | 8.3
36 | NC 24 | 176 | 5.7632 |  | 8.3
39 | NC 14 | 177 | 5.5669 |  | 8.2
39 | NC 14 | 178 | 5.5063 |  | 8.2
39 | NC 14 | 179 | 5.5249 |  | 8.1
39 | NC 14 | 180 | 5.6077 |  | 8.1
39 | NC 14 | 181 | 5.7942 |  | 8.3
40 | NC 15 | 182 | 5.5841 |  | 8.1
40 | NC 15 | 183 | 5.5728 |  | 8.1
40 | NC 15 | 184 | 5.5477 |  | 8.1
40 | NC 15 | 185 | 5.5077 |  | 8.1
40 | NC 15 | 186 | 5.518 |  | 8.2
40 | NC 15 | 187 | 5.617 |  | 8.3
40 | NC 15 | 188 | 5.7684 |  | 8.5
41 | NC 16 | 189 | 5.7245 |  | 8.1
41 | NC 16 | 190 | 5.6136 |  | 8.2
41 | NC 16 | 191 | 5.5236 |  | 8.1
41 | NC 16 | 192 | 5.4615 |  | 8.1
41 | NC 16 | 193 | 5.4548 |  | 8.1
41 | NC 16 | 194 | 5.4794 |  | 8.1
41 | NC 16 | 195 | 5.5228 |  | 8.1
41 | NC 16 | 196 | 5.619 |  | 8.3
42 | NC 17 | 197 | 5.4086 | 4 | 8.1 | OXY: Titration completed at 75.5%. Endpoint accuracy questionable. Value seems low in depth profile
42 | NC 17 | 198 | 5.4936 | 4 | 8.1 | OXY: Titrationonly up to 94%. No end point. Looks to have been something blocking the light path
42 | NC 17 | 199 | 5.6661 | 6 | #N/A
42 | NC 17 | 200 | 5.613 |  | 8
42 | NC 17 | 201 | 5.5425 | 4 | 8.1 | OXY: Titration incomplete. No end point.
42 | NC 17 | 202 | 5.4235 | 4 | 8.1 | OXY: Titration incomplete. No end point.
42 | NC 17 | 203 | 5.4539 |  | 8.1
42 | NC 17 | 204 | 5.4137 |  | 8.1
42 | NC 17 | 205 | 5.4543 |  | 8.2
42 | NC 17 | 206 | 5.7226 |  | 8.4
42 | NC 17 | 207 | 5.786 |  | 8.5
48 | NC 19 | 208 | 5.8654 |  | 8
48 | NC 19 | 209 | 5.8289 |  | 8
48 | NC 19 | 210 | 5.8116 |  | 8
48 | NC 19 | 211 | 5.8081 |  | 8
48 | NC 19 | 212 | 5.7213 |  | 8.1
48 | NC 19 | 213 | 5.6368 |  | 8.1
48 | NC 19 | 214 | 5.5831 |  | 8.1
48 | NC 19 | 215 | 5.5584 | 6 | #N/A
48 | NC 19 | 216 | 5.534 |  | 8.1
48 | NC 19 | 217 | 5.4833 |  | 8.1
48 | NC 19 | 218 | 5.5903 |  | 8.3
50 | NC 20 | 219 | 5.8235 |  | 8.1
50 | NC 20 | 220 | 5.8191 |  | 8
50 | NC 20 | 221 | 5.7719 |  | 8.1
50 | NC 20 | 222 | 5.7123 | 6 | #N/A
50 | NC 20 | 223 | 5.6689 |  | 8.1
50 | NC 20 | 224 | 5.6268 | 3 | 8.1 | OXY: Endpoint ambiguous but it looked to be quite late. Adjusted
50 | NC 20 | 225 | 5.6171 |  | 8.1
50 | NC 20 | 226 | 5.5855 |  | 8.1
50 | NC 20 | 227 | 5.6181 |  | 8.1
50 | NC 20 | 228 | 5.8337 |  | 8.2
51 | NC 21 | 229 | 5.7069 |  | 8.1
51 | NC 21 | 230 | 5.6993 |  | 8.1
51 | NC 21 | 231 | 5.6615 |  | 8.2
51 | NC 21 | 232 | 5.6386 |  | 8.2
51 | NC 21 | 233 | 5.6364 |  | 8.2
51 | NC 21 | 234 | 5.6367 |  | 8.2
51 | NC 21 | 235 | 5.8715 |  | 8.3
51 | NC 21 | 236 | 6.395 |  | 8.6
52 | NC 22 | 237 | 5.6758 |  | 8.3
52 | NC 22 | 238 | 5.6343 |  | 8.2
52 | NC 22 | 239 | 5.6364 |  | 8.2
52 | NC 22 | 240 | 5.7046 |  | 8.2
52 | NC 22 | 241 | 5.8904 |  | 8.3
52 | NC 22 | 242 | 6.5236 |  | 8.6
57 | NC 25 | 243 | 5.6287 |  | 8.1
57 | NC 25 | 244 | 5.6067 |  | 8.1
57 | NC 25 | 245 | 5.557 |  | 8.1
57 | NC 25 | 246 | 5.4576 |  | 8.1
57 | NC 25 | 247 | 5.4167 |  | 8.1
57 | NC 25 | 248 | 5.4034 |  | 8.2
57 | NC 25 | 249 | 5.7459 |  | 8.3
57 | NC 25 | 250 | 6.1636 |  | 8.5
61 | NC 26 | 251 | 5.7189 |  | 8.1
61 | NC 26 | 252 | 5.6343 |  | 8.1
61 | NC 26 | 253 | 5.542 |  | 8.1
61 | NC 26 | 254 | 5.516 |  | 8.1
61 | NC 26 | 255 | 5.4414 |  | 8.2
61 | NC 26 | 256 | 5.4288 |  | 8.2
61 | NC 26 | 257 | 5.4909 |  | 8.2
61 | NC 26 | 258 | 5.7585 |  | 8.4
66 | NC 27 | 259 | 5.6964 |  | 8.3
66 | NC 27 | 260 | 5.6677 |  | 8.3
66 | NC 27 | 261 | 5.6562 |  | 8.3
66 | NC 27 | 262 | 5.6364 |  | 8.3
66 | NC 27 | 263 | 5.4969 |  | 8.3
66 | NC 27 | 264 | 5.4095 |  | 8.4
66 | NC 27 | 265 | 5.4605 |  | 8.4
66 | NC 27 | 266 | 5.5693 |  | 8.5
66 | NC 27 | 267 | 5.8507 |  | 8.6
67 | NC 28 | 268 | 5.4446 | 3 | -99 | OXY: Seems low. Not accurate. Add 1 + 2 titer. Zeroed colorimeter in sample. No end point
67 | NC 28 | 269 | 5.49 |  | 8.4
67 | NC 28 | 270 | 5.429 |  | 8.4
67 | NC 28 | 271 | 5.4661 |  | 8.5
67 | NC 28 | 272 | 5.6659 |  | 8.6
67 | NC 28 | 273 | 5.7779 |  | 8.6
68 | NC 29 | 274 | 5.6025 |  | 8.4
68 | NC 29 | 275 | 5.4388 |  | 8.4
68 | NC 29 | 276 | 5.4295 |  | 8.5
68 | NC 29 | 277 | 5.4544 |  | 8.5
68 | NC 29 | 278 | 5.3149 |  | 8.5
68 | NC 29 | 279 | 5.646 |  | 8.6

## template
LOC:EVENT_NUMBER | Pressure | Temperature:Primary | Temperature:Secondary | Salinity:T0:C0 | Salinity:T1:C1 | Oxygen:saturation' | Oxygen:Voltage:SBE | Bottle:Position | Bottle_Number | Sample_Number | Phi | DObot/Phi | Oxygen:Dissolved | Flag:Oxygen:Dissolved | diff from fit | 1st fit m | 1st fit b | Titrated DO | Titrated DO Flag | Sample # | Diff in samp #
[23 empty] | m | b | Soc | Voffset
|  |  | This is a bit finicky, so useful to save at different stages in case a CUT/PASTE or SORT step goes badly | [18 empty] | Original
|  |  | This is a template with formulae for Phi, DO bot/Phi, Diff from fit, Soc, Voffset | [18 empty] | all data |  |  | 0 | #DIV/0!
|  |  | Substitute appropriate values for A, B, C in the Phi (col.L) formula, then save this sheet with new filename | [18 empty] | all data - outliers |  |  | 0 | #DIV/0!
Cut and paste values from the CTD bottle spreadsheet, then add DO, DO Flag and Sample #s from oxygen titration spreadsheet
Use "Diff in samp #" to realign data so sample #s match and get rid of cases where there is no sample, or no bottle data to match - then get rid of extra sample # and copy DO and flag data to columns N & O
Do a fit of column M to column H
When the first fit is done with all values, fill in the m and b values in columns Q and R - same value for every row (could base Q/R on a fit after worst outliers removed)
Then you can order on the P column to remove outliers
Plots of difference versus event # or pressure or sample # can establish if there is time-dependence or pressure dependence and identify major outliers