## Notes
Analysis: | Hugh MacLean, SOG onboard CCGS Vector, 10-APRIL-12 to 18-APRIL-12 using IOS Kit #1. Data processing by Kenny Scozzafava
Chief Scientist: | Terri Sutherland
2012-04-25 |  | KS
Standards were only analyzed twice for the cruise (~ 3 days apart). Although there was no significant change in thio concentration, I would have
liked to have seen one more standard run near the end of the cruise just for confirmation that the system was stable and accurate.
Very many titrations were quite noisy and wavy. Characteristic of bubbles swirling around in the flask. The analyst likely did not take care to remove bubbles before
the titration began.
Many casts were sampled with no duplicates. The requirement is that at least one duplicate be sampled per cast. 16 duplicates ÷ 279 samples = only 6%
#17 - A poor titration with no discernable endpoint. It seems characteristic of a bubble in the light path but no notes provided by the analyst. Please check
2012-04-30 |  | KS
#30- Was not included in the duplicate analysis due to known problems with the analysis. Dup 2: Air bubble on lens. Transmissivity only to 90%. Titration incomplete.
Out of line in depth profile as well. Only dup 1 is reported on the QF worksheet
#77 - Duplicate titrations were both problematic and flagged as a 4. They were flagged for rejection by the Chauvenet criteria.
#147 - Was not included in the duplicate analysis due to known problems with the analysis. Dup 1: End point failure at 85%. Value way off. Confirmed in
titration plot, depth profile and with duplicate sample value. Only Dup 2 is reported in the QF worksheet.
Cast 42 - Four samples were problematic. The profile seems funny looking to me but its difficult to say without the comparison. See plots worksheet.
2012-05-04 |  | KS
#1, #30 dup 2, #31, #129, #147 dup 1 - were lost due to various problems with the program and Dosimat and failure to take care of bubbles in the sample during the titrations
2012-06-29 |  | KS
Sample 11 - Added to comment that value is a slight outlier in COMPARE. CTD data was noisy therefore not much weight can be placed on the comparison
Sample 30 - I missed flagging the value as per the new flag clarifications put into place recently. Flagged as a 2.
Sample 82 - Temperature and salinity profiles similar to DO therefore the value being out of line is not unique. I amended the comment accordingly but left the flag as a 3
due to the problems associated with the titration.
Sample 197 - Added to comment that value is out in COMPARE.
Sample 198 - Added to comment that value is out in COMPARE.
Sample 202 - Added to comment that value is out in COMPARE.
Sample 268 - Added to comment that value is out in COMPARE.

## QF 2012-15 OXY FINAL
HEADER: | Oxygen samples were analyzed at sea using an automated Winkler titration system
(Brinkmann Dosimat model 765 and a PC 950 Colorimeter controlled by AutoOxy software V3.4)
with modifications based on Carpenter (1965) and adhering to WOCE protocols (Culberson 1991). For details including a duplicate analysis, see document QF2012-15OXY.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Oxygen:Dissolved | Flag:Oxygen:Dissolved | Temperature:Draw | Comment
2012-15 | 1 | NC 01 | 1 | -99 | 5 | 8 | OXY: Dosimat ran on. Endpoint overshot. Sample lost.
2012-15 | 1 | NC 01 | 2 | 5.6808 |  | 7.9
2012-15 | 1 | NC 01 | 3 | 5.7525 |  | 7.9
2012-15 | 1 | NC 01 | 4 | 5.7449 |  | 7.9
2012-15 | 1 | NC 01 | 5 | 5.7539 |  | 7.9
2012-15 | 1 | NC 01 | 6 | 5.7421 |  | 7.8
2012-15 | 1 | NC 01 | 7 | 5.7484 |  | 7.8
2012-15 | 1 | NC 01 | 8 | 5.7566 |  | 7.8
2012-15 | 1 | NC 01 | 9 | 5.7508 | 6 | 7.8
2012-15 | 1 | NC 01 | 10 | 5.749 |  | 7.8
2012-15 | 1 | NC 01 | 11 | 5.8015 | 4 | 7.9 | OXY: Air bubble on lens. Transmissivity only to 93%. No clear end point. Slight outlier in COMPARE (by 0.05 mL/L)
2012-15 | 1 | NC 01 | 12 | 5.7151 |  | 7.9
2012-15 | 1 | NC 01 | 13 | 5.6985 |  | 7.9
2012-15 | 1 | NC 01 | 14 | 5.6814 |  | 7.9
2012-15 | 2 | NC 02 | 15 | 5.8292 |  | 7.9
2012-15 | 2 | NC 02 | 16 | 5.8207 |  | 7.9
2012-15 | 2 | NC 02 | 17 | 5.7934 | 4 | 7.9 | OXY: A poor titration with no discernable endpoint. That it is not terribly out of line is a fluke.
2012-15 | 2 | NC 02 | 18 | 5.7954 |  | 7.9
2012-15 | 2 | NC 02 | 19 | 5.7875 |  | 7.9
2012-15 | 2 | NC 02 | 20 | 5.7846 |  | 7.9
2012-15 | 2 | NC 02 | 21 | 5.7718 |  | 7.9
2012-15 | 2 | NC 02 | 22 | 5.7606 |  | 7.9
2012-15 | 2 | NC 02 | 23 | 5.7345 |  | 8
2012-15 | 2 | NC 02 | 24 | 5.6758 | 6 | 8
2012-15 | 2 | NC 02 | 25 | 5.6381 | 3 | 8 | OXY: Air bubble on lens. Transmissivity only to 96%.
2012-15 | 2 | NC 02 | 26 | 5.5962 |  | 8
2012-15 | 2 | NC 02 | 27 | 5.7423 |  | 8
2012-15 | 3 | NC 03 | 28 | 6.3441 |  | 8.1
2012-15 | 3 | NC 03 | 29 | 5.7771 |  | 8
2012-15 | 3 | NC 03 | 30 | 5.768 | 2 | 8 | OXY: Duplicate value not used. See notes.
2012-15 | 3 | NC 03 | 31 | -99 | 5 | 8 | OXY: Air bubble on lens. Transmissivity only to 81%. Titration incomplete. Out of line in depth profile.
2012-15 | 3 | NC 03 | 32 | 5.7729 |  | 8
2012-15 | 3 | NC 03 | 33 | 5.7676 |  | 8
2012-15 | 3 | NC 03 | 34 | 5.7519 |  | 8
2012-15 | 3 | NC 03 | 35 | 5.6729 |  | 8
2012-15 | 3 | NC 03 | 36 | 5.479 |  | 8.1
2012-15 | 3 | NC 03 | 37 | 5.4938 |  | 8.2
2012-15 | 3 | NC 03 | 38 | 5.5718 |  | 8.2
2012-15 | 3 | NC 03 | 39 | 5.9546 |  | 8.2
2012-15 | 4 | NC 4 | 40 | 5.8062 |  | 8.2
2012-15 | 4 | NC 4 | 41 | 5.77 |  | 8.1
2012-15 | 4 | NC 4 | 42 | 5.7775 |  | 8.1
2012-15 | 4 | NC 4 | 43 | 5.7586 |  | 8.1
2012-15 | 4 | NC 4 | 44 | 5.7383 |  | 8.1
2012-15 | 4 | NC 4 | 45 | 5.7222 |  | 8.1
2012-15 | 4 | NC 4 | 46 | 5.6059 | 6 | 8.2
2012-15 | 4 | NC 4 | 47 | 5.5347 |  | 8.2
2012-15 | 4 | NC 4 | 48 | 5.47 |  | 8.3
2012-15 | 4 | NC 4 | 49 | 5.5525 |  | 8.3
2012-15 | 4 | NC 4 | 50 | 5.6978 |  | 8.3
2012-15 | 4 | NC 4 | 51 | 6.07 |  | 8.4
2012-15 | 9 | NC 5 | 52 | 5.7718 |  | 8.2
2012-15 | 9 | NC 5 | 53 | 5.7666 |  | 8.1
2012-15 | 9 | NC 5 | 54 | 5.7665 |  | 8.1
2012-15 | 9 | NC 5 | 55 | 5.7497 |  | 8.1
2012-15 | 9 | NC 5 | 56 | 5.706 |  | 8.1
2012-15 | 9 | NC 5 | 57 | 5.663 |  | 8.1
2012-15 | 9 | NC 5 | 58 | 5.5129 |  | 8.2
2012-15 | 9 | NC 5 | 59 | 5.4762 |  | 8.2
2012-15 | 9 | NC 5 | 60 | 5.4179 |  | 8.2
2012-15 | 9 | NC 5 | 61 | 5.4157 | 6 | 8.2
2012-15 | 9 | NC 5 | 62 | 5.504 |  | 8.2
2012-15 | 9 | NC 5 | 63 | 5.6656 |  | 8.2
2012-15 | 10 | NC 6 | 64 | 5.7851 |  | 8.1
2012-15 | 10 | NC 6 | 65 | 5.7756 | 36 | 8.1 | OXY: Both duplicate endpoints ambiguous. Both looked to be late but adjusted
2012-15 | 10 | NC 6 | 66 | 5.7785 |  | 8
2012-15 | 10 | NC 6 | 67 | 5.7482 |  | 8.1
2012-15 | 10 | NC 6 | 68 | 5.7301 |  | 8.1
2012-15 | 10 | NC 6 | 69 | 5.6964 |  | 8.1
2012-15 | 10 | NC 6 | 70 | 5.6197 |  | 8.1
2012-15 | 10 | NC 6 | 71 | 5.5052 |  | 8.2
2012-15 | 10 | NC 6 | 72 | 5.4079 |  | 8.3
2012-15 | 10 | NC 6 | 73 | 5.4187 |  | 8.3
2012-15 | 10 | NC 6 | 74 | 5.5359 |  | 8.3
2012-15 | 10 | NC 6 | 75 | 5.6275 |  | 8.3
2012-15 | 11 | NC 7 | 76 | 5.7885 |  | 8.1
2012-15 | 11 | NC 7 | 77 | 5.7675 | 46 | 8.1 | OXY: Replicate outliers according to the Chauvenet criteria. Duplicate titrations were both quite problematic.
2012-15 | 11 | NC 7 | 78 | 5.7638 |  | 8.1
2012-15 | 11 | NC 7 | 79 | 5.7431 |  | 8.1
2012-15 | 11 | NC 7 | 80 | 5.7103 | 3 | 8.1 | OXY: Wavy, ugly curve. Endpoint ambiguous.
2012-15 | 11 | NC 7 | 81 | 5.6253 |  | 8.1
2012-15 | 11 | NC 7 | 82 | 5.6226 | 3 | 8.2 | OXY: Messy curve. What seems to be the best endpoint was chosen by AutoOxy.
2012-15 | 11 | NC 7 | 83 | 5.6156 |  | 8.1
2012-15 | 11 | NC 7 | 84 | 5.3969 |  | 8.3
2012-15 | 11 | NC 7 | 85 | 5.4542 |  | 8.3
2012-15 | 11 | NC 7 | 86 | 5.488 |  | 8.3
2012-15 | 11 | NC 7 | 87 | 5.5438 |  | 8.3
2012-15 | 12 | NC 8 | 88 | 5.7718 |  | 8.2
2012-15 | 12 | NC 8 | 89 | 5.727 | 6 | 8.2
2012-15 | 12 | NC 8 | 90 | 5.6722 |  | 8.2
2012-15 | 12 | NC 8 | 91 | 5.6638 |  | 8.2
2012-15 | 12 | NC 8 | 92 | 5.6488 |  | 8.2
2012-15 | 12 | NC 8 | 93 | 5.6261 |  | 8.2
2012-15 | 12 | NC 8 | 94 | 5.484 |  | 8.3
2012-15 | 12 | NC 8 | 95 | 5.4497 |  | 8.3
2012-15 | 12 | NC 8 | 96 | 5.3796 |  | 8.3
2012-15 | 12 | NC 8 | 97 | 5.3615 |  | 8.3
2012-15 | 12 | NC 8 | 98 | 5.4881 |  | 8.3
2012-15 | 12 | NC 8 | 99 | 5.6436 |  | 8.3
2012-15 | 13 | NC 9 | 100 | 5.7774 |  | 8.3
2012-15 | 13 | NC 9 | 101 | 5.7745 |  | 8.2
2012-15 | 13 | NC 9 | 102 | 5.7095 | 6 | 8.2
2012-15 | 13 | NC 9 | 103 | 5.6771 |  | 8.2
2012-15 | 13 | NC 9 | 104 | 5.6642 |  | 8.2
2012-15 | 13 | NC 9 | 105 | 5.6413 |  | 8.2
2012-15 | 13 | NC 9 | 106 | 5.5694 |  | 8.2
2012-15 | 13 | NC 9 | 107 | 5.5235 |  | 8.2
2012-15 | 13 | NC 9 | 108 | 5.4293 | 3 | 8.3 | OXY: End point was unclear. No good inflection point but seemed very late. Adjusted.
2012-15 | 13 | NC 9 | 109 | 5.424 |  | 8.3
2012-15 | 13 | NC 9 | 110 | 5.4242 |  | 8.3
2012-15 | 13 | NC 9 | 111 | 5.4774 |  | 8.3
2012-15 | 13 | NC 9 | 112 | 5.617 |  | 8.3
2012-15 | 19 | NC 10 | 113 | 5.7566 |  | 8.1
2012-15 | 19 | NC 10 | 114 | 5.7187 |  | 8
2012-15 | 19 | NC 10 | 115 | 5.6826 |  | 8.2
2012-15 | 19 | NC 10 | 116 | 5.6667 | 4 | 8 | OXY: Titration incomplete
2012-15 | 19 | NC 10 | 117 | 5.6557 |  | 8
2012-15 | 19 | NC 10 | 118 | 5.5945 |  | 8
2012-15 | 19 | NC 10 | 119 | 5.5192 |  | 8.1
2012-15 | 19 | NC 10 | 120 | 5.4606 | 6 | 8.1
2012-15 | 19 | NC 10 | 121 | 5.4668 |  | 8.1
2012-15 | 19 | NC 10 | 122 | 5.4245 | 4 | 8.1 | OXY: A very messy titration with no clear endpoint
2012-15 | 19 | NC 10 | 123 | 5.416 |  | 8.2
2012-15 | 19 | NC 10 | 124 | 5.4311 |  | 8.2
2012-15 | 19 | NC 10 | 125 | 5.4395 |  | 8.2
2012-15 | 20 | NC 11 | 126 | 5.7493 |  | 8.1
2012-15 | 20 | NC 11 | 127 | 5.7545 |  | 8
2012-15 | 20 | NC 11 | 128 | 5.7533 |  | 8
2012-15 | 20 | NC 11 | 129 | -99 | 5 | 8 | OXY: Bubble on lens. Only seven points plotted in the titration. No discernable end point. A clear outlier in the depth profile
2012-15 | 20 | NC 11 | 130 | 5.7188 |  | 8
2012-15 | 20 | NC 11 | 131 | 5.6896 |  | 8
2012-15 | 20 | NC 11 | 132 | 5.6537 |  | 8
2012-15 | 20 | NC 11 | 133 | 5.6165 |  | 8
2012-15 | 20 | NC 11 | 134 | 5.51 |  | 8.1
2012-15 | 20 | NC 11 | 135 | 5.4514 | 6 | 8.1
2012-15 | 20 | NC 11 | 136 | 5.3965 |  | 8.2
2012-15 | 20 | NC 11 | 137 | 5.4223 |  | 8.2
2012-15 | 20 | NC 11 | 138 | 5.4746 |  | 8.2
2012-15 | 20 | NC 11 | 139 | 5.5526 |  | 8.2
2012-15 | 20 | NC 11 | 140 | 5.7386 |  | 8.2
2012-15 | 23 | NC 12 | 141 | 5.7685 |  | 8.1
2012-15 | 23 | NC 12 | 142 | 5.7582 |  | 8.1
2012-15 | 23 | NC 12 | 143 | 5.7355 |  | 8.1
2012-15 | 23 | NC 12 | 144 | 5.7119 |  | 8.1
2012-15 | 23 | NC 12 | 145 | 5.6944 |  | 8.1
2012-15 | 23 | NC 12 | 146 | 5.6303 |  | 8.1
2012-15 | 23 | NC 12 | 147 | 5.6266 |  | 8.1 | OXY: Duplicate value not used. See notes.
2012-15 | 23 | NC 12 | 148 | 5.5132 |  | 8.1
2012-15 | 23 | NC 12 | 149 | 5.4525 |  | 8.2
2012-15 | 23 | NC 12 | 150 | 5.4503 |  | 8.2
2012-15 | 23 | NC 12 | 151 | 5.4312 |  | 8.2
2012-15 | 23 | NC 12 | 152 | 5.4516 |  | 8.2
2012-15 | 23 | NC 12 | 153 | 5.5627 |  | 8.2
2012-15 | 23 | NC 12 | 154 | 5.6318 |  | 8.2
2012-15 | 24 | NC 13 | 155 | 5.7731 | 3 | 8.2 | OXY: Curve does an odd jump up near the end point, making it difficult to clearly identify
2012-15 | 24 | NC 13 | 156 | 5.6711 |  | 8.1
2012-15 | 24 | NC 13 | 157 | 5.6658 |  | 8.1
2012-15 | 24 | NC 13 | 158 | 5.516 | 6 | 8.1
2012-15 | 24 | NC 13 | 159 | 5.3974 |  | 8.2
2012-15 | 24 | NC 13 | 160 | 5.3371 |  | 8.3
2012-15 | 24 | NC 13 | 161 | 5.3592 |  | 8.3
2012-15 | 24 | NC 13 | 162 | 5.3469 |  | 8.3
2012-15 | 24 | NC 13 | 163 | 5.33 |  | 8.3
2012-15 | 24 | NC 13 | 164 | 5.3346 |  | 8.3
2012-15 | 24 | NC 13 | 165 | 5.8381 |  | 8.6
2012-15 | 35 | NC 23 | 166 | 5.5354 |  | 8.2
2012-15 | 35 | NC 23 | 167 | 5.5187 |  | 8.2
2012-15 | 35 | NC 23 | 168 | 5.5294 |  | 8.2
2012-15 | 35 | NC 23 | 169 | 5.5843 |  | 8.3
2012-15 | 35 | NC 23 | 170 | 5.7231 |  | 8.3
2012-15 | 36 | NC 24 | 171 | 5.5375 |  | 8.3
2012-15 | 36 | NC 24 | 172 | 5.5224 |  | 8.2
2012-15 | 36 | NC 24 | 173 | 5.5011 |  | 8.2
2012-15 | 36 | NC 24 | 174 | 5.5299 |  | 8.3
2012-15 | 36 | NC 24 | 175 | 5.6136 |  | 8.3
2012-15 | 36 | NC 24 | 176 | 5.7632 |  | 8.3
2012-15 | 39 | NC 14 | 177 | 5.5669 |  | 8.2
2012-15 | 39 | NC 14 | 178 | 5.5063 |  | 8.2
2012-15 | 39 | NC 14 | 179 | 5.5249 |  | 8.1
2012-15 | 39 | NC 14 | 180 | 5.6077 |  | 8.1
2012-15 | 39 | NC 14 | 181 | 5.7942 |  | 8.3
2012-15 | 40 | NC 15 | 182 | 5.5841 |  | 8.1
2012-15 | 40 | NC 15 | 183 | 5.5728 |  | 8.1
2012-15 | 40 | NC 15 | 184 | 5.5477 |  | 8.1
2012-15 | 40 | NC 15 | 185 | 5.5077 |  | 8.1
2012-15 | 40 | NC 15 | 186 | 5.518 |  | 8.2
2012-15 | 40 | NC 15 | 187 | 5.617 |  | 8.3
2012-15 | 40 | NC 15 | 188 | 5.7684 |  | 8.5
2012-15 | 41 | NC 16 | 189 | 5.7245 |  | 8.1
2012-15 | 41 | NC 16 | 190 | 5.6136 |  | 8.2
2012-15 | 41 | NC 16 | 191 | 5.5236 |  | 8.1
2012-15 | 41 | NC 16 | 192 | 5.4615 |  | 8.1
2012-15 | 41 | NC 16 | 193 | 5.4548 |  | 8.1
2012-15 | 41 | NC 16 | 194 | 5.4794 |  | 8.1
2012-15 | 41 | NC 16 | 195 | 5.5228 |  | 8.1
2012-15 | 41 | NC 16 | 196 | 5.619 |  | 8.3
2012-15 | 42 | NC 17 | 197 | 5.4086 | 4 | 8.1 | OXY: Titration completed at 75.5%. Endpoint accuracy questionable. Value seems low in depth profile. Outlier in COMPARE (by 0.37mL/L)
2012-15 | 42 | NC 17 | 198 | 5.4936 | 4 | 8.1 | OXY: Titrationonly up to 94%. No end point. Looks to have been something blocking the light path. Outlier in COMPARE (by 0.23mL/L)
2012-15 | 42 | NC 17 | 199 | 5.6661 | 6 | 8.05
2012-15 | 42 | NC 17 | 200 | 5.613 |  | 8
2012-15 | 42 | NC 17 | 201 | 5.5425 | 4 | 8.1 | OXY: Titration incomplete. No end point.
2012-15 | 42 | NC 17 | 202 | 5.4235 | 4 | 8.1 | OXY: Titration incomplete. No end point. Slight outlier in COMPARE (by 0.08 mL/L)
2012-15 | 42 | NC 17 | 203 | 5.4539 |  | 8.1
2012-15 | 42 | NC 17 | 204 | 5.4137 |  | 8.1
2012-15 | 42 | NC 17 | 205 | 5.4543 |  | 8.2
2012-15 | 42 | NC 17 | 206 | 5.7226 |  | 8.4
2012-15 | 42 | NC 17 | 207 | 5.786 |  | 8.5
2012-15 | 48 | NC 19 | 208 | 5.8654 |  | 8
2012-15 | 48 | NC 19 | 209 | 5.8289 |  | 8
2012-15 | 48 | NC 19 | 210 | 5.8116 |  | 8
2012-15 | 48 | NC 19 | 211 | 5.8081 |  | 8
2012-15 | 48 | NC 19 | 212 | 5.7213 |  | 8.1
2012-15 | 48 | NC 19 | 213 | 5.6368 |  | 8.1
2012-15 | 48 | NC 19 | 214 | 5.5831 |  | 8.1
2012-15 | 48 | NC 19 | 215 | 5.5584 | 6 | 8.1
2012-15 | 48 | NC 19 | 216 | 5.534 |  | 8.1
2012-15 | 48 | NC 19 | 217 | 5.4833 |  | 8.1
2012-15 | 48 | NC 19 | 218 | 5.5903 |  | 8.3
2012-15 | 50 | NC 20 | 219 | 5.8235 |  | 8.1
2012-15 | 50 | NC 20 | 220 | 5.8191 |  | 8
2012-15 | 50 | NC 20 | 221 | 5.7719 |  | 8.1
2012-15 | 50 | NC 20 | 222 | 5.7123 | 6 | 8.1
2012-15 | 50 | NC 20 | 223 | 5.6689 |  | 8.1
2012-15 | 50 | NC 20 | 224 | 5.6268 | 3 | 8.1 | OXY: Endpoint ambiguous but it looked to be quite late. Adjusted
2012-15 | 50 | NC 20 | 225 | 5.6171 |  | 8.1
2012-15 | 50 | NC 20 | 226 | 5.5855 |  | 8.1
2012-15 | 50 | NC 20 | 227 | 5.6181 |  | 8.1
2012-15 | 50 | NC 20 | 228 | 5.8337 |  | 8.2
2012-15 | 51 | NC 21 | 229 | 5.7069 |  | 8.1
2012-15 | 51 | NC 21 | 230 | 5.6993 |  | 8.1
2012-15 | 51 | NC 21 | 231 | 5.6615 |  | 8.2
2012-15 | 51 | NC 21 | 232 | 5.6386 |  | 8.2
2012-15 | 51 | NC 21 | 233 | 5.6364 |  | 8.2
2012-15 | 51 | NC 21 | 234 | 5.6367 |  | 8.2
2012-15 | 51 | NC 21 | 235 | 5.8715 |  | 8.3
2012-15 | 51 | NC 21 | 236 | 6.395 |  | 8.6
2012-15 | 52 | NC 22 | 237 | 5.6758 |  | 8.3
2012-15 | 52 | NC 22 | 238 | 5.6343 |  | 8.2
2012-15 | 52 | NC 22 | 239 | 5.6364 |  | 8.2
2012-15 | 52 | NC 22 | 240 | 5.7046 |  | 8.2
2012-15 | 52 | NC 22 | 241 | 5.8904 |  | 8.3
2012-15 | 52 | NC 22 | 242 | 6.5236 |  | 8.6
2012-15 | 57 | NC 25 | 243 | 5.6287 |  | 8.1
2012-15 | 57 | NC 25 | 244 | 5.6067 |  | 8.1
2012-15 | 57 | NC 25 | 245 | 5.557 |  | 8.1
2012-15 | 57 | NC 25 | 246 | 5.4576 |  | 8.1
2012-15 | 57 | NC 25 | 247 | 5.4167 |  | 8.1
2012-15 | 57 | NC 25 | 248 | 5.4034 |  | 8.2
2012-15 | 57 | NC 25 | 249 | 5.7459 |  | 8.3
2012-15 | 57 | NC 25 | 250 | 6.1636 |  | 8.5
2012-15 | 61 | NC 26 | 251 | 5.7189 |  | 8.1
2012-15 | 61 | NC 26 | 252 | 5.6343 |  | 8.1
2012-15 | 61 | NC 26 | 253 | 5.542 |  | 8.1
2012-15 | 61 | NC 26 | 254 | 5.516 |  | 8.1
2012-15 | 61 | NC 26 | 255 | 5.4414 |  | 8.2
2012-15 | 61 | NC 26 | 256 | 5.4288 |  | 8.2
2012-15 | 61 | NC 26 | 257 | 5.4909 |  | 8.2
2012-15 | 61 | NC 26 | 258 | 5.7585 |  | 8.4
2012-15 | 66 | NC 27 | 259 | 5.6964 |  | 8.3
2012-15 | 66 | NC 27 | 260 | 5.6677 |  | 8.3
2012-15 | 66 | NC 27 | 261 | 5.6562 |  | 8.3
2012-15 | 66 | NC 27 | 262 | 5.6364 |  | 8.3
2012-15 | 66 | NC 27 | 263 | 5.4969 |  | 8.3
2012-15 | 66 | NC 27 | 264 | 5.4095 |  | 8.4
2012-15 | 66 | NC 27 | 265 | 5.4605 |  | 8.4
2012-15 | 66 | NC 27 | 266 | 5.5693 |  | 8.5
2012-15 | 66 | NC 27 | 267 | 5.8507 |  | 8.6
2012-15 | 67 | NC 28 | 268 | 5.4446 | 3 | -99 | OXY: Seems low. Not accurate. Add 1 + 2 titer. Zeroed colorimeter in sample. No end point. Slight outlier in COMPARE (by 0.07mL/L)
2012-15 | 67 | NC 28 | 269 | 5.49 |  | 8.4
2012-15 | 67 | NC 28 | 270 | 5.429 |  | 8.4
2012-15 | 67 | NC 28 | 271 | 5.4661 |  | 8.5
2012-15 | 67 | NC 28 | 272 | 5.6659 |  | 8.6
2012-15 | 67 | NC 28 | 273 | 5.7779 |  | 8.6
2012-15 | 68 | NC 29 | 274 | 5.6025 |  | 8.4
2012-15 | 68 | NC 29 | 275 | 5.4388 |  | 8.4
2012-15 | 68 | NC 29 | 276 | 5.4295 |  | 8.5
2012-15 | 68 | NC 29 | 277 | 5.4544 |  | 8.5
2012-15 | 68 | NC 29 | 278 | 5.3149 |  | 8.5
2012-15 | 68 | NC 29 | 279 | 5.646 |  | 8.6

## Duplicates
Precision Statement for Replicate Samples Drawn From a Single Niskin Bottle:
Sp for Dissolved Oxygen for the range 5.409 to 5.778 mL/L was 0.005, k = 14 with 1 outlier(s) removed
Precision Calculation for Duplicate Samples:
Precision was determined by analyzing replicate samples drawn from one Niskin
Pooled standard deviation of pairs of samples (Sp) was calculated by:
Sp = SQRT{sum (d*d)/2k}
where k = no. of pairs and d = difference between pairs
Determination of Outliers
Outliers are discarded on the basis of Chauvenet's criteria. The statistic is calculated by the difference between the outlier and the mean, divided by the stdev.
If this absolute value is greater than the critical value of the Chauvenet criterion for the given n, the data point can be discarded
Duplicate Analysis for Samples Drawn From a Single Niskin Bottle:
Event | Station | Sample_ Number | Depth | DO1 | DO2 | DOAvg | Flask1 | Flask2 | Draw T1 | Draw T2 | Draw TAvg | Difference | Diff2 | Chauvenet | Reject? | DO1 Comment | DO2 Comment | Critical values for Chauvenet statistic
1 | NC 01 | 9 | 75 | 5.7514 | 5.7503 | 5.7508 | 1475 | 1476 | 7.8 | 7.8 | 7.8 | -0.0010552 | 1.1134e-06 | 0.16835 |  |  |  | n | Chauvenet
2 | NC 02 | 24 | 30 | 5.6793 | 5.6723 | 5.6758 | 1055 | 1059 | 8 | 8 | 8 | -0.0069258 | 4.7966e-05 | 0.50632 |  |  |  | 3 | 1.37
4 | NC 4 | 46 | 75 | 5.6052 | 5.6065 | 5.6059 | 1374 | 1377 | 8.2 | 8.2 | 8.2 | 0.001265 | 1.6003e-06 | 0.435 |  |  |  | 4 | 1.53
9 | NC 5 | 61 | 20 | 5.4223 | 5.4091 | 5.4157 | 1386 | 1387 | 8.2 | 8.2 | 8.2 | -0.013193 | 0.00017405 | 1.2265 |  |  |  | 5 | 1.64
10 | NC 6 | 65 | 200 | 5.774 | 5.7773 | 5.7756 | 1421 | 1459 | 8.1 | 8.1 | 8.1 | 0.0032819 | 1.0771e-05 | 0.66679 |  | OXY: Endpoint ambiguous. Looked to be late by using regression analysis. Corrected | OXY: Endpoint late. Corrected | 6 | 1.73
11 | NC 7 | 77 | 200 | 5.7565 | 5.7785 | 5.7675 | 196 | 371 | 8.1 | 8.1 | 8.1 | 0.021924 | 0.00048064 | 2.8092 | Y | OXY: Endpoint ambiguous. Curve does not flatten nicely | OXY: Endpoint was very early. Adjusted | 7 | 1.8
12 | NC 8 | 89 | 200 | 5.7275 | 5.7264 | 5.727 | 657 | 668 | 8.2 | 8.2 | 8.2 | -0.0011254 | 1.2665e-06 | 0.16028 |  |  |  | 8 | 1.87
13 | NC 9 | 102 | 200 | 5.712 | 5.707 | 5.7095 | 1247 | 1248 | 8.2 | 8.2 | 8.2 | -0.005 | 2.5e-05 | 0.285 |  |  |  | 9 | 1.91
19 | NC 10 | 120 | 75 | 5.4609 | 5.4604 | 5.4606 | 1377 | 1393 | 8.1 | 8.1 | 8.1 | -0.00044749 | 2.0025e-07 | 0.23819 |  |  |  | 10 | 1.97
20 | NC 11 | 135 | 75 | 5.4581 | 5.4446 | 5.4514 | 1307 | 1309 | 8.1 | 8.1 | 8.1 | -0.013468 | 0.00018139 | 1.2582 |  |  |  | 20 | 2.24
24 | NC 13 | 158 | 125 | 5.5161 | 5.5158 | 5.516 | 396 | 468 | 8.1 | 8.1 | 8.1 | -0.0002548 | 6.4925e-08 | 0.26034 |  |  |  | 30 | 2.39
42 | NC 17 | 199 | 150 | 5.67 | 5.6623 | 5.6661 | 1247 | 1248 | 8.1 | 8 | 8.05 | -0.0077681 | 6.0343e-05 | 0.60312 |  |  |  | 50 | 2.57
48 | NC 19 | 215 | 30 | 5.5603 | 5.5564 | 5.5584 | 1300 | 1306 | 8.1 | 8.1 | 8.1 | -0.0039665 | 1.5733e-05 | 0.16623 |  |  |  | 80 | 2.74
50 | NC 20 | 222 | 100 | 5.7165 | 5.708 | 5.7123 | 708 | 712 | 8.1 | 8.1 | 8.1 | -0.0085476 | 7.3062e-05 | 0.69271 |  |  |  | 100 | 2.81
[18 empty] | 150 | 2.93
[11 empty] | Mean Diff | -0.0025201 | [5 empty] | 200 | 3.02
[11 empty] | Stdev | 0.0087014
[11 empty] | Count |  | 14
[11 empty] | Sum |  | 0.0010732
[11 empty] | Sp |  | 0.006191
[11 empty] | MIN/MAX | 5.4091 | 5.7785
x outlier(s) excluded
[11 empty] | Count |  | 13
[11 empty] | Sum |  | 0.00059256
[11 empty] | Sp |  | 0.004774

## Flags
Institute of Ocean Sciences -Quality Flags for Water Sampling | Last Updated: April 2012
Double click document to use.

## Oxy Data Processing
[5 empty] | Total # of samples: | 295
File Name | Event | Station | Sample_ Number | Depth | Raw Oxygen: Dissolved | AutoOxy Thiosulfate_ titer | Regression Thiosulfate_ titer | Flask_ Number | Flask_ Volume | Blank | Standard | KIO3 Std volume ml | KIO3 Normality | THIO N | AutoOxy Oxygen: Dissolved | AutoOxy Oxygen: Dissolved | Flag | Temperature: Draw | Comment
2012-150001.oxy | 1 | NC 01 | 1 | 335 | 0 | 0 | 0 | 1414 | 132.62 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | -99 | -99 | 5 | 8 | OXY: Dosimat ran on. Endpoint overshot. Sample lost.
2012-150001.oxy | 1 | NC 01 | 2 | 300 | 5.679 | 0.997 | 0.9972 | 1421 | 140.09 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6797 | 5.6808 |  | 7.9
2012-150001.oxy | 1 | NC 01 | 3 | 250 | 5.756 | 0.958 | 0.9575 | 1459 | 132.92 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7556 | 5.7525 |  | 7.9
2012-150001.oxy | 1 | NC 01 | 4 | 200 | 5.748 | 0.944 | 0.9435 | 1464 | 131.17 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.748 | 5.7449 |  | 7.9
2012-150001.oxy | 1 | NC 01 | 5 | 175 | 5.754 | 1.014 | 1.014 | 1465 | 140.64 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7539 | 5.7539 |  | 7.9
2012-150001.oxy | 1 | NC 01 | 6 | 150 | 5.742 | 0.961 | 0.961 | 1470 | 133.64 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7421 | 5.7421 |  | 7.8
2012-150001.oxy | 1 | NC 01 | 7 | 125 | 5.745 | 1.031 | 1.0316 | 1472 | 143.19 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.745 | 5.7484 |  | 7.8
2012-150001.oxy | 1 | NC 01 | 8 | 100 | 5.76 | 0.95 | 0.9495 | 1473 | 131.73 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7597 | 5.7566 |  | 7.8
2012-150001.oxy | 1 | NC 01 | 9 | 75 | 5.752 | 0.938 | 0.938 | 1475 | 130.27 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7514 | 5.7514 |  | 7.8
2012-150001.oxy | 1 | NC 01 | 9 | 75 | 5.751 | 0.954 | 0.954 | 1476 | 132.49 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7503 | 5.7503 |  | 7.8
2012-150001.oxy | 1 | NC 01 | 10 | 50 | 5.751 | 0.968 | 0.9677 | 1477 | 134.4 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7508 | 5.749 |  | 7.8
2012-150001.oxy | 1 | NC 01 | 11 | 30 | 5.801 | 1.03 | 1.03 | 1478 | 141.68 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.8015 | 5.8015 | 4 | 7.9 | OXY: Air bubble on lens. Transmissivity only to 93%. No clear end point. Slight outlier in COMPARE (by 0.05 mL/L)
2012-150001.oxy | 1 | NC 01 | 12 | 20 | 5.716 | 0.943 | 0.9429 | 1481 | 131.76 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7158 | 5.7151 |  | 7.9
2012-150001.oxy | 1 | NC 01 | 13 | 10 | 5.699 | 0.945 | 0.945 | 1483 | 132.43 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6985 | 5.6985 |  | 7.9
2012-150001.oxy | 1 | NC 01 | 14 | 0 | 5.681 | 0.947 | 0.947 | 1485 | 133.1 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6814 | 5.6814 |  | 7.9
2012-150002.oxy | 2 | NC 02 | 15 | 318 | 5.834 | 0.99 | 0.9892 | 657 | 135.49 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.834 | 5.8292 |  | 7.9
2012-150002.oxy | 2 | NC 02 | 16 | 300 | 5.825 | 0.949 | 0.9483 | 668 | 130.14 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.825 | 5.8207 |  | 7.9
2012-150002.oxy | 2 | NC 02 | 17 | 250 | 5.793 | 0.959 | 0.959 | 690 | 132.2 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7934 | 5.7934 | 4 | 7.9 | OXY: A poor titration with no discernable endpoint. That it is not terribly out of line is a fluke
2012-150002.oxy | 2 | NC 02 | 18 | 200 | 5.795 | 1.016 | 1.016 | 708 | 139.92 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7954 | 5.7954 |  | 7.9
2012-150002.oxy | 2 | NC 02 | 19 | 175 | 5.788 | 0.938 | 0.938 | 712 | 129.47 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7875 | 5.7875 |  | 7.9
2012-150002.oxy | 2 | NC 02 | 20 | 150 | 5.782 | 1.006 | 1.0065 | 715 | 138.88 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7818 | 5.7846 |  | 7.9
2012-150002.oxy | 2 | NC 02 | 21 | 100 | 5.772 | 0.94 | 0.94 | 745 | 130.09 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7718 | 5.7718 |  | 7.9
2012-150002.oxy | 2 | NC 02 | 22 | 75 | 5.757 | 1.038 | 1.0386 | 1002 | 143.85 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7573 | 5.7606 |  | 7.9
2012-150002.oxy | 2 | NC 02 | 23 | 50 | 5.734 | 1.034 | 1.0342 | 1035 | 143.89 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7334 | 5.7345 |  | 8
2012-150002.oxy | 2 | NC 02 | 24 | 30 | 5.684 | 1.005 | 1.0042 | 1055 | 141.1 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6838 | 5.6793 |  | 8
2012-150002.oxy | 2 | NC 02 | 24 | 30 | 5.674 | 1.009 | 1.0088 | 1059 | 141.91 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6735 | 5.6723 |  | 8
2012-150002.oxy | 2 | NC 02 | 25 | 20 | 5.638 | 0.933 | 0.9331 | 1085 | 132.16 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6375 | 5.6381 | 3 | 8 | OXY: Air bubble on lens. Transmissivity only to 96%
2012-150002.oxy | 2 | NC 02 | 26 | 10 | 5.596 | 0.894 | 0.894 | 1105 | 127.62 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5962 | 5.5962 |  | 8
2012-150002.oxy | 2 | NC 02 | 27 | 0 | 5.747 | 0.955 | 0.9542 | 1106 | 132.7 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7471 | 5.7423 |  | 8
2012-150003.oxy | 3 | NC 03 | 28 | 279 | 6.349 | 1.058 | 1.0573 | 94 | 133.13 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 6.3483 | 6.3441 |  | 8.1
2012-150003.oxy | 3 | NC 03 | 29 | 200 | 5.774 | 0.988 | 0.9885 | 196 | 136.6 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7741 | 5.7771 |  | 8
2012-150003.oxy | 3 | NC 03 | 30 | 175 | 5.769 | 0.942 | 0.9418 | 371 | 130.42 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7692 | 5.768 |  | 8
2012-150003.oxy | 3 | NC 03 | 30 | 175 | 5.425 | 0.902 | 0.902 | 396 | 132.76 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4245 | -99 | 5 | 8 | OXY: Air bubble on lens. Transmissivity only to 90%. Titration incomplete. Out of line in depth profile
2012-150003.oxy | 3 | NC 03 | 31 | 150 | 4.745 | 0.802 | 0.802 | 468 | 134.84 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 4.7453 | -99 | 5 | 8 | OXY: Air bubble on lens. Transmissivity only to 81%. Titration incomplete. Out of line in depth profile
2012-150003.oxy | 3 | NC 03 | 32 | 125 | 5.776 | 0.973 | 0.9725 | 471 | 134.51 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7758 | 5.7729 |  | 8
2012-150003.oxy | 3 | NC 03 | 33 | 100 | 5.78 | 0.946 | 0.944 | 505 | 130.73 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7799 | 5.7676 |  | 8 | End point was way late. Corrected
2012-150003.oxy | 3 | NC 03 | 34 | 75 | 5.748 | 0.964 | 0.9646 | 510 | 133.91 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7483 | 5.7519 |  | 8
2012-150003.oxy | 3 | NC 03 | 35 | 50 | 5.671 | 0.922 | 0.9223 | 528 | 129.86 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.671 | 5.6729 |  | 8
2012-150003.oxy | 3 | NC 03 | 36 | 30 | 5.476 | 0.904 | 0.9045 | 552 | 131.82 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.476 | 5.479 |  | 8.1
2012-150003.oxy | 3 | NC 03 | 37 | 20 | 5.493 | 0.932 | 0.9322 | 557 | 135.45 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4926 | 5.4938 |  | 8.2
2012-150003.oxy | 3 | NC 03 | 38 | 10 | 5.569 | 0.947 | 0.9475 | 559 | 135.75 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5688 | 5.5718 |  | 8.2
2012-150003.oxy | 3 | NC 03 | 39 | 0 | 5.957 | 1.02 | 1.0196 | 569 | 136.71 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.9569 | 5.9546 |  | 8.2
2012-150004.oxy | 4 | NC 4 | 40 | 251 | 5.807 | 0.983 | 0.9828 | 1335 | 135.15 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.8074 | 5.8062 |  | 8.2
2012-150004.oxy | 4 | NC 4 | 41 | 200 | 5.769 | 1.027 | 1.0272 | 1337 | 142.06 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7689 | 5.77 |  | 8.1
2012-150004.oxy | 4 | NC 4 | 42 | 175 | 5.777 | 0.95 | 0.95 | 1338 | 131.33 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7775 | 5.7775 |  | 8.1
2012-150004.oxy | 4 | NC 4 | 43 | 150 | 5.759 | 0.979 | 0.979 | 1365 | 135.73 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7586 | 5.7586 |  | 8.1
2012-150004.oxy | 4 | NC 4 | 44 | 125 | 5.738 | 0.98 | 0.98 | 1367 | 136.34 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7383 | 5.7383 |  | 8.1
2012-150004.oxy | 4 | NC 4 | 45 | 100 | 5.722 | 0.969 | 0.969 | 1373 | 135.2 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7222 | 5.7222 |  | 8.1
2012-150004.oxy | 4 | NC 4 | 46 | 75 | 5.608 | 0.945 | 0.9445 | 1374 | 134.53 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6082 | 5.6052 |  | 8.2
2012-150004.oxy | 4 | NC 4 | 46 | 75 | 5.604 | 0.924 | 0.9244 | 1377 | 131.67 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6041 | 5.6065 |  | 8.2
2012-150004.oxy | 4 | NC 4 | 47 | 50 | 5.536 | 0.919 | 0.9187 | 1378 | 132.54 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5365 | 5.5347 |  | 8.2
2012-150004.oxy | 4 | NC 4 | 48 | 30 | 5.469 | 0.921 | 0.9211 | 1386 | 134.43 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4694 | 5.47 |  | 8.3
2012-150004.oxy | 4 | NC 4 | 49 | 20 | 5.555 | 0.941 | 0.9405 | 1387 | 135.22 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5554 | 5.5525 |  | 8.3
2012-150004.oxy | 4 | NC 4 | 50 | 10 | 5.701 | 0.977 | 0.9765 | 1389 | 136.81 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7007 | 5.6978 |  | 8.3
2012-150004.oxy | 4 | NC 4 | 51 | 0 | 6.07 | 1 | 1 | 1392 | 131.6 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 6.07 | 6.07 |  | 8.4
2012-150009.oxy | 9 | NC 5 | 52 | 218 | 5.772 | 0.977 | 0.977 | 1335 | 135.15 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7718 | 5.7718 |  | 8.2
2012-150009.oxy | 9 | NC 5 | 53 | 200 | 5.769 | 1.027 | 1.0266 | 1337 | 142.06 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7689 | 5.7666 |  | 8.1
2012-150009.oxy | 9 | NC 5 | 54 | 175 | 5.765 | 0.948 | 0.9482 | 1338 | 131.33 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7653 | 5.7665 |  | 8.1
2012-150009.oxy | 9 | NC 5 | 55 | 150 | 5.753 | 0.978 | 0.9775 | 1365 | 135.73 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7527 | 5.7497 |  | 8.1
2012-150009.oxy | 9 | NC 5 | 56 | 125 | 5.709 | 0.975 | 0.9745 | 1367 | 136.34 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7089 | 5.706 |  | 8.1
2012-150009.oxy | 9 | NC 5 | 57 | 100 | 5.663 | 0.959 | 0.959 | 1373 | 135.2 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.663 | 5.663 |  | 8.1
2012-150009.oxy | 9 | NC 5 | 58 | 75 | 5.513 | 0.929 | 0.929 | 1374 | 134.53 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5129 | 5.5129 |  | 8.2
2012-150009.oxy | 9 | NC 5 | 59 | 50 | 5.483 | 0.904 | 0.903 | 1377 | 131.67 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4823 | 5.4762 |  | 8.2
2012-150009.oxy | 9 | NC 5 | 60 | 30 | 5.415 | 0.899 | 0.8994 | 1378 | 132.54 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4155 | 5.4179 |  | 8.2
2012-150009.oxy | 9 | NC 5 | 61 | 20 | 5.422 | 0.913 | 0.9131 | 1386 | 134.43 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4217 | 5.4223 |  | 8.2
2012-150009.oxy | 9 | NC 5 | 61 | 20 | 5.407 | 0.916 | 0.9163 | 1387 | 135.22 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4073 | 5.4091 |  | 8.2 | Check endpoint. Hugh thought it might be early. End point ok.
2012-150009.oxy | 9 | NC 5 | 62 | 10 | 5.502 | 0.943 | 0.9434 | 1389 | 136.81 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5016 | 5.504 |  | 8.2
2012-150009.oxy | 9 | NC 5 | 63 | 0 | 5.668 | 0.934 | 0.9336 | 1392 | 131.6 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.668 | 5.6656 |  | 8.2
2012-150010.oxy | 10 | NC 6 | 64 | 226 | 5.787 | 0.961 | 0.9607 | 1414 | 132.62 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7869 | 5.7851 |  | 8.1
2012-150010.oxy | 10 | NC 6 | 65 | 200 | 5.782 | 1.015 | 1.0135 | 1421 | 140.09 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7826 | 5.774 | 3 | 8.1 | OXY: Endpoint ambiguous. Looked to be late by using regression analysis. Adjusted
2012-150010.oxy | 10 | NC 6 | 65 | 200 | 5.786 | 0.963 | 0.9616 | 1459 | 132.92 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7857 | 5.7773 | 3 | 8.1 | OXY: Endpoint late. Adjusted
2012-150010.oxy | 10 | NC 6 | 66 | 175 | 5.778 | 0.949 | 0.949 | 1464 | 131.17 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7785 | 5.7785 |  | 8
2012-150010.oxy | 10 | NC 6 | 67 | 150 | 5.748 | 1.013 | 1.013 | 1465 | 140.64 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7482 | 5.7482 |  | 8.1 | Incorrect depth. Corrected
2012-150010.oxy | 10 | NC 6 | 68 | 125 | 5.73 | 0.959 | 0.959 | 1470 | 133.64 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7301 | 5.7301 |  | 8.1
2012-150010.oxy | 10 | NC 6 | 69 | 100 | 5.7 | 1.023 | 1.0223 | 1472 | 143.19 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7003 | 5.6964 |  | 8.1
2012-150010.oxy | 10 | NC 6 | 70 | 75 | 5.626 | 0.928 | 0.927 | 1473 | 131.73 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6258 | 5.6197 |  | 8.1
2012-150010.oxy | 10 | NC 6 | 71 | 50 | 5.505 | 0.898 | 0.898 | 1475 | 130.27 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5052 | 5.5052 |  | 8.2
2012-150010.oxy | 10 | NC 6 | 72 | 30 | 5.406 | 0.897 | 0.8974 | 1476 | 132.49 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4055 | 5.4079 |  | 8.3
2012-150010.oxy | 10 | NC 6 | 73 | 20 | 5.423 | 0.913 | 0.9123 | 1477 | 134.4 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4229 | 5.4187 |  | 8.3
2012-150010.oxy | 10 | NC 6 | 74 | 10 | 5.53 | 0.982 | 0.983 | 1478 | 141.68 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5302 | 5.5359 |  | 8.3
2012-150010.oxy | 10 | NC 6 | 75 | 0 | 5.625 | 0.928 | 0.9285 | 1481 | 131.76 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6245 | 5.6275 |  | 8.3
2012-150011.oxy | 11 | NC 7 | 76 | 226 | 5.789 | 0.965 | 0.965 | 94 | 133.13 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7885 | 5.7885 |  | 8.1
2012-150011.oxy | 11 | NC 7 | 77 | 200 | 5.757 | 0.985 | 0.985 | 196 | 136.6 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7565 | 5.7565 | 4 | 8.1 | OXY: Endpoint ambiguous. Curve does not flatten nicely
2012-150011.oxy | 11 | NC 7 | 77 | 200 | 5.769 | 0.942 | 0.9435 | 371 | 130.42 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7692 | 5.7785 | 3 | 8.1 | OXY: Endpoint was very early. Adjusted
2012-150011.oxy | 11 | NC 7 | 78 | 175 | 5.769 | 0.959 | 0.9582 | 396 | 132.76 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7686 | 5.7638 |  | 8.1
2012-150011.oxy | 11 | NC 7 | 79 | 150 | 5.738 | 0.969 | 0.9699 | 468 | 134.84 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7377 | 5.7431 |  | 8.1
2012-150011.oxy | 11 | NC 7 | 80 | 125 | 5.71 | 0.962 | 0.962 | 471 | 134.51 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7103 | 5.7103 | 3 | 8.1 | OXY: Wavy, ugly curve. Endpoint ambiguous.
2012-150011.oxy | 11 | NC 7 | 81 | 100 | 5.627 | 0.921 | 0.9208 | 505 | 130.73 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6266 | 5.6253 |  | 8.1
2012-150011.oxy | 11 | NC 7 | 82 | 75 | 5.623 | 0.943 | 0.943 | 510 | 133.91 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6226 | 5.6226 | 3 | 8.2 | OXY: Messy curve. What seems to be the best endpoint was chosen by AutoOxy.
2012-150011.oxy | 11 | NC 7 | 83 | 50 | 5.443 | 0.913 | 0.913 | 528 | 129.858 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6156 | 5.6156 | 3 | 8.1 | OXY: Messy curve. What seems to be the best endpoint was chosen by AutoOxy. Seems out of line in the depth profile.
2012-150011.oxy | 11 | NC 7 | 84 | 30 | 5.397 | 0.891 | 0.891 | 552 | 131.82 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.3969 | 5.3969 |  | 8.3
2012-150011.oxy | 11 | NC 7 | 85 | 20 | 5.463 | 0.927 | 0.9255 | 557 | 135.45 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4631 | 5.4542 |  | 8.3 | Endpoint was too late
2012-150011.oxy | 11 | NC 7 | 86 | 10 | 5.486 | 0.933 | 0.9333 | 559 | 135.75 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4862 | 5.488 |  | 8.3
2012-150011.oxy | 11 | NC 7 | 87 | 0 | 5.547 | 0.95 | 0.9495 | 569 | 136.71 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5467 | 5.5438 |  | 8.3
2012-150012.oxy | 12 | NC 8 | 88 | 255 | 5.766 | 0.939 | 0.94 | 745 | 130.09 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7656 | 5.7718 |  | 8.2
2012-150012.oxy | 12 | NC 8 | 89 | 200 | 5.727 | 0.972 | 0.972 | 657 | 135.49 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7275 | 5.7275 |  | 8.2
2012-150012.oxy | 12 | NC 8 | 89 | 200 | 5.726 | 0.933 | 0.933 | 668 | 130.14 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7264 | 5.7264 |  | 8.2
2012-150012.oxy | 12 | NC 8 | 90 | 175 | 5.678 | 0.94 | 0.939 | 690 | 132.2 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6782 | 5.6722 |  | 8.2
2012-150012.oxy | 12 | NC 8 | 91 | 150 | 5.664 | 0.993 | 0.993 | 708 | 139.92 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6638 | 5.6638 |  | 8.2
2012-150012.oxy | 12 | NC 8 | 92 | 125 | 5.651 | 0.916 | 0.9156 | 712 | 129.47 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6512 | 5.6488 |  | 8.2
2012-150012.oxy | 12 | NC 8 | 93 | 100 | 5.626 | 0.979 | 0.979 | 715 | 138.88 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6261 | 5.6261 |  | 8.2
2012-150012.oxy | 12 | NC 8 | 94 | 75 | 5.484 | 0.989 | 0.9889 | 1002 | 143.85 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4846 | 5.484 |  | 8.3
2012-150012.oxy | 12 | NC 8 | 95 | 50 | 5.45 | 0.983 | 0.983 | 1035 | 143.89 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4497 | 5.4497 |  | 8.3
2012-150012.oxy | 12 | NC 8 | 96 | 30 | 5.383 | 0.952 | 0.9514 | 1055 | 141.1 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.383 | 5.3796 |  | 8.3
2012-150012.oxy | 12 | NC 8 | 97 | 20 | 5.358 | 0.953 | 0.9537 | 1059 | 141.91 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.3575 | 5.3615 |  | 8.3
2012-150012.oxy | 12 | NC 8 | 98 | 10 | 5.49 | 0.877 | 0.8768 | 1105 | 127.62 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4894 | 5.4881 |  | 8.3
2012-150012.oxy | 12 | NC 8 | 99 | 0 | 5.638 | 0.933 | 0.934 | 1085 | 132.16 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6375 | 5.6436 |  | 8.3
2012-150013.oxy | 13 | NC 9 | 100 | 278 | 5.777 | 0.983 | 0.983 | 1213 | 135.84 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7774 | 5.7774 |  | 8.3
2012-150013.oxy | 13 | NC 9 | 101 | 250 | 5.775 | 1.043 | 1.043 | 1218 | 144.11 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7745 | 5.7745 |  | 8.2
2012-150013.oxy | 13 | NC 9 | 102 | 200 | 5.712 | 1.032 | 1.032 | 1247 | 144.14 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7122 | 5.7122 |  | 8.2
2012-150013.oxy | 13 | NC 9 | 102 | 200 | 5.757 | 1.04 | 1.04 | 1248 | 145.38 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.7068 | 5.7068 |  | 8.2 | Horrible duplicate was due to an incorrect flask ID and volume. Corrected
2012-150013.oxy | 13 | NC 9 | 103 | 175 | 5.677 | 0.936 | 0.936 | 1261 | 131.67 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6771 | 5.6771 |  | 8.2
2012-150013.oxy | 13 | NC 9 | 104 | 150 | 5.664 | 1.012 | 1.0121 | 1268 | 142.57 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6637 | 5.6642 |  | 8.2
2012-150013.oxy | 13 | NC 9 | 105 | 125 | 5.64 | 1.01 | 1.0103 | 1298 | 142.89 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6396 | 5.6413 |  | 8.2
2012-150013.oxy | 13 | NC 9 | 106 | 100 | 5.575 | 0.956 | 0.955 | 1300 | 136.87 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5753 | 5.5694 |  | 8.2
2012-150013.oxy | 13 | NC 9 | 107 | 75 | 5.519 | 0.92 | 0.9207 | 1306 | 133.09 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.5193 | 5.5235 |  | 8.2
2012-150013.oxy | 13 | NC 9 | 108 | 50 | 5.447 | 0.908 | 0.905 | 1307 | 133.08 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4474 | 5.4293 | 3 | 8.3 | OXY: End point was unclear. No good inflection point but seemed very late. Adjusted
2012-150013.oxy | 13 | NC 9 | 109 | 30 | 5.424 | 0.914 | 0.914 | 1309 | 134.52 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.424 | 5.424 |  | 8.3
2012-150013.oxy | 13 | NC 9 | 110 | 20 | 5.421 | 0.884 | 0.8845 | 1311 | 130.22 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4212 | 5.4242 |  | 8.3
2012-150013.oxy | 13 | NC 9 | 111 | 10 | 5.477 | 0.912 | 0.912 | 1312 | 132.94 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.4774 | 5.4774 |  | 8.3
2012-150013.oxy | 13 | NC 9 | 112 | 0 | 5.62 | 0.96 | 0.9595 | 1313 | 136.36 | 0.0013 | 0.8625 | 10.0277 | 0.012111 | 0.14101 | 5.6199 | 5.617 |  | 8.3
2012-150019.oxy | 19 | NC 10 | 113 | 306 | 5.757 | 0.974 | 0.974 | 1335 | 135.15 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7566 | 5.7566 |  | 8.1
2012-150019.oxy | 19 | NC 10 | 114 | 250 | 5.715 | 1.017 | 1.0177 | 1337 | 142.06 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7148 | 5.7187 |  | 8
2012-150019.oxy | 19 | NC 10 | 115 | 200 | 5.683 | 0.934 | 0.934 | 1338 | 131.33 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6826 | 5.6826 |  | 8.2
2012-150019.oxy | 19 | NC 10 | 116 | 175 | 5.667 | 0.963 | 0.963 | 1365 | 135.73 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6667 | 5.6667 | 4 | 8 | OXY: Titration incomplete
2012-150019.oxy | 19 | NC 10 | 117 | 150 | 5.659 | 0.966 | 0.9655 | 1367 | 136.34 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6586 | 5.6557 |  | 8
2012-150019.oxy | 19 | NC 10 | 118 | 125 | 5.6 | 0.948 | 0.947 | 1373 | 135.2 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6004 | 5.5945 |  | 8
2012-150019.oxy | 19 | NC 10 | 119 | 100 | 5.522 | 0.93 | 0.9296 | 1374 | 134.53 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5216 | 5.5192 |  | 8.1
2012-150019.oxy | 19 | NC 10 | 120 | 75 | 5.455 | 0.899 | 0.9 | 1377 | 131.67 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4548 | 5.4609 |  | 8.1
2012-150019.oxy | 19 | NC 10 | 120 | 75 | 5.466 | 0.97 | 0.969 | 1393 | 141.65 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4661 | 5.4604 |  | 8.1
2012-150019.oxy | 19 | NC 10 | 121 | 50 | 5.467 | 0.907 | 0.907 | 1378 | 132.54 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4668 | 5.4668 |  | 8.1
2012-150019.oxy | 19 | NC 10 | 122 | 30 | 5.43 | 0.914 | 0.913 | 1386 | 134.43 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4304 | 5.4245 | 4 | 8.1 | OXY: A very messy titration with no clear endpoint
2012-150019.oxy | 19 | NC 10 | 123 | 20 | 5.416 | 0.917 | 0.917 | 1387 | 135.22 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.416 | 5.416 |  | 8.2
2012-150019.oxy | 19 | NC 10 | 124 | 10 | 5.434 | 0.931 | 0.9305 | 1389 | 136.81 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4341 | 5.4311 |  | 8.2
2012-150019.oxy | 19 | NC 10 | 125 | 0 | 5.229 | 0.896 | 0.896 | 1392 | 131.6 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4395 | 5.4395 |  | 8.2 | Wrond flask ID and volume. Corrected
2012-150020.oxy | 20 | NC 11 | 126 | 322 | 5.75 | 0.978 | 0.9778 | 1213 | 135.84 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7505 | 5.7493 |  | 8.1
2012-150020.oxy | 20 | NC 11 | 127 | 322 | 5.755 | 1.039 | 1.039 | 1218 | 144.11 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7545 | 5.7545 |  | 8
2012-150020.oxy | 20 | NC 11 | 128 | 300 | 5.753 | 1.039 | 1.039 | 1247 | 144.14 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7533 | 5.7533 |  | 8
2012-150020.oxy | 20 | NC 11 | 129 | 250 | 5.588 | 1.018 | 1.018 | 1248 | 145.38 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.588 | -99 | 5 | 8 | OXY: Bubble on lens. Only seven points plotted in the titration. No discernable end point. A clear outlier in the depth profile
2012-150020.oxy | 20 | NC 11 | 130 | 200 | 5.717 | 0.942 | 0.9424 | 1261 | 131.67 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7164 | 5.7188 |  | 8
2012-150020.oxy | 20 | NC 11 | 131 | 175 | 5.689 | 1.016 | 1.0162 | 1268 | 142.57 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6885 | 5.6896 |  | 8
2012-150020.oxy | 20 | NC 11 | 132 | 150 | 5.659 | 1.013 | 1.0121 | 1298 | 142.89 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6587 | 5.6537 |  | 8
2012-150020.oxy | 20 | NC 11 | 133 | 125 | 5.619 | 0.963 | 0.9626 | 1300 | 136.87 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6188 | 5.6165 |  | 8
2012-150020.oxy | 20 | NC 11 | 134 | 100 | 5.51 | 0.918 | 0.918 | 1306 | 133.09 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.51 | 5.51 |  | 8.1
2012-150020.oxy | 20 | NC 11 | 135 | 75 | 5.462 | 0.91 | 0.9093 | 1307 | 133.08 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4623 | 5.4581 |  | 8.1
2012-150020.oxy | 20 | NC 11 | 135 | 75 | 5.445 | 0.917 | 0.917 | 1309 | 134.52 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4446 | 5.4446 |  | 8.1
2012-150020.oxy | 20 | NC 11 | 136 | 50 | 5.4 | 0.88 | 0.8795 | 1311 | 130.22 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.3995 | 5.3965 |  | 8.2
2012-150020.oxy | 20 | NC 11 | 137 | 30 | 5.426 | 0.903 | 0.9024 | 1312 | 132.94 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.426 | 5.4223 |  | 8.2
2012-150020.oxy | 20 | NC 11 | 138 | 20 | 5.476 | 0.935 | 0.9348 | 1313 | 136.36 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4758 | 5.4746 |  | 8.2
2012-150020.oxy | 20 | NC 11 | 139 | 10 | 5.556 | 0.945 | 0.9445 | 1314 | 135.85 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5556 | 5.5526 |  | 8.2
2012-150020.oxy | 20 | NC 11 | 140 | 0 | 5.739 | 0.961 | 0.961 | 1316 | 133.78 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7386 | 5.7386 |  | 8.2
2012-150023.oxy | 23 | NC 12 | 141 | 318 | 5.771 | 0.979 | 0.9785 | 657 | 135.49 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7715 | 5.7685 |  | 8.1
2012-150023.oxy | 23 | NC 12 | 142 | 300 | 5.76 | 0.938 | 0.9377 | 668 | 130.14 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.76 | 5.7582 |  | 8.1
2012-150023.oxy | 23 | NC 12 | 143 | 250 | 5.742 | 0.95 | 0.949 | 690 | 132.2 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7416 | 5.7355 |  | 8.1
2012-150023.oxy | 23 | NC 12 | 144 | 200 | 5.712 | 1.001 | 1.001 | 708 | 139.92 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7119 | 5.7119 |  | 8.1
2012-150023.oxy | 23 | NC 12 | 145 | 175 | 5.698 | 0.923 | 0.9225 | 712 | 129.47 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6975 | 5.6944 |  | 8.1
2012-150023.oxy | 23 | NC 12 | 146 | 150 | 5.634 | 0.98 | 0.9793 | 715 | 138.88 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6343 | 5.6303 |  | 8.1
2012-150023.oxy | 23 | NC 12 | 147 | 125 | 5.048 | 0.822 | 0.822 | 745 | 130.09 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.0478 | -99 | 5 | 8.1 | OXY: End point failure at 85%. Value way off. Confirmed in titration plot, depth profile and with duplicate sample value
2012-150023.oxy | 23 | NC 12 | 147 | 125 | 5.626 | 1.014 | 1.0141 | 1002 | 143.85 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.626 | 5.6266 |  | 8.1
2012-150023.oxy | 23 | NC 12 | 148 | 100 | 5.513 | 0.994 | 0.994 | 1035 | 143.89 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5132 | 5.5132 |  | 8.1
2012-150023.oxy | 23 | NC 12 | 149 | 75 | 5.454 | 0.964 | 0.9638 | 1055 | 141.1 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4536 | 5.4525 |  | 8.2
2012-150023.oxy | 23 | NC 12 | 150 | 50 | 5.45 | 0.969 | 0.969 | 1059 | 141.91 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4503 | 5.4503 |  | 8.2
2012-150023.oxy | 23 | NC 12 | 151 | 30 | 5.428 | 0.898 | 0.8985 | 1085 | 132.16 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4282 | 5.4312 |  | 8.2
2012-150023.oxy | 23 | NC 12 | 152 | 20 | 5.455 | 0.871 | 0.8705 | 1105 | 127.62 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4548 | 5.4516 |  | 8.2
2012-150023.oxy | 23 | NC 12 | 153 | 10 | 5.563 | 0.924 | 0.924 | 1106 | 132.7 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5627 | 5.5627 |  | 8.2
2012-150023.oxy | 23 | NC 12 | 154 | 0 | 5.635 | 0.959 | 0.9585 | 1123 | 135.93 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6347 | 5.6318 |  | 8.2
2012-150024.oxy | 24 | NC 13 | 155 | 300 | 5.773 | 0.962 | 0.962 | 94 | 133.13 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7731 | 5.7731 | 3 | 8.2 | OXY: Curve does an odd jump up near the end point, making it difficult to clearly identify
2012-150024.oxy | 24 | NC 13 | 156 | 175 | 5.671 | 0.97 | 0.97 | 196 | 136.6 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6711 | 5.6711 |  | 8.1
2012-150024.oxy | 24 | NC 13 | 157 | 150 | 5.661 | 0.924 | 0.9247 | 371 | 130.42 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6615 | 5.6658 |  | 8.1
2012-150024.oxy | 24 | NC 13 | 158 | 125 | 5.524 | 0.918 | 0.9167 | 396 | 132.76 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5239 | 5.5161 |  | 8.1
2012-150024.oxy | 24 | NC 13 | 158 | 125 | 5.521 | 0.932 | 0.9312 | 468 | 134.84 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5206 | 5.5158 |  | 8.1
2012-150024.oxy | 24 | NC 13 | 159 | 100 | 5.397 | 0.909 | 0.909 | 471 | 134.51 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.3974 | 5.3974 |  | 8.2
2012-150024.oxy | 24 | NC 13 | 160 | 75 | 5.335 | 0.873 | 0.8733 | 505 | 130.73 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.3353 | 5.3371 |  | 8.3
2012-150024.oxy | 24 | NC 13 | 161 | 50 | 5.356 | 0.898 | 0.8985 | 510 | 133.91 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.3562 | 5.3592 |  | 8.3
2012-150024.oxy | 24 | NC 13 | 162 | 30 | 5.353 | 0.87 | 0.869 | 528 | 129.86 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.353 | 5.3469 |  | 8.3
2012-150024.oxy | 24 | NC 13 | 163 | 20 | 5.333 | 0.88 | 0.8795 | 552 | 131.82 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.333 | 5.33 |  | 8.3
2012-150024.oxy | 24 | NC 13 | 164 | 10 | 5.33 | 0.904 | 0.9048 | 557 | 135.45 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.3298 | 5.3346 |  | 8.3
2012-150024.oxy | 24 | NC 13 | 165 | 0 | 5.843 | 0.993 | 0.9922 | 559 | 135.75 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8428 | 5.8381 |  | 8.6
2012-150035.oxy | 35 | NC 23 | 166 | 57 | 5.533 | 0.922 | 0.9225 | 94 | 133.13 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5324 | 5.5354 |  | 8.2
2012-150035.oxy | 35 | NC 23 | 167 | 30 | 5.519 | 0.944 | 0.944 | 196 | 136.6 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5187 | 5.5187 |  | 8.2
2012-150035.oxy | 35 | NC 23 | 168 | 20 | 5.532 | 0.903 | 0.9025 | 371 | 130.42 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5324 | 5.5294 |  | 8.2
2012-150035.oxy | 35 | NC 23 | 169 | 10 | 5.584 | 0.928 | 0.928 | 396 | 132.76 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5843 | 5.5843 |  | 8.3
2012-150035.oxy | 35 | NC 23 | 170 | 0 | 5.723 | 0.966 | 0.9661 | 468 | 134.84 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7225 | 5.7231 |  | 8.3
2012-150036.oxy | 36 | NC 24 | 171 | 55 | 5.538 | 0.906 | 0.906 | 505 | 130.73 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5375 | 5.5375 |  | 8.3
2012-150036.oxy | 36 | NC 24 | 172 | 50 | 5.524 | 0.926 | 0.9258 | 510 | 133.91 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5236 | 5.5224 |  | 8.2
2012-150036.oxy | 36 | NC 24 | 173 | 30 | 5.501 | 0.894 | 0.894 | 528 | 129.86 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5011 | 5.5011 |  | 8.2
2012-150036.oxy | 36 | NC 24 | 174 | 20 | 5.527 | 0.912 | 0.9124 | 552 | 131.82 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5275 | 5.5299 |  | 8.3
2012-150036.oxy | 36 | NC 24 | 175 | 10 | 5.614 | 0.952 | 0.952 | 557 | 135.45 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6136 | 5.6136 |  | 8.3
2012-150036.oxy | 36 | NC 24 | 176 | 0 | 5.766 | 0.98 | 0.9795 | 559 | 135.75 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7662 | 5.7632 |  | 8.3
2012-150039.oxy | 39 | NC 14 | 177 | 63 | 5.567 | 0.953 | 0.953 | 569 | 136.71 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5669 | 5.5669 |  | 8.2
2012-150039.oxy | 39 | NC 14 | 178 | 30 | 5.506 | 0.97 | 0.97 | 570 | 140.63 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5063 | 5.5063 |  | 8.2
2012-150039.oxy | 39 | NC 14 | 179 | 20 | 5.526 | 0.911 | 0.9108 | 611 | 131.71 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5261 | 5.5249 |  | 8.1
2012-150039.oxy | 39 | NC 14 | 180 | 10 | 5.611 | 0.939 | 0.9385 | 620 | 133.69 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6107 | 5.6077 |  | 8.1
2012-150039.oxy | 39 | NC 14 | 181 | 0 | 5.795 | 0.945 | 0.9448 | 621 | 130.31 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7954 | 5.7942 |  | 8.3
2012-150040.oxy | 40 | NC 15 | 182 | 95 | 5.582 | 0.947 | 0.9473 | 657 | 135.49 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5824 | 5.5841 |  | 8.1
2012-150040.oxy | 40 | NC 15 | 183 | 75 | 5.575 | 0.908 | 0.9076 | 668 | 130.14 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5753 | 5.5728 |  | 8.1
2012-150040.oxy | 40 | NC 15 | 184 | 50 | 5.548 | 0.918 | 0.918 | 690 | 132.2 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5477 | 5.5477 |  | 8.1
2012-150040.oxy | 40 | NC 15 | 185 | 30 | 5.506 | 0.965 | 0.9653 | 708 | 139.92 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.506 | 5.5077 |  | 8.1
2012-150040.oxy | 40 | NC 15 | 186 | 20 | 5.518 | 0.894 | 0.894 | 712 | 129.47 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.518 | 5.518 |  | 8.2
2012-150040.oxy | 40 | NC 15 | 187 | 10 | 5.617 | 0.977 | 0.977 | 715 | 138.88 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.617 | 5.617 |  | 8.3
2012-150040.oxy | 40 | NC 15 | 188 | 0 | 5.768 | 0.939 | 0.939 | 745 | 130.09 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7684 | 5.7684 |  | 8.5
2012-150041.oxy | 41 | NC 16 | 189 | 165 | 5.725 | 1.032 | 1.032 | 1035 | 143.89 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7245 | 5.7245 |  | 8.1
2012-150041.oxy | 41 | NC 16 | 190 | 100 | 5.612 | 0.992 | 0.9922 | 1055 | 141.1 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6124 | 5.6136 |  | 8.2
2012-150041.oxy | 41 | NC 16 | 191 | 75 | 5.524 | 0.982 | 0.982 | 1059 | 141.91 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5236 | 5.5236 |  | 8.1
2012-150041.oxy | 41 | NC 16 | 192 | 50 | 5.465 | 0.904 | 0.9035 | 1085 | 132.16 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4645 | 5.4615 |  | 8.1
2012-150041.oxy | 41 | NC 16 | 193 | 30 | 5.455 | 0.871 | 0.871 | 1105 | 127.62 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4548 | 5.4548 |  | 8.1
2012-150041.oxy | 41 | NC 16 | 194 | 20 | 5.484 | 0.911 | 0.9102 | 1106 | 132.7 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4842 | 5.4794 |  | 8.1
2012-150041.oxy | 41 | NC 16 | 195 | 10 | 5.529 | 0.941 | 0.94 | 1123 | 135.93 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5287 | 5.5228 |  | 8.1
2012-150041.oxy | 41 | NC 16 | 196 | 0 | 5.621 | 0.936 | 0.9356 | 1128 | 133.02 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6214 | 5.619 |  | 8.3
2012-150042.oxy | 42 | NC 17 | 197 | 221 | 5.409 | 0.92 | 0.92 | 1213 | 135.84 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4086 | 5.4086 | 4 | 8.1 | OXY: Titration completed at 75.5%. Endpoint accuracy questionable. Value seems low in depth profile. Outlier in COMPARE (by 0.37mL/L)
2012-150042.oxy | 42 | NC 17 | 198 | 175 | 5.494 | 0.992 | 0.992 | 1218 | 144.11 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4936 | 5.4936 | 4 | 8.1 | OXY: Titrationonly up to 94%. No end point. Looks to have been something blocking the light path. Outlier in COMPARE (by 0.23mL/L)
2012-150042.oxy | 42 | NC 17 | 199 | 150 | 5.67 | 1.024 | 1.024 | 1247 | 144.14 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.67 | 5.67 |  | 8.1
2012-150042.oxy | 42 | NC 17 | 199 | 150 | 5.66 | 1.031 | 1.0315 | 1248 | 145.38 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6595 | 5.6623 |  | 8
2012-150042.oxy | 42 | NC 17 | 200 | 125 | 5.613 | 0.925 | 0.925 | 1261 | 131.67 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.613 | 5.613 |  | 8
2012-150042.oxy | 42 | NC 17 | 201 | 100 | 5.543 | 0.99 | 0.99 | 1268 | 142.57 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5425 | 5.5425 | 4 | 8.1 | OXY: Titration incomplete. No end point. Looks to have been something blocking the light path
2012-150042.oxy | 42 | NC 17 | 202 | 75 | 5.424 | 0.971 | 0.971 | 1298 | 142.89 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4235 | 5.4235 | 4 | 8.1 | OXY: Titration incomplete. No end point. Slight outlier in COMPARE (by 0.08 mL/L)
2012-150042.oxy | 42 | NC 17 | 203 | 50 | 5.455 | 0.935 | 0.9348 | 1300 | 136.87 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.455 | 5.4539 |  | 8.1
2012-150042.oxy | 42 | NC 17 | 204 | 30 | 5.414 | 0.902 | 0.902 | 1306 | 133.09 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4137 | 5.4137 |  | 8.1
2012-150042.oxy | 42 | NC 17 | 205 | 20 | 5.454 | 0.905 | 0.905 | 1318 | 132.55 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4543 | 5.4543 |  | 8.2
2012-150042.oxy | 42 | NC 17 | 206 | 10 | 5.724 | 0.964 | 0.9637 | 1309 | 134.52 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7244 | 5.7226 |  | 8.4
2012-150042.oxy | 42 | NC 17 | 207 | 0 | 5.787 | 0.943 | 0.9428 | 1311 | 130.22 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7872 | 5.786 |  | 8.5
2012-150048.oxy | 48 | NC 19 | 208 | 201 | 5.868 | 0.998 | 0.9975 | 1213 | 135.84 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8684 | 5.8654 |  | 8
2012-150048.oxy | 48 | NC 19 | 209 | 175 | 5.827 | 1.052 | 1.0524 | 1218 | 144.11 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8267 | 5.8289 |  | 8
2012-150048.oxy | 48 | NC 19 | 210 | 150 | 5.809 | 1.049 | 1.0495 | 1247 | 144.14 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8088 | 5.8116 |  | 8
2012-150048.oxy | 48 | NC 19 | 211 | 125 | 5.808 | 1.058 | 1.058 | 1248 | 145.38 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8081 | 5.8081 |  | 8
2012-150048.oxy | 48 | NC 19 | 212 | 100 | 5.723 | 0.943 | 0.9428 | 1261 | 131.67 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7225 | 5.7213 |  | 8.1
2012-150048.oxy | 48 | NC 19 | 213 | 75 | 5.632 | 1.006 | 1.0068 | 1268 | 142.57 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6323 | 5.6368 |  | 8.1
2012-150048.oxy | 48 | NC 19 | 214 | 50 | 5.586 | 1 | 0.9995 | 1298 | 142.89 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5859 | 5.5831 |  | 8.1
2012-150048.oxy | 48 | NC 19 | 215 | 30 | 5.56 | 0.953 | 0.953 | 1300 | 136.87 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5603 | 5.5603 |  | 8.1
2012-150048.oxy | 48 | NC 19 | 215 | 30 | 5.558 | 0.926 | 0.9257 | 1306 | 133.09 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5582 | 5.5564 |  | 8.1
2012-150048.oxy | 48 | NC 19 | 216 | 20 | 5.533 | 0.918 | 0.9182 | 1318 | 132.55 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5328 | 5.534 |  | 8.1
2012-150048.oxy | 48 | NC 19 | 217 | 10 | 5.486 | 0.924 | 0.9235 | 1309 | 134.52 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4863 | 5.4833 |  | 8.1
2012-150048.oxy | 48 | NC 19 | 218 | 0 | 5.59 | 0.911 | 0.911 | 1311 | 130.22 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5903 | 5.5903 |  | 8.3
2012-150050.oxy | 50 | NC 20 | 219 | 152 | 5.825 | 0.988 | 0.9878 | 657 | 135.49 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8247 | 5.8235 |  | 8.1
2012-150050.oxy | 50 | NC 20 | 220 | 150 | 5.821 | 0.948 | 0.9476 | 668 | 130.14 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8216 | 5.8191 |  | 8
2012-150050.oxy | 50 | NC 20 | 221 | 125 | 5.778 | 0.956 | 0.955 | 690 | 132.2 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.778 | 5.7719 |  | 8.1
2012-150050.oxy | 50 | NC 20 | 222 | 100 | 5.718 | 1.002 | 1.0018 | 708 | 139.92 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7177 | 5.7165 |  | 8.1
2012-150050.oxy | 50 | NC 20 | 222 | 100 | 5.71 | 0.925 | 0.9247 | 712 | 129.47 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7098 | 5.708 |  | 8.1
2012-150050.oxy | 50 | NC 20 | 223 | 75 | 5.669 | 0.986 | 0.986 | 715 | 138.88 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6689 | 5.6689 |  | 8.1
2012-150050.oxy | 50 | NC 20 | 224 | 50 | 5.639 | 0.918 | 0.916 | 745 | 130.09 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6391 | 5.6268 | 3 | 8.1 | OXY: Endpoint ambiguous but it looked to be quite late. Adjusted
2012-150050.oxy | 50 | NC 20 | 225 | 30 | 5.615 | 1.012 | 1.0124 | 1002 | 143.85 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6149 | 5.6171 |  | 8.1
2012-150050.oxy | 50 | NC 20 | 226 | 20 | 5.586 | 1.007 | 1.007 | 1035 | 143.89 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5855 | 5.5855 |  | 8.1
2012-150050.oxy | 50 | NC 20 | 227 | 10 | 5.618 | 0.993 | 0.993 | 1055 | 141.1 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6181 | 5.6181 |  | 8.1
2012-150050.oxy | 50 | NC 20 | 228 | 0 | 5.834 | 1.037 | 1.037 | 1059 | 141.91 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8337 | 5.8337 |  | 8.2
2012-150051.oxy | 51 | NC 21 | 229 | 107 | 5.707 | 0.951 | 0.951 | 94 | 133.13 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7069 | 5.7069 |  | 8.1
2012-150051.oxy | 51 | NC 21 | 230 | 100 | 5.87 | 0.978 | 0.9748 | 196 | 136.6 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.718 | 5.6993 |  | 8.1 | Wrong sample number, flask ID, flask volume and depth (Corrected). End point was crazy late too. Adjusted
2012-150051.oxy | 51 | NC 21 | 231 | 75 | 5.661 | 0.924 | 0.924 | 371 | 130.42 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6615 | 5.6615 |  | 8.2
2012-150051.oxy | 51 | NC 21 | 232 | 50 | 5.639 | 0.937 | 0.937 | 396 | 132.76 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6386 | 5.6386 |  | 8.2
2012-150051.oxy | 51 | NC 21 | 233 | 30 | 5.639 | 0.952 | 0.9515 | 468 | 134.84 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6394 | 5.6364 |  | 8.2
2012-150051.oxy | 51 | NC 21 | 234 | 20 | 5.641 | 0.95 | 0.9492 | 471 | 134.51 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6415 | 5.6367 |  | 8.2
2012-150051.oxy | 51 | NC 21 | 235 | 10 | 5.881 | 0.962 | 0.9605 | 505 | 130.73 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8807 | 5.8715 |  | 8.3
2012-150051.oxy | 51 | NC 21 | 236 | 0 | 6.397 | 1.072 | 1.0717 | 510 | 133.91 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 6.3968 | 6.395 |  | 8.6
2012-150052.oxy | 52 | NC 22 | 237 | 45 | 5.673 | 0.936 | 0.9364 | 552 | 131.82 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6733 | 5.6758 |  | 8.3
2012-150052.oxy | 52 | NC 22 | 238 | 45 | 5.637 | 0.956 | 0.9555 | 557 | 135.45 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6372 | 5.6343 |  | 8.2
2012-150052.oxy | 52 | NC 22 | 239 | 30 | 5.636 | 0.958 | 0.958 | 559 | 135.75 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6364 | 5.6364 |  | 8.2
2012-150052.oxy | 52 | NC 22 | 240 | 20 | 5.708 | 0.977 | 0.9765 | 569 | 136.71 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7075 | 5.7046 |  | 8.2
2012-150052.oxy | 52 | NC 22 | 241 | 10 | 5.893 | 1.038 | 1.0375 | 570 | 140.63 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8933 | 5.8904 |  | 8.3
2012-150052.oxy | 52 | NC 22 | 242 | 0 | 6.524 | 1.075 | 1.075 | 611 | 131.71 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 6.5236 | 6.5236 |  | 8.6
2012-150057.oxy | 57 | NC 25 | 243 | 126 | 5.629 | 0.938 | 0.938 | 94 | 133.13 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6287 | 5.6287 |  | 8.1
2012-150057.oxy | 57 | NC 25 | 244 | 100 | 5.607 | 0.959 | 0.959 | 196 | 136.6 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6067 | 5.6067 |  | 8.1
2012-150057.oxy | 57 | NC 25 | 245 | 75 | 5.557 | 0.907 | 0.907 | 371 | 130.42 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.557 | 5.557 |  | 8.1
2012-150057.oxy | 57 | NC 25 | 246 | 50 | 5.458 | 0.907 | 0.907 | 396 | 132.76 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4576 | 5.4576 |  | 8.1
2012-150057.oxy | 57 | NC 25 | 247 | 30 | 5.426 | 0.916 | 0.9145 | 468 | 134.84 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4256 | 5.4167 |  | 8.1
2012-150057.oxy | 57 | NC 25 | 248 | 20 | 5.403 | 0.91 | 0.91 | 471 | 134.51 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4034 | 5.4034 |  | 8.2
2012-150057.oxy | 57 | NC 25 | 249 | 10 | 5.746 | 0.94 | 0.94 | 505 | 130.73 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7459 | 5.7459 |  | 8.3
2012-150057.oxy | 57 | NC 25 | 250 | 0 | 6.169 | 1.034 | 1.033 | 510 | 133.91 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 6.1696 | 6.1636 |  | 8.5
2012-150061.oxy | 61 | NC 26 | 251 | 132 | 5.722 | 0.944 | 0.9435 | 552 | 131.82 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7219 | 5.7189 |  | 8.1
2012-150061.oxy | 61 | NC 26 | 252 | 100 | 5.637 | 0.956 | 0.9555 | 557 | 135.45 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6372 | 5.6343 |  | 8.1
2012-150061.oxy | 61 | NC 26 | 253 | 75 | 5.542 | 0.942 | 0.942 | 559 | 135.75 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.542 | 5.542 |  | 8.1
2012-150061.oxy | 61 | NC 26 | 254 | 50 | 5.514 | 0.944 | 0.9443 | 569 | 136.71 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5142 | 5.516 |  | 8.1
2012-150061.oxy | 61 | NC 26 | 255 | 30 | 5.438 | 0.958 | 0.9586 | 570 | 140.63 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.438 | 5.4414 |  | 8.2
2012-150061.oxy | 61 | NC 26 | 256 | 20 | 5.423 | 0.894 | 0.895 | 611 | 131.71 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4227 | 5.4288 |  | 8.2
2012-150061.oxy | 61 | NC 26 | 257 | 10 | 5.491 | 0.919 | 0.919 | 620 | 133.69 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4909 | 5.4909 |  | 8.2
2012-150061.oxy | 61 | NC 26 | 258 | 0 | 5.758 | 0.939 | 0.939 | 621 | 130.31 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7585 | 5.7585 |  | 8.4
2012-150066.oxy | 66 | NC 27 | 259 | 174 | 5.695 | 0.966 | 0.9663 | 657 | 135.49 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6947 | 5.6964 |  | 8.3
2012-150066.oxy | 66 | NC 27 | 260 | 125 | 5.674 | 0.924 | 0.923 | 668 | 130.14 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6738 | 5.6677 |  | 8.3
2012-150066.oxy | 66 | NC 27 | 261 | 100 | 5.657 | 0.936 | 0.9359 | 690 | 132.2 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6568 | 5.6562 |  | 8.3
2012-150066.oxy | 66 | NC 27 | 262 | 75 | 5.638 | 0.988 | 0.9878 | 708 | 139.92 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6376 | 5.6364 |  | 8.3
2012-150066.oxy | 66 | NC 27 | 263 | 50 | 5.5 | 0.891 | 0.8906 | 712 | 129.47 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4994 | 5.4969 |  | 8.3
2012-150066.oxy | 66 | NC 27 | 264 | 30 | 5.41 | 0.941 | 0.941 | 715 | 138.88 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4095 | 5.4095 |  | 8.4
2012-150066.oxy | 66 | NC 27 | 265 | 20 | 5.46 | 0.889 | 0.889 | 745 | 130.09 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4605 | 5.4605 |  | 8.4
2012-150066.oxy | 66 | NC 27 | 266 | 10 | 5.57 | 1.004 | 1.0038 | 1002 | 143.85 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.5704 | 5.5693 |  | 8.5
2012-150066.oxy | 66 | NC 27 | 267 | 0 | 5.853 | 1.055 | 1.0547 | 1035 | 143.89 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.8524 | 5.8507 |  | 8.6
2012-150067.oxy | 67 | NC 28 | 268 | 68 | 5.445 | 0.968 | 0.968 | 1059 | 141.91 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4446 | 5.4446 | 3 | -99 | OXY: Seems low. Not accurate. Add 1 + 2 titer. Zeroed colorimeter in sample. No end point. Slight outlier in COMPARE (by 0.07mL/L)
2012-150067.oxy | 67 | NC 28 | 269 | 50 | 5.489 | 0.908 | 0.9082 | 1085 | 132.16 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4888 | 5.49 |  | 8.4
2012-150067.oxy | 67 | NC 28 | 270 | 30 | 5.43 | 0.867 | 0.8669 | 1105 | 127.62 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4297 | 5.429 |  | 8.4
2012-150067.oxy | 67 | NC 28 | 271 | 20 | 5.466 | 0.908 | 0.908 | 1106 | 132.7 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4661 | 5.4661 |  | 8.5
2012-150067.oxy | 67 | NC 28 | 272 | 10 | 5.664 | 0.964 | 0.9643 | 1123 | 135.93 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6642 | 5.6659 |  | 8.6
2012-150067.oxy | 67 | NC 28 | 273 | 0 | 5.784 | 0.963 | 0.962 | 1128 | 133.02 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.7839 | 5.7779 |  | 8.6
2012-150068.oxy | 68 | NC 29 | 274 | 96 | 5.603 | 0.948 | 0.948 | 1335 | 135.15 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6025 | 5.6025 |  | 8.4
2012-150068.oxy | 68 | NC 29 | 275 | 50 | 5.439 | 0.968 | 0.968 | 1337 | 142.06 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4388 | 5.4388 |  | 8.4
2012-150068.oxy | 68 | NC 29 | 276 | 30 | 5.433 | 0.893 | 0.8925 | 1338 | 131.33 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4325 | 5.4295 |  | 8.5
2012-150068.oxy | 68 | NC 29 | 277 | 20 | 5.64 | 0.927 | 0.927 | 1365 | 135.728 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.4544 | 5.4544 |  | 8.5 | Wrong sample #, Flask ID and volume. Corrected
2012-150068.oxy | 68 | NC 29 | 278 | 10 | 5.315 | 0.946 | 0.946 | 1337 | 142.06 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.3149 | 5.3149 |  | 8.5
2012-150068.oxy | 68 | NC 29 | 279 | 0 | 5.648 | 0.956 | 0.9557 | 1373 | 135.2 | 0.0003 | 0.862 | 10.0277 | 0.012111 | 0.1411 | 5.6478 | 5.646 |  | 8.6

## Raw AutoOxy Data
File Name | LOC:EVENT_NUMBER | LOC:STATION | Sample_number | Depth | Oxygen:Dissolved | Thiosulfate_titer | Flask_Number | Flask_Volume | Blank | Standard
2012-150001.oxy | 1 | NC 01 | 2 | 300 | 0 | 0 | 1421 | 140.09 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 2 | 300 | 5.679 | 0.997 | 1421 | 140.09 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 3 | 250 | 5.756 | 0.958 | 1459 | 132.92 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 4 | 200 | 5.748 | 0.944 | 1464 | 131.17 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 5 | 175 | 5.754 | 1.014 | 1465 | 140.64 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 6 | 150 | 5.742 | 0.961 | 1470 | 133.64 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 7 | 125 | 5.745 | 1.031 | 1472 | 143.19 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 8 | 100 | 5.76 | 0.95 | 1473 | 131.73 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 9 | 75 | 5.752 | 0.938 | 1475 | 130.27 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 9 | 75 | 5.751 | 0.954 | 1476 | 132.49 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 10 | 50 | 5.751 | 0.968 | 1477 | 134.4 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 11 | 30 | 5.801 | 1.03 | 1478 | 141.68 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 12 | 20 | 5.716 | 0.943 | 1481 | 131.76 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 13 | 10 | 5.699 | 0.945 | 1483 | 132.43 | 0.0013 | 0.8625
2012-150001.oxy | 1 | NC 01 | 14 | 0 | 5.681 | 0.947 | 1485 | 133.1 | 0.0013 | 0.8625
2012-150001_1.oxy | 1 | NC 01 | 1 | 335 | 1.502 | 0.252 | 1414 | 132.62 | 0.0013 | 0.8625
2012-150001_1.oxy | 1 | NC 01 | 1 | 335 | 3.14 | 0.523 | 1414 | 132.62 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 15 | 318 | 5.834 | 0.99 | 657 | 135.49 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 16 | 300 | 5.825 | 0.949 | 668 | 130.14 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 17 | 250 | 5.793 | 0.959 | 690 | 132.2 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 18 | 200 | 5.795 | 1.016 | 708 | 139.92 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 19 | 175 | 5.788 | 0.938 | 712 | 129.47 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 20 | 150 | 5.782 | 1.006 | 715 | 138.88 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 21 | 100 | 5.772 | 0.94 | 745 | 130.09 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 22 | 75 | 5.757 | 1.038 | 1002 | 143.85 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 23 | 50 | 5.734 | 1.034 | 1035 | 143.89 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 24 | 30 | 5.684 | 1.005 | 1055 | 141.1 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 24 | 30 | 5.674 | 1.009 | 1059 | 141.91 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 25 | 20 | 5.638 | 0.933 | 1085 | 132.16 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 26 | 10 | 5.596 | 0.894 | 1105 | 127.62 | 0.0013 | 0.8625
2012-150002.oxy | 2 | NC 02 | 27 | 0 | 5.747 | 0.955 | 1106 | 132.7 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 28 | 279 | 6.349 | 1.058 | 94 | 133.13 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 29 | 200 | 5.774 | 0.988 | 196 | 136.6 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 30 | 175 | 5.769 | 0.942 | 371 | 130.42 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 30 | 175 | 5.425 | 0.902 | 396 | 132.76 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 31 | 150 | 4.745 | 0.802 | 468 | 134.84 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 32 | 125 | 5.776 | 0.973 | 471 | 134.51 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 33 | 100 | 5.78 | 0.946 | 505 | 130.73 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 34 | 75 | 5.748 | 0.964 | 510 | 133.91 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 35 | 50 | 5.671 | 0.922 | 528 | 129.86 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 36 | 30 | 5.476 | 0.904 | 552 | 131.82 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 37 | 20 | 5.493 | 0.932 | 557 | 135.45 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 38 | 10 | 5.569 | 0.947 | 559 | 135.75 | 0.0013 | 0.8625
2012-150003.oxy | 3 | NC 03 | 39 | 0 | 5.957 | 1.02 | 569 | 136.71 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 40 | 251 | 5.807 | 0.983 | 1335 | 135.15 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 41 | 200 | 5.769 | 1.027 | 1337 | 142.06 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 42 | 175 | 5.777 | 0.95 | 1338 | 131.33 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 43 | 150 | 5.759 | 0.979 | 1365 | 135.73 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 44 | 125 | 5.738 | 0.98 | 1367 | 136.34 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 45 | 100 | 5.722 | 0.969 | 1373 | 135.2 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 46 | 75 | 5.608 | 0.945 | 1374 | 134.53 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 46 | 75 | 5.604 | 0.924 | 1377 | 131.67 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 47 | 50 | 5.536 | 0.919 | 1378 | 132.54 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 48 | 30 | 5.469 | 0.921 | 1386 | 134.43 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 49 | 20 | 5.555 | 0.941 | 1387 | 135.22 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 50 | 10 | 5.701 | 0.977 | 1389 | 136.81 | 0.0013 | 0.8625
2012-150004.oxy | 4 | NC 4 | 51 | 0 | 6.07 | 1 | 1392 | 131.6 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 52 | 218 | 5.772 | 0.977 | 1335 | 135.15 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 53 | 200 | 5.769 | 1.027 | 1337 | 142.06 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 54 | 175 | 5.765 | 0.948 | 1338 | 131.33 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 55 | 150 | 5.753 | 0.978 | 1365 | 135.73 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 56 | 125 | 5.709 | 0.975 | 1367 | 136.34 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 57 | 100 | 5.663 | 0.959 | 1373 | 135.2 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 58 | 75 | 5.513 | 0.929 | 1374 | 134.53 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 59 | 50 | 5.483 | 0.904 | 1377 | 131.67 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 60 | 30 | 5.415 | 0.899 | 1378 | 132.54 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 61 | 20 | 5.422 | 0.913 | 1386 | 134.43 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 61 | 20 | 5.407 | 0.916 | 1387 | 135.22 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 62 | 10 | 5.502 | 0.943 | 1389 | 136.81 | 0.0013 | 0.8625
2012-150009.oxy | 9 | NC 5 | 63 | 0 | 5.668 | 0.934 | 1392 | 131.6 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 64 | 226 | 5.787 | 0.961 | 1414 | 132.62 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 65 | 200 | 5.782 | 1.015 | 1421 | 140.09 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 65 | 200 | 5.786 | 0.963 | 1459 | 132.92 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 66 | 175 | 5.778 | 0.949 | 1464 | 131.17 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 67 | 175 | 5.748 | 1.013 | 1465 | 140.64 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 68 | 125 | 5.73 | 0.959 | 1470 | 133.64 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 69 | 100 | 5.7 | 1.023 | 1472 | 143.19 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 70 | 75 | 5.626 | 0.928 | 1473 | 131.73 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 71 | 50 | 5.505 | 0.898 | 1475 | 130.27 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 72 | 30 | 5.406 | 0.897 | 1476 | 132.49 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 73 | 20 | 5.423 | 0.913 | 1477 | 134.4 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 74 | 10 | 5.53 | 0.982 | 1478 | 141.68 | 0.0013 | 0.8625
2012-150010.oxy | 10 | NC 6 | 75 | 0 | 5.625 | 0.928 | 1481 | 131.76 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 76 | 226 | 5.789 | 0.965 | 94 | 133.13 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 77 | 200 | 5.757 | 0.985 | 196 | 136.6 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 77 | 200 | 5.769 | 0.942 | 371 | 130.42 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 78 | 175 | 5.769 | 0.959 | 396 | 132.76 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 79 | 150 | 5.738 | 0.969 | 468 | 134.84 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 80 | 125 | 5.71 | 0.962 | 471 | 134.51 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 81 | 100 | 5.627 | 0.921 | 505 | 130.73 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 82 | 75 | 5.623 | 0.943 | 510 | 133.91 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 83 | 50 | 5.443 | 0.913 | 510 | 133.91 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 84 | 30 | 5.397 | 0.891 | 552 | 131.82 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 85 | 20 | 5.463 | 0.927 | 557 | 135.45 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 86 | 10 | 5.486 | 0.933 | 559 | 135.75 | 0.0013 | 0.8625
2012-150011.oxy | 11 | NC 7 | 87 | 0 | 5.547 | 0.95 | 569 | 136.71 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 88 | 255 | 5.766 | 0.939 | 745 | 130.09 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 89 | 200 | 5.727 | 0.972 | 657 | 135.49 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 89 | 200 | 5.726 | 0.933 | 668 | 130.14 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 90 | 175 | 5.678 | 0.94 | 690 | 132.2 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 91 | 150 | 5.664 | 0.993 | 708 | 139.92 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 92 | 125 | 5.651 | 0.916 | 712 | 129.47 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 93 | 100 | 5.626 | 0.979 | 715 | 138.88 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 94 | 75 | 5.484 | 0.989 | 1002 | 143.85 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 95 | 50 | 5.45 | 0.983 | 1035 | 143.89 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 96 | 30 | 5.383 | 0.952 | 1055 | 141.1 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 97 | 20 | 5.358 | 0.953 | 1059 | 141.91 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 98 | 10 | 5.49 | 0.877 | 1105 | 127.62 | 0.0013 | 0.8625
2012-150012.oxy | 12 | NC 8 | 99 | 10 | 5.638 | 0.933 | 1085 | 132.16 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 100 | 278 | 5.777 | 0.983 | 1213 | 135.84 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 101 | 250 | 5.775 | 1.043 | 1218 | 144.11 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 102 | 200 | 5.712 | 1.032 | 1247 | 144.14 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 102 | 200 | 5.757 | 1.04 | 1247 | 144.14 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 103 | 175 | 5.677 | 0.936 | 1261 | 131.67 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 104 | 150 | 5.664 | 1.012 | 1268 | 142.57 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 105 | 125 | 5.64 | 1.01 | 1298 | 142.89 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 106 | 100 | 5.575 | 0.956 | 1300 | 136.87 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 107 | 75 | 5.519 | 0.92 | 1306 | 133.09 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 108 | 50 | 5.447 | 0.908 | 1307 | 133.08 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 109 | 30 | 5.424 | 0.914 | 1309 | 134.52 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 110 | 20 | 5.421 | 0.884 | 1311 | 130.22 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 111 | 10 | 5.477 | 0.912 | 1312 | 132.94 | 0.0013 | 0.8625
2012-150013.oxy | 13 | NC 9 | 112 | 0 | 5.62 | 0.96 | 1313 | 136.36 | 0.0013 | 0.8625
2012-150019.oxy | 19 | NC 10 | 113 | 306 | 5.757 | 0.974 | 1335 | 135.15 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 114 | 250 | 5.715 | 1.017 | 1337 | 142.06 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 115 | 200 | 5.683 | 0.934 | 1338 | 131.33 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 116 | 175 | 5.667 | 0.963 | 1365 | 135.73 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 117 | 150 | 5.659 | 0.966 | 1367 | 136.34 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 118 | 125 | 5.6 | 0.948 | 1373 | 135.2 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 119 | 100 | 5.522 | 0.93 | 1374 | 134.53 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 120 | 75 | 5.455 | 0.899 | 1377 | 131.67 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 120 | 75 | 5.466 | 0.97 | 1393 | 141.65 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 121 | 50 | 5.467 | 0.907 | 1378 | 132.54 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 122 | 30 | 5.43 | 0.914 | 1386 | 134.43 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 123 | 20 | 5.416 | 0.917 | 1387 | 135.22 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 124 | 10 | 5.434 | 0.931 | 1389 | 136.81 | 0.0003 | 0.862
2012-150019.oxy | 19 | NC 10 | 125 | 0 | 5.229 | 0.896 | 1389 | 136.81 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 126 | 322 | 5.75 | 0.978 | 1213 | 135.84 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 127 | 322 | 5.755 | 1.039 | 1218 | 144.11 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 128 | 300 | 5.753 | 1.039 | 1247 | 144.14 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 129 | 250 | 5.588 | 1.018 | 1248 | 145.38 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 130 | 200 | 5.717 | 0.942 | 1261 | 131.67 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 131 | 175 | 5.689 | 1.016 | 1268 | 142.57 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 132 | 150 | 5.659 | 1.013 | 1298 | 142.89 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 133 | 125 | 5.619 | 0.963 | 1300 | 136.87 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 134 | 100 | 5.51 | 0.918 | 1306 | 133.09 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 135 | 75 | 5.462 | 0.91 | 1307 | 133.08 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 135 | 75 | 5.445 | 0.917 | 1309 | 134.52 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 136 | 50 | 5.4 | 0.88 | 1311 | 130.22 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 137 | 30 | 5.426 | 0.903 | 1312 | 132.94 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 1371 | 20 | 5.476 | 0.935 | 1313 | 136.36 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 139 | 10 | 5.556 | 0.945 | 1314 | 135.85 | 0.0003 | 0.862
2012-150020.oxy | 20 | NC 11 | 140 | 0 | 5.739 | 0.961 | 1316 | 133.78 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 141 | 318 | 5.771 | 0.979 | 657 | 135.49 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 142 | 300 | 5.76 | 0.938 | 668 | 130.14 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 143 | 250 | 5.742 | 0.95 | 690 | 132.2 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 144 | 200 | 5.712 | 1.001 | 708 | 139.92 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 145 | 175 | 5.698 | 0.923 | 712 | 129.47 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 146 | 150 | 5.634 | 0.98 | 715 | 138.88 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 147 | 125 | 5.048 | 0.822 | 745 | 130.09 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 147 | 125 | 5.626 | 1.014 | 1002 | 143.85 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 148 | 100 | 5.513 | 0.994 | 1035 | 143.89 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 149 | 75 | 5.454 | 0.964 | 1055 | 141.1 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 150 | 50 | 5.45 | 0.969 | 1059 | 141.91 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 151 | 30 | 5.428 | 0.898 | 1085 | 132.16 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 152 | 20 | 5.455 | 0.871 | 1105 | 127.62 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 153 | 10 | 5.563 | 0.924 | 1106 | 132.7 | 0.0003 | 0.862
2012-150023.oxy | 23 | NC 12 | 154 | 0 | 5.635 | 0.959 | 1123 | 135.93 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 155 | 300 | 5.773 | 0.962 | 94 | 133.13 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 156 | 175 | 5.671 | 0.97 | 196 | 136.6 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 157 | 150 | 5.661 | 0.924 | 371 | 130.42 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 158 | 125 | 5.524 | 0.918 | 396 | 132.76 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 158 | 125 | 5.521 | 0.932 | 468 | 134.84 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 159 | 100 | 5.397 | 0.909 | 471 | 134.51 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 160 | 75 | 5.335 | 0.873 | 505 | 130.73 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 161 | 50 | 5.356 | 0.898 | 510 | 133.91 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 162 | 30 | 5.353 | 0.87 | 528 | 129.86 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 163 | 20 | 5.333 | 0.88 | 552 | 131.82 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 164 | 10 | 5.33 | 0.904 | 557 | 135.45 | 0.0003 | 0.862
2012-150024.oxy | 24 | NC 13 | 165 | 0 | 5.843 | 0.993 | 559 | 135.75 | 0.0003 | 0.862
2012-150035.oxy | 35 | NC 23 | 166 | 57 | 5.533 | 0.922 | 94 | 133.13 | 0.0003 | 0.862
2012-150035.oxy | 35 | NC 23 | 167 | 30 | 5.519 | 0.944 | 196 | 136.6 | 0.0003 | 0.862
2012-150035.oxy | 35 | NC 23 | 168 | 20 | 5.532 | 0.903 | 371 | 130.42 | 0.0003 | 0.862
2012-150035.oxy | 35 | NC 23 | 169 | 10 | 5.584 | 0.928 | 396 | 132.76 | 0.0003 | 0.862
2012-150035.oxy | 35 | NC 23 | 170 | 0 | 5.723 | 0.966 | 468 | 134.84 | 0.0003 | 0.862
2012-150036.oxy | 36 | NC 24 | 171 | 55 | 5.538 | 0.906 | 505 | 130.73 | 0.0003 | 0.862
2012-150036.oxy | 36 | NC 24 | 172 | 50 | 5.524 | 0.926 | 510 | 133.91 | 0.0003 | 0.862
2012-150036.oxy | 36 | NC 24 | 173 | 30 | 5.501 | 0.894 | 528 | 129.86 | 0.0003 | 0.862
2012-150036.oxy | 36 | NC 24 | 174 | 20 | 5.527 | 0.912 | 552 | 131.82 | 0.0003 | 0.862
2012-150036.oxy | 36 | NC 24 | 175 | 10 | 5.614 | 0.952 | 557 | 135.45 | 0.0003 | 0.862
2012-150036.oxy | 36 | NC 24 | 176 | 0 | 5.766 | 0.98 | 559 | 135.75 | 0.0003 | 0.862
2012-150039.oxy | 39 | NC 14 | 177 | 63 | 5.567 | 0.953 | 569 | 136.71 | 0.0003 | 0.862
2012-150039.oxy | 39 | NC 14 | 178 | 30 | 5.506 | 0.97 | 570 | 140.63 | 0.0003 | 0.862
2012-150039.oxy | 39 | NC 14 | 179 | 20 | 5.526 | 0.911 | 611 | 131.71 | 0.0003 | 0.862
2012-150039.oxy | 39 | NC 14 | 180 | 10 | 5.611 | 0.939 | 620 | 133.69 | 0.0003 | 0.862
2012-150039.oxy | 39 | NC 14 | 181 | 0 | 5.795 | 0.945 | 621 | 130.31 | 0.0003 | 0.862
2012-150040.oxy | 40 | NC 15 | 182 | 95 | 5.582 | 0.947 | 657 | 135.49 | 0.0003 | 0.862
2012-150040.oxy | 40 | NC 15 | 183 | 75 | 5.575 | 0.908 | 668 | 130.14 | 0.0003 | 0.862
2012-150040.oxy | 40 | NC 15 | 184 | 50 | 5.548 | 0.918 | 690 | 132.2 | 0.0003 | 0.862
2012-150040.oxy | 40 | NC 15 | 185 | 30 | 5.506 | 0.965 | 708 | 139.92 | 0.0003 | 0.862
2012-150040.oxy | 40 | NC 15 | 186 | 20 | 5.518 | 0.894 | 712 | 129.47 | 0.0003 | 0.862
2012-150040.oxy | 40 | NC 15 | 187 | 10 | 5.617 | 0.977 | 715 | 138.88 | 0.0003 | 0.862
2012-150040.oxy | 40 | NC 15 | 188 | 0 | 5.768 | 0.939 | 745 | 130.09 | 0.0003 | 0.862
2012-150041.oxy | 41 | NC 16 | 189 | 165 | 5.725 | 1.032 | 1035 | 143.89 | 0.0003 | 0.862
2012-150041.oxy | 41 | NC 16 | 190 | 100 | 5.612 | 0.992 | 1055 | 141.1 | 0.0003 | 0.862
2012-150041.oxy | 41 | NC 16 | 191 | 75 | 5.524 | 0.982 | 1059 | 141.91 | 0.0003 | 0.862
2012-150041.oxy | 41 | NC 16 | 192 | 50 | 5.465 | 0.904 | 1085 | 132.16 | 0.0003 | 0.862
2012-150041.oxy | 41 | NC 16 | 193 | 30 | 5.455 | 0.871 | 1105 | 127.62 | 0.0003 | 0.862
2012-150041.oxy | 41 | NC 16 | 194 | 20 | 5.484 | 0.911 | 1106 | 132.7 | 0.0003 | 0.862
2012-150041.oxy | 41 | NC 16 | 195 | 10 | 5.529 | 0.941 | 1123 | 135.93 | 0.0003 | 0.862
2012-150041.oxy | 41 | NC 16 | 196 | 0 | 5.621 | 0.936 | 1128 | 133.02 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 197 | 221 | 5.409 | 0.92 | 1213 | 135.84 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 198 | 175 | 5.494 | 0.992 | 1218 | 144.11 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 199 | 150 | 5.67 | 1.024 | 1247 | 144.14 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 199 | 150 | 5.66 | 1.031 | 1248 | 145.38 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 200 | 125 | 5.613 | 0.925 | 1261 | 131.67 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 201 | 100 | 5.543 | 0.99 | 1268 | 142.57 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 202 | 75 | 5.424 | 0.971 | 1298 | 142.89 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 203 | 50 | 5.455 | 0.935 | 1300 | 136.87 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 204 | 30 | 5.414 | 0.902 | 1306 | 133.09 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 205 | 20 | 5.454 | 0.905 | 1318 | 132.55 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 206 | 10 | 5.724 | 0.964 | 1309 | 134.52 | 0.0003 | 0.862
2012-150042.oxy | 42 | NC 17 | 207 | 0 | 5.787 | 0.943 | 1311 | 130.22 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 208 | 201 | 5.868 | 0.998 | 1213 | 135.84 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 209 | 175 | 5.827 | 1.052 | 1218 | 144.11 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 210 | 150 | 5.809 | 1.049 | 1247 | 144.14 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 211 | 125 | 5.808 | 1.058 | 1248 | 145.38 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 212 | 100 | 5.723 | 0.943 | 1261 | 131.67 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 213 | 75 | 5.632 | 1.006 | 1268 | 142.57 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 214 | 50 | 5.586 | 1 | 1298 | 142.89 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 215 | 30 | 5.56 | 0.953 | 1300 | 136.87 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 215 | 30 | 5.558 | 0.926 | 1306 | 133.09 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 216 | 20 | 5.533 | 0.918 | 1318 | 132.55 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 217 | 10 | 5.486 | 0.924 | 1309 | 134.52 | 0.0003 | 0.862
2012-150048.oxy | 48 | NC 19 | 218 | 0 | 5.59 | 0.911 | 1311 | 130.22 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 219 | 152 | 5.825 | 0.988 | 657 | 135.49 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 220 | 150 | 5.821 | 0.948 | 668 | 130.14 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 221 | 125 | 5.778 | 0.956 | 690 | 132.2 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 222 | 100 | 5.718 | 1.002 | 708 | 139.92 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 222 | 100 | 5.71 | 0.925 | 712 | 129.47 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 223 | 75 | 5.669 | 0.986 | 715 | 138.88 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 224 | 50 | 5.639 | 0.918 | 745 | 130.09 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 225 | 30 | 5.615 | 1.012 | 1002 | 143.85 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 226 | 20 | 5.586 | 1.007 | 1035 | 143.89 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 227 | 10 | 5.618 | 0.993 | 1055 | 141.1 | 0.0003 | 0.862
2012-150050.oxy | 50 | NC 20 | 228 | 0 | 5.834 | 1.037 | 1059 | 141.91 | 0.0003 | 0.862
2012-150051.oxy | 51 | NC 21 | 229 | 107 | 5.707 | 0.951 | 94 | 133.13 | 0.0003 | 0.862
2012-150051.oxy | 51 | NC 21 | 229 | 107 | 5.87 | 0.978 | 94 | 133.13 | 0.0003 | 0.862
2012-150051.oxy | 51 | NC 21 | 231 | 75 | 5.661 | 0.924 | 371 | 130.42 | 0.0003 | 0.862
2012-150051.oxy | 51 | NC 21 | 232 | 50 | 5.639 | 0.937 | 396 | 132.76 | 0.0003 | 0.862
2012-150051.oxy | 51 | NC 21 | 233 | 30 | 5.639 | 0.952 | 468 | 134.84 | 0.0003 | 0.862
2012-150051.oxy | 51 | NC 21 | 234 | 20 | 5.641 | 0.95 | 471 | 134.51 | 0.0003 | 0.862
2012-150051.oxy | 51 | NC 21 | 235 | 10 | 5.881 | 0.962 | 505 | 130.73 | 0.0003 | 0.862
2012-150051.oxy | 51 | NC 21 | 236 | 0 | 6.397 | 1.072 | 510 | 133.91 | 0.0003 | 0.862
2012-150052.oxy | 52 | NC 22 | 237 | 45 | 5.673 | 0.936 | 552 | 131.82 | 0.0003 | 0.862
2012-150052.oxy | 52 | NC 22 | 238 | 45 | 5.637 | 0.956 | 557 | 135.45 | 0.0003 | 0.862
2012-150052.oxy | 52 | NC 22 | 239 | 30 | 5.636 | 0.958 | 559 | 135.75 | 0.0003 | 0.862
2012-150052.oxy | 52 | NC 22 | 240 | 20 | 5.708 | 0.977 | 569 | 136.71 | 0.0003 | 0.862
2012-150052.oxy | 52 | NC 22 | 241 | 10 | 5.893 | 1.038 | 570 | 140.63 | 0.0003 | 0.862
2012-150052.oxy | 52 | NC 22 | 242 | 0 | 6.524 | 1.075 | 611 | 131.71 | 0.0003 | 0.862
2012-150057.oxy | 57 | NC 25 | 243 | 126 | 5.629 | 0.938 | 94 | 133.13 | 0.0003 | 0.862
2012-150057.oxy | 57 | NC 25 | 244 | 100 | 5.607 | 0.959 | 196 | 136.6 | 0.0003 | 0.862
2012-150057.oxy | 57 | NC 25 | 245 | 75 | 5.557 | 0.907 | 371 | 130.42 | 0.0003 | 0.862
2012-150057.oxy | 57 | NC 25 | 246 | 50 | 5.458 | 0.907 | 396 | 132.76 | 0.0003 | 0.862
2012-150057.oxy | 57 | NC 25 | 247 | 30 | 5.426 | 0.916 | 468 | 134.84 | 0.0003 | 0.862
2012-150057.oxy | 57 | NC 25 | 248 | 20 | 5.403 | 0.91 | 471 | 134.51 | 0.0003 | 0.862
2012-150057.oxy | 57 | NC 25 | 249 | 10 | 5.746 | 0.94 | 505 | 130.73 | 0.0003 | 0.862
2012-150057.oxy | 57 | NC 25 | 250 | 0 | 6.169 | 1.034 | 510 | 133.91 | 0.0003 | 0.862
2012-150061.oxy | 61 | NC 26 | 251 | 132 | 5.722 | 0.944 | 552 | 131.82 | 0.0003 | 0.862
2012-150061.oxy | 61 | NC 26 | 252 | 100 | 5.637 | 0.956 | 557 | 135.45 | 0.0003 | 0.862
2012-150061.oxy | 61 | NC 26 | 253 | 75 | 5.542 | 0.942 | 559 | 135.75 | 0.0003 | 0.862
2012-150061.oxy | 61 | NC 26 | 254 | 50 | 5.514 | 0.944 | 569 | 136.71 | 0.0003 | 0.862
2012-150061.oxy | 61 | NC 26 | 255 | 30 | 5.438 | 0.958 | 570 | 140.63 | 0.0003 | 0.862
2012-150061.oxy | 61 | NC 26 | 256 | 20 | 5.423 | 0.894 | 611 | 131.71 | 0.0003 | 0.862
2012-150061.oxy | 61 | NC 26 | 257 | 10 | 5.491 | 0.919 | 620 | 133.69 | 0.0003 | 0.862
2012-150061.oxy | 61 | NC 26 | 258 | 0 | 5.758 | 0.939 | 621 | 130.31 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 259 | 174 | 5.695 | 0.966 | 657 | 135.49 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 260 | 125 | 5.674 | 0.924 | 668 | 130.14 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 261 | 100 | 5.657 | 0.936 | 690 | 132.2 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 262 | 75 | 5.638 | 0.988 | 708 | 139.92 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 263 | 50 | 5.5 | 0.891 | 712 | 129.47 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 264 | 30 | 5.41 | 0.941 | 715 | 138.88 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 265 | 20 | 5.46 | 0.889 | 745 | 130.09 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 266 | 10 | 5.57 | 1.004 | 1002 | 143.85 | 0.0003 | 0.862
2012-150066.oxy | 66 | NC 27 | 267 | 0 | 5.853 | 1.055 | 1035 | 143.89 | 0.0003 | 0.862
2012-150067.oxy | 67 | NC 28 | 268 | 68 | 0 | 0 | 1059 | 141.91 | 0.0003 | 0.862
2012-150067.oxy | 67 | NC 28 | 268 | 68 | 5.445 | 0.968 | 1059 | 141.91 | 0.0003 | 0.862
2012-150067.oxy | 67 | NC 28 | 269 | 50 | 5.489 | 0.908 | 1085 | 132.16 | 0.0003 | 0.862
2012-150067.oxy | 67 | NC 28 | 270 | 30 | 5.43 | 0.867 | 1105 | 127.62 | 0.0003 | 0.862
2012-150067.oxy | 67 | NC 28 | 271 | 20 | 5.466 | 0.908 | 1106 | 132.7 | 0.0003 | 0.862
2012-150067.oxy | 67 | NC 28 | 272 | 10 | 5.664 | 0.964 | 1123 | 135.93 | 0.0003 | 0.862
2012-150067.oxy | 67 | NC 28 | 273 | 0 | 5.784 | 0.963 | 1128 | 133.02 | 0.0003 | 0.862
2012-150068.oxy | 68 | NC 29 | 274 | 96 | 5.603 | 0.948 | 1335 | 135.15 | 0.0003 | 0.862
2012-150068.oxy | 68 | NC 29 | 275 | 50 | 5.439 | 0.968 | 1337 | 142.06 | 0.0003 | 0.862
2012-150068.oxy | 68 | NC 29 | 276 | 30 | 5.433 | 0.893 | 1338 | 131.33 | 0.0003 | 0.862
2012-150068.oxy | 68 | NC 29 | 265 | 20 | 5.64 | 0.927 | 1338 | 131.33 | 0.0003 | 0.862
2012-150068.oxy | 68 | NC 29 | 278 | 10 | 5.315 | 0.946 | 1337 | 142.06 | 0.0003 | 0.862
2012-150068.oxy | 68 | NC 29 | 279 | 0 | 5.648 | 0.956 | 1373 | 135.2 | 0.0003 | 0.862

## Plots
Sample #17 was an incomplete titration but it seems that the value is not too far out
Samples 30 and 31 clearly out to lunch
2 | 0.002 | 5.4 | 0.002
5 | 0.2 | 6.8 | 0.202
8 | 0.2 | 12.7 | 0.402
11 | 0.2 | 23.4 | 0.602 |  | titration incomplete
14 | 0.1 | 39.4 | 0.702
17 | 0.1 | 56.5 | 0.802
19 | 0.02 | 62.4 | 0.822
21 | 0.02 | 65.9 | 0.842
23 | 0.02 | 70 | 0.862
25 | 0.02 | 74.3 | 0.882
27 | 0.008 | 79 | 0.89
29 | 0.008 | 80.6 | 0.898
31 | 0.004 | 51 | 0.902
3 | 0.002 | 5.5 | 0.002
6 | 0.2 | 6.5 | 0.202
9 | 0.2 | 12.6 | 0.402
12 | 0.2 | 22.6 | 0.602
15 | 0.1 | 38.1 | 0.702 |  | titration incomplete
18 | 0.1 | 53.6 | 0.802
20 | 0.02 | 60.8 | 0.822
22 | 0.02 | 64.1 | 0.842
23 | 0.02 | 68.1 | 0.862
25 | 0.02 | 70.7 | 0.882
27 | 0.008 | 76.6 | 0.89
29 | 0.008 | 78.2 | 0.898
32 | 0.008 | 81.6 | 0.906
34 | 0.004 | 66.5 | 0.91
Sample #82 was a very ugly curve with no clear end point. Seems slight out of line
Sample #83 was a very ugly curve with no clear end point. Seems slight out of line
Sample #77 Iffy duplicates
Sample #129 clearly out of line due to an incomplete titration
3 | 0.002 | 5.8 | 0.002
6 | 0.2 | 7.6 | 0.202
9 | 0.2 | 14.6 | 0.402
12 | 0.2 | 27.1 | 0.602
15 | 0.1 | 45.2 | 0.702
18 | 0.02 | 53.6 | 0.722
20 | 0.02 | 54.4 | 0.742
22 | 0.02 | 58 | 0.762
24 | 0.02 | 61.9 | 0.782
26 | 0.02 | 65.9 | 0.802
28 | 0.02 | 70.2 | 0.822
30 | 0.008 | 74.8 | 0.83
32 | 0.008 | 76.5 | 0.838
34 | 0.008 | 78.4 | 0.846
36 | 0.008 | 80.4 | 0.854
38 | 0.004 | 82.5 | 0.858
40 | 0.004 | 83.5 | 0.862
42 | 0.004 | 84.5 | 0.866
44 | 0.004 | 51.4 | 0.87
Sample 147 Dup 1 out to lunch
3 | 0.002 | 4.4 | 0.002
6 | 0.2 | 5.4 | 0.202
9 | 0.2 | 10.1 | 0.402
12 | 0.2 | 18.8 | 0.602
15 | 0.1 | 32.3 | 0.702
18 | 0.1 | 44.7 | 0.802
21 | 0.05 | 57.4 | 0.852
23 | 0.02 | 62.2 | 0.872
25 | 0.02 | 65.9 | 0.892
27 | 0.02 | 70.3 | 0.912
29 | 0.008 | 74.1 | 0.92
31 | 0.008 | 73.6 | 0.928
33 | 0.008 | 73.7 | 0.936
35 | 0.008 | 74.1 | 0.944
37 | 0.008 | 74.1 | 0.952
39 | 0.008 | 74.4 | 0.96
41 | 0.008 | 74.6 | 0.968
43 | 0.008 | 74.4 | 0.976
3 | 0.002 | 5.5 | 0.002
6 | 0.2 | 6.3 | 0.202
9 | 0.2 | 11.2 | 0.402
12 | 0.2 | 19.9 | 0.602
15 | 0.1 | 33.1 | 0.702
18 | 0.1 | 44.3 | 0.802
21 | 0.05 | 34.7 | 0.852
24 | 0.1 | 66.5 | 0.952
26 | 0.02 | 80.9 | 0.972
28 | 0.004 | 84.7 | 0.976
30 | 0.004 | 86.3 | 0.98
32 | 0.004 | 87 | 0.984
34 | 0.004 | 87.9 | 0.988
36 | 0.004 | 88.9 | 0.992
38 | 0.004 | 89.4 | 0.996
40 | 0.004 | 90.9 | 1
42 | 0.002 | 89.9 | 1.002
44 | 0.004 | 92.2 | 1.006
46 | 0.002 | 93.9 | 1.008
48 | 0.002 | 94.1 | 1.01
50 | 0.002 | 94.5 | 1.012
52 | 0.002 | 85.8 | 1.014
#202 may be an outlier.
#201 a clear problem with titration but value doesnt look way out
#198 seems low
#197 seems low
#224 slightly out of line?
#68 seems low