## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 March 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2012-16

Agency: MEAD

Location: Baynes Sound

Project: Baynes Sound

Party Chief: Sutherland T.

Platform: Vector

Date: July 7, 2012 – July 13, 2012

Processed by: Germaine Gatien

Date of Processing: 21 December 2012 – 16 January 2013

Number of original HEX files: 34

Number of CTD files: 34

Number of bottle casts:	30

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1176), a SeaPoint Fluorometer (#2356) with a 3X cable, a Wet Labs Eco-AFL/FL Fluorometer (#2215), a Biospherical QSP-400 PAR sensor (#4601), a surface PAR (#16504) and an altimeter (no serial # available).

The deck unit was a Seabird model 11, serial #0425. The logging computer was #2.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, but the altimeter serial number was missing from the log and from the configuration file and there was no mention of the surface PAR. Since the SPAR is often in the equipment list but not actually mounted, it is wise to have it listed in the log to ensure the channel gets converted. The CTD and rosette logs were in good order with useful notes about problems encountered. 

The files from the first cast were misnamed with the HEX file saved with extension XMLCON and the BL file with extension XMBL. After renaming these files the data was successfully converted.

Two fluorometers were used for this cruise but there was no extracted chlorophyll sampling. The ratio of WetLabs ECO fluorescence to SeaPoint fluorescence was roughly 1 when the SeaPoint values were between 1 and 2.5ug/L, with the SeaPoint usually a little higher than the ECO. When fluorescence was <1 or >3ug/L the ECO read significantly higher than the SeaPoint. It is difficult to determine a dark value for the ECO fluorometer because the signal is very noisy, but subtracting an estimated dark value results in a ratio that is closer to 1 at the lower end of the range. At the higher end (>2.5ug/L) it appears that the SeaPoint rises much more slowly than the ECO no matter which dark values are used. Further study is needed before we can understand the differences between these two types of fluorometers.

There were 3 yo-yo tows. In each case the first downcast was processed for inclusion in the OSD data archive. The complete files were partly processed for the use of the chief scientist.

There was a problem with all sensors mounted on the secondary pump for the upcast of event #19. The primary channels were chosen for temperature and salinity, but pad values had to be inserted for dissolved oxygen data in the bottle file between 2db and 30db.

The comparison between CTD and bottle salinity was very noisy due to shallow sampling and the fact that the samples had not been analyzed until more than 4 months after collection. The results are not trusted. Results from the previous use of the sensors during 2012-05 suggest that the salinity was low by about 0.0013, but there were salinometer stability problems for that analysis. No recalibration has been applied to the salinity.

Another problem was noted with near-surface salinity samples; they were consist major outliers in comparisons with CTD salinity during this cruise and other 2012 cruises that sampled shallow, low salinity waters. This may be due to non-linearity in the Autosal.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6 mL/L from 0 to 10db

±0.3 mL/L from 10 to 45db

±0.1 mL/L below 45db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen and salinity. No CTD equipment problems were noted but there were some problems with bottles and the CTD may have hit the bottom during cast 15.

The last 3 events were yo-yo tows with the CTD moving up and down repeatedly between 0 and 20m. The speed of the tows varies.

Nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. All had been recalibrated in March or April 2011.

The calibration constants were checked for all instruments and no errors were found, but the pressure offset has been adjusted recently based on 2012-05 results. So that value was changed from +0.24db to +0.74db and the file was saved as 2012-16-ctd.xmlcon.

File 2012-16-0001.xmlcon is not really a configuration file; it is large and must be the hex file. There is no hex file in the raw data for that cast. The file was renamed as a HEX file. There is also a file 2012-16-0001.xmbl which was renamed as *.bl. After those changes were made conversion was successful.

##### BOTTLE FILE PREPARATION

The ROS files were recreated using file 2012-16-ctd.xmlcon. They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. All data looked fine.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample #s were added to the ADDSAMP file based on the rosette sheet information. Files 3 and 4 have the same sample numbers because there was a problem on the event #3 so that samples were not taken. So no hydro file should be prepared for that event. File #4 is at the same site.

The addsamp.csv file was then converted to CST files. The CST files will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Because file #21 has sample numbers that are not increasing consecutively, the file was edited to put the last bottle at the top. This will enable merging based on sample numbers.

(Late in processing the SAMAVG file for cast #19 was edited to replace SBE Dissolved Oxygen data with pad values between 2db and 30db due to the discovery of a malfunction in the secondary pump during that part of the upcast. The merge process was then repeated for that cast.)

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2012-16-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2012-16oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2012-16oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2012-16SAL.xls. The analysis was done within 4 to 10 days of collection. The files were simplified and saved as 2012-16sal.csv. That file was then converted to individual SAL files. Sample #120 was changed to #1120. Sample #10 was edited to show it was from event #4. There was a problem with event #3 and no samples were taken; a repeat cast at that station was called #4.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2012-16nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers, the bucket samples were removed and the file was saved as 2012-16-nuts.csv. The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure is scattered with more cases of the CTD being higher than the bottles than lower. The 2 deepest bottles had been flagged by the analyst, so were removed from the comparison, although the problems noted (poor seals) would be expected to lead to evaporation and hence high values but they are, in fact, lower than the CTD (by 0.13 and 0.07), not higher. These were both bottles fired right at the bottom of the cast. This is not recommended because we usually get poor comparisons at the bottom of casts; this may be due to bottom currents or sediments in samples, or imperfect Niskin flushing. It is notable that the only other bottle fired at the bottom had CTD salinity higher than the bottle by 0.10.

When all unflagged bottles below 15db are included in the fit the primary salinity is low by 0.011. If we exclude the only unflagged bottom bottle the CTD is high by 0.005. To judge whether the local salinity gradient was a significant factor, an estimate was made (from downcast data) of the difference between the CTD salinity at the depth of the stop and the CTD salinity 1.5m above the CTD. If we then exclude cases with gradient estimate >0.5, the primary CTD salinity is lower than bottles by 0.0003. For the secondary we get very similar results. But the scatter is very high even if a lower gradient cut-off is used.

The comparisons are very similar for both T/C pairs, so CTD malfunction is unlikely to explain the noise in the comparison.  But there are 2 factors that lead to errors of opposite signs and those probably account for the scatter:

- Recent experience has shown that waiting more than 2 months before analyzing salinity bottles with the type of liner and cap used for this cruise leads to high bottle values. The error is random, some bottles affected more than others, but errors of at least +0.01 are likely. This leads to the CTD looking lower than the bottles.

- These data are all from shallow waters where the local gradients are usually high. The 1.5m distance between the Niskin bottle and the CTD is significant in the presence of a strong gradient and leads to the CTD looking higher than the bottles. Even the lowest gradient found could lead to the CTD looking high by 0.025.

Nine out of 10 samples from 10db were severe outliers with the CTD lower than the bottles by from 0.15 to 1.0. The only other 10db sample was a major outlier in the opposite direction. Most of these differences are larger than we would expect from evaporation of samples due to the long wait for analysis. Moreover, we expect that sort of error to be random with pressure, not grouped in this way. Similar problems were noted during 2012-57 with the same equipment, though for that cruise the differences were not as large but were seen as deep as 25m. Possible explanations include contamination of samples or non-linearity in the Autosal since the biggest problems are associated with low salinity water. For cruise 2012-71 in October, despite quick analysis of samples, there were similar high salinity values near the surface. There is little likelihood of contamination by particles at that time of year. These outliers ere flagged “3” with comment “Problem suspected in all samples above 15db based on   comparison with CTD salinity and results during cruise 2012-57; possibly due to Autosal non-linearity..”  (For more details see document CTD550-June-June-2012-sal-study.doc.)

The large outlier from 10m at cast #4 was investigated further. The salinity gradient is high, but not enough for the 1.5m offset between bottle and CTD sensor to explain the difference. It could be a case of a mis-sample as it looks more likely to be from 5m. There were 2 casts at the site (only 1 has bottle samples) and the CTD salinity is consistent for both. The bottle has salinity of ~23.4 and the CTD ~25.4. No other 10m salinity sample from the cruise has a value <25. Since the bottle value looks too low, it cannot be explained by evaporation.

Sample #13, cast #4 should be flagged “3” as an outlier and possible mis-sample.

Samples #6, 13, 19, 24, 45, 93, 107, 114, 119, 1120, 159 and 179 were flagged “3” as major outliers that are likely due to Autosal problem or sample contamination.

The 2 samples flagged by the analyst do look a little out of line but in the opposite direction to expectations based on the comments. However, they are both bottom bottles and those are often associated with errors likely to be of opposite sign. The flags will be left as “3”. There is too much scatter to consider these outliers.

This comparison is not suitable for judging the salinity accuracy.

For full details for the COMPARE run see file 2012-16-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. Following recent recommendations from SeaBird recalibration will be based on a fit of differences between bottles and CTD DO with the offset forced to equal 0. SeaBird warn that this will not fit well for values <2mL/L due to limitations in Winkler titrations at low values, but there were no values <3mL/L from this cruise.

The scatter is greater than usual but this is not too surprising given the shallow casts, and active mixing in the region being sampled. If data are removed that have the sensor reading higher than bottles, cases where bottles were flagged “3” or “4” and 2 cases of CTD lower than the bottles by >1.5mL/L, a reasonable fit emerges with the offset forced to be 0. A few more outliers were identified based on residuals and the fit was:

Bottle DO = 1.0516 * CTD DO

This is a lower slope than found during 2012-05 in June (1.0704) though a 2nd recalibration was applied to that cruise to bring downcast into better agreement with bottles. Combining those two corrections brought the effective slope down to ~1.052. While that is remarkably close to the fit for this cruise, it may not be significant since the 2nd correction is expected to be partly correcting for imperfect response of the sensor while moving, so should not apply to the rosette files. Given the shallow casts with high temporal variation, the results from 2012-05 are likely not appropriate. So applying the 1.0516 slope looks like the best choice.

Flags – All samples flagged by the analyst were checked. There is too much scatter to comment on the samples flagged “2”. The samples flagged 3 and 4 look a little out of line, but that may not be significant given the noise in the comparison.

There were 2 outliers that had not been flagged by the analyst: one had very noisy CTD DO data; for the other it looks like there was a problem with the pump and the CTD DO is not reliable. From 2m to 30m all CTD DO data will be replaced with pad values, since the sensor was mounted on the secondary pump.

For full details of the comparisons see files 2012-16-dox-comp1.xls.

Fluorescence

There are no extracted chlorophyll data but we can compare the two fluorometers during bottle stops using the MRG files. There is one very high value – 60ug/L for the ECO but the SeaPoint goes off scale at just below 50ug/L since a 3X cable was in use, so that point will be excluded from the comparison. All other values are <10ug/L. The ratio of the 2 fluorometer readings was calculated and plotted against ECO fluorescence.

- There are 2 outliers with very low SeaPoint values and much higher ECO values. They come from different casts and different depths. (Press=52db, SP=0.016, ECO= 0.259 and Press=25db, SP=0.004, ECO=0.724). A small error in the SeaPoint would greatly affect the ratio so full-cast plots were examined and there is considerable noise at these two bottom bottles, so these should not be included in the comparison.

- An examination of the full plots was also made to see if the “dark values” are sufficiently different to have a notable effect on the ratio. There is only one cast deep enough to be useful and for that the SeaPoint achieves steady values of about 0.049ug/L ± 0.01. The ECO is much noisier with values of ~0.16 ±0.05ug/L. In the Strait of Georgia there is a rise in fluorescence near the bottom, so this comparison is between 150 and 200db, with the bottom at about 240db. These values can be used to determine if the relationship between the two sensors is better with those values subtracted.

- When the off-scale value and the 2 outliers discussed above are excluded, a pattern emerges with good correspondence between the two fluorometers when the values are between 1 and 3ug/L, with the SeaPoint usually a little higher than the ECO but sometimes the ECO is higher, with a ratio of roughly 1 ± 10%. Below and above that range, the ECO reads higher than the SeaPoint. When the dark values are subtracted from each fluorescence value, the general pattern remains the same but the ratios are notably lower at the lower end of the range. In the 1-2.5ug/L the SeaPoint is often higher than the ECO by about 10%. Given the noise in the ECO at low values, it is difficult to determine the dark value. If we assume it to be higher, than the ratio will become close to one at low values, but then the SeaPoint is higher than the ECO by a more significant amount in the middle range. So the differences at the low end of the range may be due to a poor choice of dark value for ECO or there may just be too much noise in the ECO to get reliable values. At the higher end (>2.5ug/L) it appears that the SeaPoint rises much more slowly than the ECO no matter which dark values are used.

- For now, no recalibration of fluorescence will be applied, since different cruises are leading to contrary conclusions and it is hard to determine the dark value for the ECO because it is so noisy and there is only 1 deep cast.

Plots of Titrated DO and CTD DO against CTD salinity were examined and outliers were investigated. No further flags appear justified.

At this point the data from the MRGCLN2 files were exported to a spreadsheet for comparison with the rosette files to ensure no data were misplaced or missing. No problems were found.

##### Conversion of Full Files from Raw Data

All files were converted using 2012-16-ctd-new.con.

A few casts were examined and all expected channels are present. The descent rate is moderate to high and mostly quite steady.

The two temperature and conductivity channels are fairly close during the downcasts and, as usual, much farther apart during the upcasts. Even during stops there is a lot of noise, sometimes more in the primary channels, sometimes in the secondary. However, the casts are quite shallow and in an area where tidal currents may be significant, so these variations may be real. The deepest cast was a test cast and it does look quieter at depth. There are some very odd excursions in the secondary temperature and conductivity in some upcasts, most notably between 10 and 25db in cast #19. The dissolved oxygen is also bad for cast #19, so the issue is likely with the pump or plumbing.

Altimetry is sometimes spiky, but there appears to be a useful signal at the bottom, and fluorescence, PAR and transmissivity look normal. The ECO readings are higher than the SeaPoint below 20m and lower or close above that.

The Surface PAR looks more believable than during its last use during 2012-05.

The yoyo transects look reasonable.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on a few casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge because the temperature is so noisy on the upcast. A setting of +3.0s was found appropriate for cruises 2012-05 and 2012-57 which bracketed this cruise. When downcasts only were compared from this cruise, a setting of 3s produces a reasonable match of T and DO features.

ALIGNCTD was used to advance the DO Voltage by 3.0s relative to the pressure.

##### CELLTM

The upcast data are noisy so the usual tests for CELLTM settings are difficult to interpret. The same equipment was used during six 2011 cruises. The tests for 2011-26 in June 2011 were reasonably clear and the best choice overall was found to be (α = 0.02, β=7) for the primary and (α = 0.03, β=9) for the secondary. Those results also appeared useful for the other 5 cruises when the same equipment was used. But for 2012-05 the results suggest the reverse is best, though differences were small. Perhaps there was a change of pumps and/or the way the equipment was mounted between 2011 and 2012. For 2012-57 which followed this cruise there are only 2 casts and they are noisy, but the results looked similar to 2012-05. Tests on 2 casts showed the values used for 2012-05 improved the data, but that the secondary channels looked even better with a choice of α = 0.03, β=9.

CELLTM was run using (α = 0.03, β=9) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

- on 1 cast to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. That was placed in a test directory and will not be archived.

##### Test Plots and Channel Check

There is only 1 cast deep enough for even a rough check for agreement between the pairs of T and C sensors. That was the first cast which only sampled to 240db. For comparison one cast from 2012-05 and 2 from 2012-57 are shown; the same T and C sensors were used.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2012-05-0032 | 250 | 0.0001 | +0.00015 | +0.0018 | Steady, high |
| 2012-16-0001 | 220 | -0.0001 | +0.00025 | +0.0026 | Steady, moderate |
| 2012-57-0011 | 310 | -0.0001 | +0.00025 | +0.0026 | Steady, high |
| 2012-57-0013 | 270 | -0.0002 | +0.00025 | +0.0027 | Steady, high |

The differences resemble 2012-57 but are higher than for 2012-05 suggesting drift in at least one T/C pair. However, it is hard to judge this from shallow casts.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are a few negative values in pressure, but on examination they all occur in upcasts, mostly when pumps were off. The transmissivity goes to zero within a few records of the pressure going below 0. The conductivity is clearly out of water within a second; given the pumps were off there would be some delay. Looking at other casts the conductivity is very low when pressure is +0.4db so that is likely very near the surface. Overall it looks like the pressure is close to correct, but it is possible a further correction will be needed soon. There are some negative values in the ECO fluorescence. Those are likely to disappear in processing.

Surface check was run and shows an average surface pressure for the cruise was 2.4db which looks reasonable for the Vector.

The cross-reference check was compared with the log book and problems were found in station names for casts 1, 13 and 29-32. There was no bottle file created for cast #3 because of problems with the Niskin bottles and cast #4 was a repeat. But the downcast file is fine, so it will be processed.

The station names were also fixed in the MRG and SAMAVG files.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. There were many changes in water depth made in the log book but not in the headers. The headers were changed to match the log. Most of the changes are minor but for cast #15 there was likely a typo in the header.

Plots were made and the altimetry headers all look appropriate with the exception of the yo-yo casts for which almost all values <15m look like spikes. The altimetry header will be removed from these casts.

##### Shift

Fluorescence

Tests were run on the only cast with no stops for bottles to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence traces look like that of the temperature trace. The results are difficult to judge because the upcast temperature is noisy but an advance of +6 to +19 records looks best for the ECO and +24 for the SeaPoint. For other recent cruises the best choices were found to be from +6 to +12 for the ECO and +24 for the SeaPoint.

SHIFT was run twice on all casts to advance the SeaPoint fluorescence by +24 records and the ECO fluorescence channel by +12 records.

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on 2 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. Only the deepest cast had the sort of features that prove useful in the comparison. The results looked best overall when a shift of -0.7s was applied to the primary and a shift of -1.2 to the secondary conductivity.

SHIFT was run twice on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

Header Check was repeated on the DEL files and there were no negative fluorescence values.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late March 2011 and were used for 6 cruises in 2011 and 2012-05 just before this cruise. For 4 of these cruises there were few calibration samples and/or the bottle calibration was not trusted. No corrections were applied. For 2011-27 both sensors were low by about 0.0015 but the samples had been stored for 2.5 months which can lead to evaporation, making the CTD look low compared to bottles. For 2011-17 which was analyzed within a month the primary was found to be high by just 0.0001 and the secondary low by 0.0004. This supports the notion that evaporation affected the earlier analysis. During 2012-05 the samples were analysed very quickly but there was a possible problem with the salinometer; the primary salinity was found to be low by 0.0013 and the secondary high by 0.0004. There was also a cruise after this one, 2011-57, which has been partly processed, and it shows both sensors being low, by 0.008 and 0.006, but again the analysis was done more than 2 months after collection.

Dissolved Oxygen

The sensor has been used extensively since its last recalibration in March 2011, but the method for calibrating the DO sensors changed recently, so comparison is difficult.

For 2012-05 and 2012-57 which bracketed this cruise the slope of the first correction was about 1.07 and 1.055, which does not look like calibration drift from which we would expect increasing slope. However, there was a 2nd calibration applied to 2012-05 downcast CTD DO that produces an effective slope of 1.052. There was also a 2nd recalibration for the CTD DO during 2012-57 but that corrected for pressure dependence in the fit of downcast CTD DO versus upcast bottles. This may have been due to response time problems in some very unusual DO profiles. It makes comparison of the fits of little value.

Pressure

The sensor was recalibrated in April 2011 and was used for 6 cruises in 2011 as well as 2012-05 and 2012-57 before and after this cruise. The offset was increased from +0.24 to +0.74db for cruises starting from June 2012, based on observations during 2012-57 which sampled very close to the surface. That value looked fine for 2012-05 as well.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The temperature profiles all fell within those ranges, but there were many small excursions towards low salinity between 20 and 40db and at the bottom of a few casts. Most of these outliers were in the northern part of the Sound and a few at the southern end of the study region and during the test cast in the Strait of Georgia. These could be related to run-off from the Fraser and other rivers. This climatology is considered too severe for near-shore casts and this particular area is not well represented in the data base. Moreover, low salinity values have been noted from other 2012 cruises using a variety of sensors. The excursions are not considered evidence of instrumental problems.

Repeat Casts – Cast #4 was a repeat of cast #3 but they are too shallow to be used to establish how well the sensors were performing. The plots are similar, but there were clearly some changes. Along lines of constant density differences in temperature and salinity below 15m were less than 0.02C° and 0.005, respectively, which is reasonable for this region.

Post-Cruise Calibration - There were no post-cruise calibrations available.

##### CHANNEL CHOICE STUDY

There is confusing evidence about which temperature/salinity pair to choose for editing. In other recent uses of this equipment the secondary looked best, but for these data both pairs have significant noise. However, the secondary is very bad for at least one cast, #19. The COMPARE results make the issue even more confusing, though it is such a noisy comparison that the results are suspect. The standard deviations of the bottles included in fits are high even when many bottles are rejected due to high local gradients or being bottom bottles. The standard deviation in the secondary salinity during bottle stops is higher than in the primary though both are very high, not unusual in such shallow waters amid active mixing.

Given the severe problems in secondary pumped channels for at least one cast, and slightly noisier traces on average, the primary T/C pair was selected for archiving, and thus editing.

##### DETAILED EDITING

The primary temperature and salinity channels were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records.

All casts required light editing, mostly at the top and bottom of casts.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure does not appear to require recalibration

No salinity recalibration will be applied at this time. The results of 3 cruises in 2012 show the primary to be lower than the secondary and the difference increased between 2012-05 and 2012-57 though that could indicate an alignment problem rather than calibration drift. For 2012-05 which had the salinity analysis run very promptly after collection, the primary was found to be low by 0.0013, but there was some concern about the salinometer performance. The other comparisons are definitely not trustworthy. There may be a post-cruise calibration soon, so if there has been significant change in the primary sensor, a post-processing recalibration can be applied.

SBE Dissolved Oxygen was recalibrated by applying the following equation:

Corrected DO = 1.0516 * SBE DO

CALIBRATE was run using file 2012-16-recal1.ccf to apply the corrections to the SAM files and COMPARE was rerun to ensure the correction worked well and it did. The results show values slightly low at the lower end of the range and slightly high at the higher end. When forced through the origin the fit was excellent. (See 2012-16-DO-comp2.xls.)

CALIBRATE was then applied to the EDT and MRGCLN2 files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure the CTD looked high near the bottom, and generally got higher with decreasing pressure, except at 5db where they look low. There is a lot of scatter so that the way outliers are identified has a large influence on the fit. The following fit looks satisfactory overall:

CTD DO– Bottle DO = 0.0009 * CTD Pressure - 0.1174

and the average differences was +0.095mL/L.

That correction was applied to the thinned files and COMPARE was rerun.

The results looked good in both plots of differences against pressure and DO concentration.

(See 2012-16-dox-comp3.xls and 2012-16-dox-comp4.xls for details.)

CALIBRATE was run on the COR1 files to apply the above correction. (Output: COR2)

This correction does not apply to the bottle files.

##### Special Fluorometer Processing

A median filter, size 11, was applied to the two fluorescence channels in the COR2 files.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. T-S plots looked fine. Profiles turned up no problems.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6 mL/L from 0 to 10db

±0.3 mL/L from 10 to 45db

±0.1 mL/L below 45db

For details on the processing see processing report: 2012-16-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. There was a lot of variability which may reflect local mixing. The test cast in open waters and most of the more northerly casts had saturations between 90% and 110%, but the values were between 110% and 130% to the south. A few cases of higher saturations were checked to see if the SBE DO near-surface waters were in reasonable agreement with bottles and they were overall, with bottles sometimes higher than CTD and sometimes lower.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and no errors were found.

A header check was run on the final files and no problems were found.

Track plots look ok.

Plots of each file were examined to ensure no problems had crept in and none were found. A note was put into the header of cast #19 to explain why pad values were inserted for some of the SBE Dissolved Oxygen data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### 26 Yo-Yo casts

The complete cast files for events #32-24 have been saved. The CLN files have been through steps described in sections 1-12. The SHFC1 files have been through the steps described in section 13 and the ATC files have date and time channels added to them after the shift steps. These files will not be placed in the OSD Data Archive, but will be provided to the chief scientist and may be processed further in future.

# Particulars (Notes from log)

# 3. Repeat of BS2 at 1618 – first try 1547 but vents/spigots open.

# 7. Up to surface and down again for 5m bottle.

# 14. Bottle #6 not fired.

# 15. CTD may have touched bottom but going slow.

# 19. Bit quick at bottom.

# 21. Samples 117-121 duplicated from cast #20 so changed to 1117 – 1121.

26. Bottom drifting deeper through cast.

28. Bottle 1 did not close.

29. Station should be BS27.

30. Station should be BS28.

31. Station should be BS29

32. Toyo cast 0-20m 1m/s

33-34. Toyo casts 0-20m 0.5m/s.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 1Apr2011 | Factory |  |  |
| Conductivity | 3396 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2668 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2754 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | Jan2012 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 1Apr11 | Factory |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| SBE Fluorometer | 2356 |  |  |  |  |
| Eco-AFL Fluorometer | 2215 | Aug2012 |  |  |  |
| Pressure Sensor | 75636 | 13Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |