## Notes
Analysis: | Hugh Maclean, 25-APRIL-12 to 26-APRIL-12 using Water Properties' Guildline 8400B Salinometer Serial Number 68572. Data processing by Kenny Scozzafava
Chief Scientist: | Sophia Johannessen
2012-05-17 |  | KS
#8, #41, #63 - samples had liners that were not seald properly. Please check
#68 - is flagged for rejection by the Chauvenet criteria. There is no evidence to help determine which one might be off.
2012-06-12 |  | KS
Sample #68 - one replicate value was better than the other when compared against the ctd. The value 29.1481 was low by 0.35 while 29.1414 was low by 0.011. Only the 2nd replicate was reported and flagged 26
2015-03-30 |  | KS
Amended header in the QF worksheet to show that silicone inserts were used in place of the typical nylon inserts supplied by Guildline Instruments.

## QF 2012-19 SAL FINAL
Header: | Salinity samples were collected in 200 mL type ll glass bottles with reusable silicone inserts supplied by Xiamen Cuten Accessories Industry, Ltd. And
screw caps supplied by Ocean Scientific International Limited. They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer
standardized with IAPSO standard seawater 20 days after collection. For details including a duplicate analysis, see document QF2012-19SAL.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Salinity:Bottle | Flag:Salinity Bottle | Comments
2012-19 | 5 | SOGS | 1 | 30.8268 | 6
2012-19 | 5 | SOGS | 2 | 30.83
2012-19 | 5 | SOGS | 3 | 30.7259
2012-19 | 5 | SOGS | 4 | 30.7266
2012-19 | 5 | SOGS | 5 | 30.5742
2012-19 | 5 | SOGS | 6 | 30.5729
2012-19 | 5 | SOGS | 7 | 30.435 | 6
2012-19 | 5 | SOGS | 8 | 30.4364 | 3 | SAL: Liner creased
2012-19 | 5 | SOGS | 9 | 30.2032
2012-19 | 5 | SOGS | 10 | 30.2047
2012-19 | 5 | SOGS | 11 | 30.0612
2012-19 | 5 | SOGS | 12 | 30.0637
2012-19 | 5 | SOGS | 13 | 29.9087
2012-19 | 5 | SOGS | 14 | 29.9109
2012-19 | 5 | SOGS | 15 | 29.4658
2012-19 | 5 | SOGS | 16 | 29.4671
2012-19 | 5 | SOGS | 17 | 29.1869
2012-19 | 5 | SOGS | 18 | 29.1833
2012-19 | 5 | SOGS | 19 | 27.8793
2012-19 | 5 | SOGS | 20 | 27.9169
2012-19 | 5 | SOGS | 21 | 26.2406
2012-19 | 5 | SOGS | 22 | 26.2472
2012-19 | 5 | SOGS | 23 | 26.2372
2012-19 | 5 | SOGS | 24 | 26.2367 | 6
2012-19 | 7 | SOGN | 45 | 30.6481 | 6
2012-19 | 7 | SOGN | 46 | 30.6472
2012-19 | 7 | SOGN | 47 | 30.5407
2012-19 | 7 | SOGN | 48 | 30.5413
2012-19 | 7 | SOGN | 49 | 30.4983
2012-19 | 7 | SOGN | 50 | 30.4957
2012-19 | 7 | SOGN | 51 | 30.3565 | 6
2012-19 | 7 | SOGN | 52 | 30.3561
2012-19 | 7 | SOGN | 53 | 30.0286
2012-19 | 7 | SOGN | 54 | 30.027
2012-19 | 7 | SOGN | 55 | 29.4172
2012-19 | 7 | SOGN | 56 | 29.4237
2012-19 | 7 | SOGN | 57 | 29.3215
2012-19 | 7 | SOGN | 58 | 29.3216
2012-19 | 7 | SOGN | 59 | 29.232
2012-19 | 7 | SOGN | 60 | 29.2311
2012-19 | 7 | SOGN | 61 | 29.177
2012-19 | 7 | SOGN | 62 | 29.177
2012-19 | 7 | SOGN | 63 | 29.1735 | 3 | SAL: Liner creased
2012-19 | 7 | SOGN | 64 | 29.168
2012-19 | 7 | SOGN | 65 | 29.1575
2012-19 | 7 | SOGN | 66 | 29.1595
2012-19 | 7 | SOGN | 67 | 29.1494
2012-19 | 7 | SOGN | 68 | 29.1414 | 26 | SAL: Replicate value not used. See comments
2012-19 | 9 | Fr 1 | 25 | 22.7236
2012-19 | 9 | Fr 1 | 26 | 22.5147
2012-19 | 10 | Fr 2 | 27 | 23.8703
2012-19 | 10 | Fr 2 | 28 | 23.2662
2012-19 | 11 | Fr 3 | 29 | 22.9451
2012-19 | 11 | Fr 3 | 30 | 22.9576
2012-19 | 12 | Fr 4 | 31 | 25.5636
2012-19 | 12 | Fr 4 | 32 | 25.5381
2012-19 | 13 | Fr 5 | 33 | 24.9741
2012-19 | 13 | Fr 5 | 34 | 24.9693
2012-19 | 14 | Fr 6 | 35 | 24.9745
2012-19 | 14 | Fr 6 | 36 | 24.8866
2012-19 | 15 | Fr 7 | 37 | 23.7072
2012-19 | 15 | Fr 7 | 38 | 23.3272
2012-19 | 16 | Fr 8 | 39 | 24.1891
2012-19 | 16 | Fr 8 | 40 | 24.0527
2012-19 | 17 | Fr 9 | 41 | 23.1808 | 3 | SAL: Liner creased
2012-19 | 18 | Fr 10 | 43 | 26.0297

## Duplicates
Precision Statement for Replicate Samples Drawn From a Single Niskin Bottle:
Sp for Salinity for the range 26.2364 to 30.8271 PSU was 0.0012, k = 6 with 1 outlier(s) removed
Precision Calculation for Duplicate Samples:
Precision was determined by analyzing replicate samples drawn from one Niskin
Pooled standard deviation of pairs of samples (Sp) was calucalated by:
Sp = SQRT{sum (d*d)/2k}
where k = no. of pairs and d = difference between pairs
Determination of Outliers
Outliers are discarded on the basis of Chauevenet's criteria. The statistic is calculated by the difference between the outlier and the mean, divided by the stdev.
If this absolute value is greater than the critical value of the Chauevenet criterion for the given n, the datapoint can be discarded
Precision Analysis for Replicate Samples Drawn From a Single Niskin bottle:
Event | Station | Sample | SAL1 | SAL2 | SALAvg | Difference | Diff2 | Chauvenet | Reject? | SAL1 Comment | SAL2 Comment | Critical values for Chauvenet statistic
5 | SOGS | 1 | 30.8271 | 30.8264 | 30.8268 | -0.00071321 | 5.0866e-07 | 0.0027857 |  |  |  | n | Chauvenet
5 | SOGS | 7 | 30.4351 | 30.4349 | 30.435 | -0.00027326 | 7.4673e-08 | 0.12943 |  |  |  | 3 | 1.37
5 | SOGS | 24 | 26.237 | 26.2364 | 26.2367 | -0.00058692 | 3.4448e-07 | 0.035168 |  |  |  | 4 | 1.53
7 | SOGN | 45 | 30.6463 | 30.6498 | 30.6481 | 0.003494 | 1.2208e-05 | 1.2616 |  |  |  | 5 | 1.64
7 | SOGN | 51 | 30.3562 | 30.3568 | 30.3565 | 0.0005555 | 3.0858e-07 | 0.37851 |  |  |  | 6 | 1.73
7 | SOGN | 68 | 29.1481 | 29.1414 | 29.1447 | -0.0066998 | 4.4887e-05 | 1.802 | Y |  |  | 7 | 1.8
[12 empty] | 8 | 1.87
[5 empty] | mean diff | -0.00070394 | [5 empty] | 9 | 1.91
[5 empty] | stdev | 0.0033274 | [5 empty] | 10 | 1.97
[5 empty] | count |  | 6 |  |  |  |  | 20 | 2.24
[5 empty] | SUM |  | 5.8331e-05 |  |  |  |  | 30 | 2.39
[5 empty] | Sp |  | 0.0022048 |  |  |  |  | 50 | 2.57
[5 empty] | MIN/MAX | 26.2364 | 30.8271 |  |  |  |  | 80 | 2.74
[12 empty] | 100 | 2.81
[5 empty] | with 1 outlier removed | [6 empty] | 150 | 2.93
[5 empty] | count |  | 5 |  |  |  |  | 200 | 3.02
[5 empty] | SUM |  | 1.3445e-05
[5 empty] | Sp |  | 0.0011595

## Flags
Institute of Ocean Sciences - Quality Flags for Water Sampling | Last Updated: October 2011
Double click document to use.

## Salinity Data Processing
Sample ID | Event | Station | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | New Calculated Salinity | Flag | Comments
2012-19#33 | 5 | SOGS | 1 | 2012-04-25 13:40 | 24C | 0.89289 | 5.5875e-06 | 5.4836e-05 | 0.89295 | 30.8271 | 0.00021806 | #VALUE!
2012-19#43 | 5 | SOGS | 1 | 2012-04-25 14:02 | 24C | 0.89288 | 3.466e-06 | 5.4836e-05 | 0.89293 | 30.8264 | 0.00013432 | #VALUE!
2012-19#67 | 5 | SOGS | 2 | 2012-04-26 10:36 | 24C | 0.89297 | 2.475e-06 | 5.4777e-05 | 0.89302 | 30.83 | 9.533e-05 | #VALUE!
2012-19#64 | 5 | SOGS | 3 | 2012-04-26 10:29 | 24C | 0.89027 | 3.1414e-06 | 5.4717e-05 | 0.89033 | 30.7259 | 0.00012136 | #VALUE!
2012-19#42 | 5 | SOGS | 4 | 2012-04-25 14:00 | 24C | 0.89029 | 3.3588e-06 | 5.4955e-05 | 0.89035 | 30.7266 | 0.00012957 | #VALUE!
2012-19#74 | 5 | SOGS | 5 | 2012-04-26 10:53 | 24C | 0.88634 | 2.2398e-06 | 5.4717e-05 | 0.88639 | 30.5742 | 8.6157e-05 | #VALUE!
2012-19#48 | 5 | SOGS | 6 | 2012-04-25 14:13 | 24C | 0.8863 | 2.8794e-06 | 5.4955e-05 | 0.88636 | 30.5729 | 0.00011136 | #VALUE!
2012-19#32 | 5 | SOGS | 7 | 2012-04-25 13:38 | 24C | 0.88273 | 5.5967e-06 | 5.4836e-05 | 0.88278 | 30.4351 | 0.00021811 | #VALUE!
2012-19#72 | 5 | SOGS | 7 | 2012-04-26 10:48 | 24C | 0.88272 | 3.466e-06 | 5.4777e-05 | 0.88278 | 30.4349 | 0.00013467 | #VALUE!
2012-19#71 | 5 | SOGS | 8 | 2012-04-26 10:45 | 24C | 0.88276 | 2.8897e-06 | 5.4836e-05 | 0.88282 | 30.4364 | 0.00011128 | #VALUE! | 3 | SAL: Liner creased
2012-19#27 | 5 | SOGS | 9 | 2012-04-25 11:17 | 24C | 0.87671 | 3.2142e-06 | 5.4717e-05 | 0.87676 | 30.2032 | 0.00012475 | #VALUE!
2012-19#66 | 5 | SOGS | 10 | 2012-04-26 10:34 | 24C | 0.87674 | 3.4517e-06 | 5.4777e-05 | 0.8768 | 30.2047 | 0.00013244 | #VALUE!
2012-19#61 | 5 | SOGS | 11 | 2012-04-26 10:22 | 24C | 0.87301 | 2.89e-06 | 5.4717e-05 | 0.87307 | 30.0612 | 0.00011119 | #VALUE!
2012-19#1 | 5 | SOGS | 12 | 2012-04-24 14:01 | 24C | 0.87308 | 8.9291e-06 | 5.4955e-05 | 0.87313 | 30.0637 | 0.00034571 | #VALUE!
2012-19#4 | 5 | SOGS | 13 | 2012-04-24 14:52 | 24C | 0.86905 | 4.9693e-06 | 5.4896e-05 | 0.8691 | 29.9087 | 0.00019109 | #VALUE!
2012-19#18 | 5 | SOGS | 14 | 2012-04-25 10:57 | 24C | 0.86911 | 3.4516e-06 | 5.4836e-05 | 0.86916 | 29.9109 | 0.00013117 | #VALUE!
2012-19#68 | 5 | SOGS | 15 | 2012-04-26 10:38 | 24C | 0.85751 | 3.466e-06 | 5.4777e-05 | 0.85756 | 29.4658 | 0.00013272 | #VALUE!
2012-19#70 | 5 | SOGS | 16 | 2012-04-26 10:43 | 24C | 0.85754 | 3.4659e-06 | 5.4777e-05 | 0.8576 | 29.4671 | 0.00013177 | #VALUE!
2012-19#69 | 5 | SOGS | 17 | 2012-04-26 10:41 | 24C | 0.85023 | 3.1418e-06 | 5.4777e-05 | 0.85029 | 29.1869 | 0.00012001 | #VALUE!
2012-19#46 | 5 | SOGS | 18 | 2012-04-25 14:09 | 24C | 0.85014 | 3.3308e-06 | 5.4896e-05 | 0.85019 | 29.1833 | 0.00012796 | #VALUE!
2012-19#47 | 5 | SOGS | 19 | 2012-04-25 14:11 | 24C | 0.81597 | 2.7425e-06 | 5.4777e-05 | 0.81603 | 27.8793 | 0.00010447 | #VALUE!
2012-19#9 | 5 | SOGS | 20 | 2012-04-25 10:37 | 24C | 0.81696 | 3.2145e-06 | 5.4836e-05 | 0.81702 | 27.9169 | 0.00012245 | #VALUE!
2012-19#56 | 5 | SOGS | 21 | 2012-04-25 14:35 | 24C | 0.77272 | 3.8627e-06 | 5.4836e-05 | 0.77277 | 26.2406 | 0.0001433 | #VALUE!
2012-19#73 | 5 | SOGS | 22 | 2012-04-26 10:50 | 24C | 0.7729 | 4.9817e-06 | 5.4777e-05 | 0.77295 | 26.2472 | 0.0001884 | #VALUE!
2012-19#44 | 5 | SOGS | 23 | 2012-04-25 14:05 | 24C | 0.77263 | 3.1145e-06 | 5.4896e-05 | 0.77268 | 26.2372 | 0.00011881 | #VALUE!
2012-19#14 | 5 | SOGS | 24 | 2012-04-25 10:48 | 24C | 0.77262 | 4.2112e-06 | 5.4896e-05 | 0.77268 | 26.237 | 0.00016051 | #VALUE!
2012-19#20 | 5 | SOGS | 24 | 2012-04-25 11:01 | 24C | 0.77261 | 3.3587e-06 | 5.4836e-05 | 0.77266 | 26.2364 | 0.0001251 | #VALUE!
2012-19#35 | 7 | SOGN | 45 | 2012-04-25 13:44 | 24C | 0.88821 | 4.1137e-06 | 5.4777e-05 | 0.88826 | 30.6463 | 0.00015881 | #VALUE!
2012-19#15 | 7 | SOGN | 45 | 2012-04-25 10:50 | 24C | 0.8883 | 2.2725e-06 | 5.4836e-05 | 0.88835 | 30.6498 | 8.7704e-05 | #VALUE!
2012-19#10 | 7 | SOGN | 46 | 2012-04-25 10:39 | 24C | 0.88823 | 2.0803e-06 | 5.4836e-05 | 0.88829 | 30.6472 | 8.2231e-05 | #VALUE!
2012-19#6 | 7 | SOGN | 47 | 2012-04-24 14:57 | 24C | 0.88547 | 2.7433e-06 | 5.4836e-05 | 0.88552 | 30.5407 | 0.00010343 | #VALUE!
2012-19#29 | 7 | SOGN | 48 | 2012-04-25 13:32 | 24C | 0.88549 | 3.4657e-06 | 5.4836e-05 | 0.88554 | 30.5413 | 0.00013591 | #VALUE!
2012-19#12 | 7 | SOGN | 49 | 2012-04-25 10:43 | 24C | 0.88437 | 3.1999e-06 | 5.4955e-05 | 0.88442 | 30.4983 | 0.00012432 | #VALUE!
2012-19#34 | 7 | SOGN | 50 | 2012-04-25 13:42 | 24C | 0.8843 | 3.4657e-06 | 5.4836e-05 | 0.88436 | 30.4957 | 0.00013277 | #VALUE!
2012-19#8 | 7 | SOGN | 51 | 2012-04-24 15:02 | 24C | 0.88068 | 5.4767e-06 | 5.4836e-05 | 0.88074 | 30.3562 | 0.00020884 | #VALUE!
2012-19#49 | 7 | SOGN | 51 | 2012-04-25 14:16 | 24C | 0.8807 | 3.3593e-06 | 5.4836e-05 | 0.88075 | 30.3568 | 0.0001296 | #VALUE!
2012-19#41 | 7 | SOGN | 52 | 2012-04-25 13:58 | 24C | 0.88068 | 2.7262e-06 | 5.4896e-05 | 0.88073 | 30.3561 | 0.0001071 | #VALUE!
2012-19#21 | 7 | SOGN | 53 | 2012-04-25 11:03 | 24C | 0.87217 | 3.4389e-06 | 5.4955e-05 | 0.87222 | 30.0286 | 0.00013004 | #VALUE!
2012-19#11 | 7 | SOGN | 54 | 2012-04-25 10:41 | 24C | 0.87212 | 3.4516e-06 | 5.4955e-05 | 0.87218 | 30.027 | 0.00013157 | #VALUE!
2012-19#16 | 7 | SOGN | 55 | 2012-04-25 10:52 | 24C | 0.85624 | 5.0134e-06 | 5.4836e-05 | 0.8563 | 29.4172 | 0.00019543 | #VALUE!
2012-19#24 | 7 | SOGN | 56 | 2012-04-25 11:10 | 24C | 0.85641 | 3.9648e-06 | 5.4836e-05 | 0.85647 | 29.4237 | 0.00015336 | #VALUE!
2012-19#22 | 7 | SOGN | 57 | 2012-04-25 11:05 | 24C | 0.85374 | 2.4908e-06 | 5.4836e-05 | 0.8538 | 29.3215 | 9.4379e-05 | #VALUE!
2012-19#45 | 7 | SOGN | 58 | 2012-04-25 14:07 | 24C | 0.85375 | 3.8748e-06 | 5.4717e-05 | 0.8538 | 29.3216 | 0.00014661 | #VALUE!
2012-19#31 | 7 | SOGN | 59 | 2012-04-25 13:36 | 24C | 0.85141 | 4.9772e-06 | 5.4836e-05 | 0.85146 | 29.232 | 0.00019073 | #VALUE!
2012-19#59 | 7 | SOGN | 60 | 2012-04-25 14:42 | 24C | 0.85138 | 5.7678e-06 | 5.4777e-05 | 0.85144 | 29.2311 | 0.00021924 | #VALUE!
2012-19#50 | 7 | SOGN | 61 | 2012-04-25 14:19 | 24C | 0.84997 | 4.9489e-06 | 5.4836e-05 | 0.85003 | 29.177 | 0.00018909 | #VALUE!
2012-19#25 | 7 | SOGN | 62 | 2012-04-25 11:13 | 24C | 0.84997 | 3.9746e-06 | 5.4836e-05 | 0.85003 | 29.177 | 0.00015163 | #VALUE!
2012-19#37 | 7 | SOGN | 63 | 2012-04-25 13:49 | 24C | 0.84988 | 1.3712e-06 | 5.4836e-05 | 0.84994 | 29.1735 | 5.2235e-05 | #VALUE! | 3 | SAL: Liner creased
2012-19#63 | 7 | SOGN | 64 | 2012-04-26 10:27 | 24C | 0.84974 | 3.5301e-06 | 5.4717e-05 | 0.84979 | 29.168 | 0.00013521 | #VALUE!
2012-19#40 | 7 | SOGN | 65 | 2012-04-25 13:56 | 24C | 0.84946 | 3.611e-06 | 5.4896e-05 | 0.84952 | 29.1575 | 0.00013706 | #VALUE!
2012-19#30 | 7 | SOGN | 66 | 2012-04-25 13:34 | 24C | 0.84951 | 3.865e-06 | 5.4896e-05 | 0.84957 | 29.1595 | 0.00014644 | #VALUE!
2012-19#19 | 7 | SOGN | 67 | 2012-04-25 10:59 | 24C | 0.84925 | 4.322e-06 | 5.4836e-05 | 0.84931 | 29.1494 | 0.00016701 | #VALUE!
2012-19#28 | 7 | SOGN | 68 | 2012-04-25 13:30 | 24C | 0.84922 | 4.8241e-06 | 5.4836e-05 | 0.84927 | 29.1481 | 0.00018357 | #VALUE! | 3 | SAL:Replicate outliers. Was further out in COMPARE. See comments
2012-19#23 | 7 | SOGN | 68 | 2012-04-25 11:08 | 24C | 0.84904 | 4.8241e-06 | 5.4836e-05 | 0.8491 | 29.1414 | 0.00018524 | #VALUE! | 3 | SAL:Replicate outliers
2012-19#26 | 9 | Fr 1 | 25 | 2012-04-25 11:15 | 24C | 0.67863 | 3.2142e-06 | 5.4777e-05 | 0.67868 | 22.7236 | 0.00011785 | #VALUE!
2012-19#39 | 9 | Fr 1 | 26 | 2012-04-25 13:54 | 24C | 0.67298 | 4.5979e-06 | 5.4836e-05 | 0.67304 | 22.5147 | 0.00016979 | #VALUE!
2012-19#55 | 10 | Fr 2 | 27 | 2012-04-25 14:32 | 24C | 0.7095 | 2.84e-06 | 5.4836e-05 | 0.70956 | 23.8703 | 0.00010674 | #VALUE!
2012-19#38 | 10 | Fr 2 | 28 | 2012-04-25 13:51 | 24C | 0.69326 | 2.7425e-06 | 5.4896e-05 | 0.69332 | 23.2662 | 0.00010186 | #VALUE!
2012-19#3 | 11 | Fr 3 | 29 | 2012-04-24 14:50 | 24C | 0.68461 | 5.4851e-06 | 5.4836e-05 | 0.68466 | 22.9451 | 0.0002032 | #VALUE!
2012-19#52 | 11 | Fr 3 | 30 | 2012-04-25 14:26 | 24C | 0.68494 | 3.8651e-06 | 5.4777e-05 | 0.685 | 22.9576 | 0.00014305 | #VALUE!
2012-19#62 | 12 | Fr 4 | 31 | 2012-04-26 10:25 | 24C | 0.75474 | 3.4657e-06 | 5.4836e-05 | 0.7548 | 25.5635 | 0.00012973 | #VALUE!
2012-19#58 | 12 | Fr 4 | 32 | 2012-04-25 14:39 | 24C | 0.75407 | 2.8897e-06 | 5.4836e-05 | 0.75412 | 25.5381 | 0.00010976 | #VALUE!
2012-19#51 | 13 | Fr 5 | 33 | 2012-04-25 14:23 | 24C | 0.73904 | 3.9742e-06 | 5.4896e-05 | 0.73909 | 24.9741 | 0.00014965 | #VALUE!
2012-19#5 | 13 | Fr 5 | 34 | 2012-04-24 14:55 | 24C | 0.73891 | 3.4516e-06 | 5.4955e-05 | 0.73897 | 24.9693 | 0.00012941 | #VALUE!
2012-19#65 | 14 | Fr 6 | 35 | 2012-04-26 10:32 | 24C | 0.73905 | 2.2391e-06 | 5.4777e-05 | 0.73911 | 24.9745 | 8.4555e-05 | #VALUE!
2012-19#57 | 14 | Fr 6 | 36 | 2012-04-25 14:37 | 24C | 0.73671 | 4.3337e-06 | 5.4777e-05 | 0.73676 | 24.8866 | 0.00016286 | #VALUE!
2012-19#53 | 15 | Fr 7 | 37 | 2012-04-25 14:28 | 24C | 0.70512 | 4.7426e-06 | 5.4836e-05 | 0.70518 | 23.7072 | 0.0001778 | #VALUE!
2012-19#17 | 15 | Fr 7 | 38 | 2012-04-25 10:54 | 24C | 0.6949 | 3.5403e-06 | 5.4836e-05 | 0.69496 | 23.3272 | 0.00013084 | #VALUE!
2012-19#13 | 16 | Fr 8 | 39 | 2012-04-25 10:45 | 24C | 0.71805 | 4.4783e-06 | 5.4836e-05 | 0.71811 | 24.1891 | 0.0001673 | #VALUE!
2012-19#2 | 16 | Fr 8 | 40 | 2012-04-24 14:06 | 24C | 0.71439 | 6.4449e-06 | 5.4896e-05 | 0.71445 | 24.0527 | 0.00023998 | #VALUE!
2012-19#36 | 17 | Fr 9 | 41 | 2012-04-25 13:47 | 24C | 0.69096 | 3.6104e-06 | 5.4836e-05 | 0.69102 | 23.1808 | 0.00013467 | #VALUE! | 3 | SAL: Liner creased
2012-19#7 | 18 | Fr 10 | 43 | 2012-04-24 15:00 | 24C | 0.76713 | 6.5178e-06 | 5.4896e-05 | 0.76718 | 26.0297 | 0.00024752 | #VALUE!

## RAW Salinometer Output
Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 | K15 | 2012-04-24 13:59 | 24C | 0.99972 | 6.132e-06 | 5.5015e-05 | 0.99978
2012-19#1 | SOGS-12 | 2012-04-24 14:01 | 24C | 0.87308 | 8.9291e-06 | 5.4955e-05 | 0.87313 | 30.0637 | 0.00034571
2012-19#2 | Fr 8-40 | 2012-04-24 14:06 | 24C | 0.71439 | 6.4449e-06 | 5.4896e-05 | 0.71445 | 24.0527 | 0.00023998
2012-19#3 | Fr 3-29 | 2012-04-24 14:50 | 24C | 0.68461 | 5.4851e-06 | 5.4836e-05 | 0.68466 | 22.9451 | 0.0002032
2012-19#4 | SOGS-13 | 2012-04-24 14:52 | 24C | 0.86905 | 4.9693e-06 | 5.4896e-05 | 0.8691 | 29.9087 | 0.00019109
2012-19#5 | Fr 5-34 | 2012-04-24 14:55 | 24C | 0.73891 | 3.4516e-06 | 5.4955e-05 | 0.73897 | 24.9693 | 0.00012941
2012-19#6 | SOGN-47 | 2012-04-24 14:57 | 24C | 0.88547 | 2.7433e-06 | 5.4836e-05 | 0.88552 | 30.5407 | 0.00010343
2012-19#7 | Fr 10-43 | 2012-04-24 15:00 | 24C | 0.76713 | 6.5178e-06 | 5.4896e-05 | 0.76718 | 26.0297 | 0.00024752
2012-19#8 | SOGN-51 | 2012-04-24 15:02 | 24C | 0.88068 | 5.4767e-06 | 5.4836e-05 | 0.88074 | 30.3562 | 0.00020884
2012-19#9 | SOGS-20 | 2012-04-25 10:37 | 24C | 0.81696 | 3.2145e-06 | 5.4836e-05 | 0.81702 | 27.9169 | 0.00012245
2012-19#10 | SOGN-46 | 2012-04-25 10:39 | 24C | 0.88823 | 2.0803e-06 | 5.4836e-05 | 0.88829 | 30.6472 | 8.2231e-05
2012-19#11 | SOGN-54 | 2012-04-25 10:41 | 24C | 0.87212 | 3.4516e-06 | 5.4955e-05 | 0.87218 | 30.027 | 0.00013157
2012-19#12 | SOGN-49 | 2012-04-25 10:43 | 24C | 0.88437 | 3.1999e-06 | 5.4955e-05 | 0.88442 | 30.4983 | 0.00012432
2012-19#13 | Fr 8-39 | 2012-04-25 10:45 | 24C | 0.71805 | 4.4783e-06 | 5.4836e-05 | 0.71811 | 24.1891 | 0.0001673
2012-19#14 | SOGS-24 | 2012-04-25 10:48 | 24C | 0.77262 | 4.2112e-06 | 5.4896e-05 | 0.77268 | 26.237 | 0.00016051
2012-19#15 | SOGN-45-B | 2012-04-25 10:50 | 24C | 0.8883 | 2.2725e-06 | 5.4836e-05 | 0.88835 | 30.6498 | 8.7704e-05
2012-19#16 | SOGN-55 | 2012-04-25 10:52 | 24C | 0.85624 | 5.0134e-06 | 5.4836e-05 | 0.8563 | 29.4172 | 0.00019543
2012-19#17 | Fr 7-38 | 2012-04-25 10:54 | 24C | 0.6949 | 3.5403e-06 | 5.4836e-05 | 0.69496 | 23.3272 | 0.00013084
2012-19#18 | SOGS-14 | 2012-04-25 10:57 | 24C | 0.86911 | 3.4516e-06 | 5.4836e-05 | 0.86916 | 29.9109 | 0.00013117
2012-19#19 | SOGN-67 | 2012-04-25 10:59 | 24C | 0.84925 | 4.322e-06 | 5.4836e-05 | 0.84931 | 29.1494 | 0.00016701
2012-19#20 | SOGS-24 | 2012-04-25 11:01 | 24C | 0.77261 | 3.3587e-06 | 5.4836e-05 | 0.77266 | 26.2364 | 0.0001251
2012-19#21 | SOGN-53 | 2012-04-25 11:03 | 24C | 0.87217 | 3.4389e-06 | 5.4955e-05 | 0.87222 | 30.0286 | 0.00013004
2012-19#22 | SOGN-57 | 2012-04-25 11:05 | 24C | 0.85374 | 2.4908e-06 | 5.4836e-05 | 0.8538 | 29.3215 | 9.4379e-05
2012-19#23 | SOGN-68-B | 2012-04-25 11:08 | 24C | 0.84904 | 4.8241e-06 | 5.4836e-05 | 0.8491 | 29.1414 | 0.00018524
2012-19#24 | SOGN-56 | 2012-04-25 11:10 | 24C | 0.85641 | 3.9648e-06 | 5.4836e-05 | 0.85647 | 29.4237 | 0.00015336
2012-19#25 | SOGN-62 | 2012-04-25 11:13 | 24C | 0.84997 | 3.9746e-06 | 5.4836e-05 | 0.85003 | 29.177 | 0.00015163
2012-19#26 | Fr 1-25 | 2012-04-25 11:15 | 24C | 0.67863 | 3.2142e-06 | 5.4777e-05 | 0.67868 | 22.7236 | 0.00011785
2012-19#27 | SOGS-9 | 2012-04-25 11:17 | 24C | 0.87671 | 3.2142e-06 | 5.4717e-05 | 0.87676 | 30.2032 | 0.00012475
2012-19#28 | SOGN-68 | 2012-04-25 13:30 | 24C | 0.84922 | 4.8241e-06 | 5.4836e-05 | 0.84927 | 29.1481 | 0.00018357
2012-19#29 | SOGN-48 | 2012-04-25 13:32 | 24C | 0.88549 | 3.4657e-06 | 5.4836e-05 | 0.88554 | 30.5413 | 0.00013591
2012-19#30 | SOGN-66 | 2012-04-25 13:34 | 24C | 0.84951 | 3.865e-06 | 5.4896e-05 | 0.84957 | 29.1595 | 0.00014644
2012-19#31 | SOGN-59 | 2012-04-25 13:36 | 24C | 0.85141 | 4.9772e-06 | 5.4836e-05 | 0.85146 | 29.232 | 0.00019073
2012-19#32 | SOGS-7 | 2012-04-25 13:38 | 24C | 0.88273 | 5.5967e-06 | 5.4836e-05 | 0.88278 | 30.4351 | 0.00021811
2012-19#33 | SOGS-1 | 2012-04-25 13:40 | 24C | 0.89289 | 5.5875e-06 | 5.4836e-05 | 0.89295 | 30.8271 | 0.00021806
2012-19#34 | SOGN-50 | 2012-04-25 13:42 | 24C | 0.8843 | 3.4657e-06 | 5.4836e-05 | 0.88436 | 30.4957 | 0.00013277
2012-19#35 | SOGN-45-A | 2012-04-25 13:44 | 24C | 0.88821 | 4.1137e-06 | 5.4777e-05 | 0.88826 | 30.6463 | 0.00015881
2012-19#36 | Fr 9-41 | 2012-04-25 13:47 | 24C | 0.69096 | 3.6104e-06 | 5.4836e-05 | 0.69102 | 23.1808 | 0.00013467 | liner creased
2012-19#37 | SOGN-63 | 2012-04-25 13:49 | 24C | 0.84988 | 1.3712e-06 | 5.4836e-05 | 0.84994 | 29.1735 | 5.2235e-05 | liner creased
2012-19#38 | Fr 2-28 | 2012-04-25 13:51 | 24C | 0.69326 | 2.7425e-06 | 5.4896e-05 | 0.69332 | 23.2662 | 0.00010186
2012-19#39 | Fr 1-26 | 2012-04-25 13:54 | 24C | 0.67298 | 4.5979e-06 | 5.4836e-05 | 0.67304 | 22.5147 | 0.00016979
2012-19#40 | SOGN-65 | 2012-04-25 13:56 | 24C | 0.84946 | 3.611e-06 | 5.4896e-05 | 0.84952 | 29.1575 | 0.00013706
2012-19#41 | SOGN-52 | 2012-04-25 13:58 | 24C | 0.88068 | 2.7262e-06 | 5.4896e-05 | 0.88073 | 30.3561 | 0.0001071
2012-19#42 | SOGS-4 | 2012-04-25 14:00 | 24C | 0.89029 | 3.3588e-06 | 5.4955e-05 | 0.89035 | 30.7266 | 0.00012957
2012-19#43 | SOGS-1 | 2012-04-25 14:02 | 24C | 0.89288 | 3.466e-06 | 5.4836e-05 | 0.89293 | 30.8264 | 0.00013432
2012-19#44 | SOGS-23 | 2012-04-25 14:05 | 24C | 0.77263 | 3.1145e-06 | 5.4896e-05 | 0.77268 | 26.2372 | 0.00011881
2012-19#45 | SOGN-58 | 2012-04-25 14:07 | 24C | 0.85375 | 3.8748e-06 | 5.4717e-05 | 0.8538 | 29.3216 | 0.00014661
2012-19#46 | SOGS-18 | 2012-04-25 14:09 | 24C | 0.85014 | 3.3308e-06 | 5.4896e-05 | 0.85019 | 29.1833 | 0.00012796
2012-19#47 | SOGS-19 | 2012-04-25 14:11 | 24C | 0.81597 | 2.7425e-06 | 5.4777e-05 | 0.81603 | 27.8793 | 0.00010447
2012-19#48 | SOGS-6 | 2012-04-25 14:13 | 24C | 0.8863 | 2.8794e-06 | 5.4955e-05 | 0.88636 | 30.5729 | 0.00011136
2012-19#49 | SOGN-51 | 2012-04-25 14:16 | 24C | 0.8807 | 3.3593e-06 | 5.4836e-05 | 0.88075 | 30.3568 | 0.0001296
2012-19#50 | SOGN-61 | 2012-04-25 14:19 | 24C | 0.84997 | 4.9489e-06 | 5.4836e-05 | 0.85003 | 29.177 | 0.00018909
2012-19#51 | Fr 5-33 | 2012-04-25 14:23 | 24C | 0.73904 | 3.9742e-06 | 5.4896e-05 | 0.73909 | 24.9741 | 0.00014965
2012-19#52 | Fr 3-30 | 2012-04-25 14:26 | 24C | 0.68494 | 3.8651e-06 | 5.4777e-05 | 0.685 | 22.9576 | 0.00014305
2012-19#53 | Fr 7-37 | 2012-04-25 14:28 | 24C | 0.70512 | 4.7426e-06 | 5.4836e-05 | 0.70518 | 23.7072 | 0.0001778
2012-19#54 | P24 DWS | 2012-04-25 14:30 | 24C | 0.98932 | 3.2148e-06 | 5.4836e-05 | 0.98938 | 34.5822 | 0.00012554
2012-19#55 | Fr 2-27 | 2012-04-25 14:32 | 24C | 0.7095 | 2.84e-06 | 5.4836e-05 | 0.70956 | 23.8703 | 0.00010674
2012-19#56 | SOGS-21 | 2012-04-25 14:35 | 24C | 0.77272 | 3.8627e-06 | 5.4836e-05 | 0.77277 | 26.2406 | 0.0001433
2012-19#57 | Fr 6-36 | 2012-04-25 14:37 | 24C | 0.73671 | 4.3337e-06 | 5.4777e-05 | 0.73676 | 24.8866 | 0.00016286
2012-19#58 | Fr 4-32 | 2012-04-25 14:39 | 24C | 0.75407 | 2.8897e-06 | 5.4836e-05 | 0.75412 | 25.5381 | 0.00010976
2012-19#59 | SOGN-60 | 2012-04-25 14:42 | 24C | 0.85138 | 5.7678e-06 | 5.4777e-05 | 0.85144 | 29.2311 | 0.00021924
2012-19#60 | P24 DWS | 2012-04-26 10:19 | 24C | 0.98927 | 3.8622e-06 | 5.4777e-05 | 0.98932 | 34.58 | 0.00015172
2012-19#61 | SOGS-11 | 2012-04-26 10:22 | 24C | 0.87301 | 2.89e-06 | 5.4717e-05 | 0.87307 | 30.0612 | 0.00011119
2012-19#62 | Fr 4-31 | 2012-04-26 10:25 | 24C | 0.75474 | 3.4657e-06 | 5.4836e-05 | 0.7548 | 25.5635 | 0.00012973
2012-19#63 | SOGN-64 | 2012-04-26 10:27 | 24C | 0.84974 | 3.5301e-06 | 5.4717e-05 | 0.84979 | 29.168 | 0.00013521
2012-19#64 | SOGS-3 | 2012-04-26 10:29 | 24C | 0.89027 | 3.1414e-06 | 5.4717e-05 | 0.89033 | 30.7259 | 0.00012136
2012-19#65 | Fr 6-35 | 2012-04-26 10:32 | 24C | 0.73905 | 2.2391e-06 | 5.4777e-05 | 0.73911 | 24.9745 | 8.4555e-05
2012-19#66 | SOGS-10 | 2012-04-26 10:34 | 24C | 0.87674 | 3.4517e-06 | 5.4777e-05 | 0.8768 | 30.2047 | 0.00013244
2012-19#67 | SOGS-2 | 2012-04-26 10:36 | 24C | 0.89297 | 2.475e-06 | 5.4777e-05 | 0.89302 | 30.83 | 9.533e-05
2012-19#68 | SOGS-15 | 2012-04-26 10:38 | 24C | 0.85751 | 3.466e-06 | 5.4777e-05 | 0.85756 | 29.4658 | 0.00013272
2012-19#69 | SOGS-17 | 2012-04-26 10:41 | 24C | 0.85023 | 3.1418e-06 | 5.4777e-05 | 0.85029 | 29.1869 | 0.00012001
2012-19#70 | SOGS-16 | 2012-04-26 10:43 | 24C | 0.85754 | 3.4659e-06 | 5.4777e-05 | 0.8576 | 29.4671 | 0.00013177
2012-19#71 | SOGS-8 | 2012-04-26 10:45 | 24C | 0.88276 | 2.8897e-06 | 5.4836e-05 | 0.88282 | 30.4364 | 0.00011128 | liner creased
2012-19#72 | SOGS-7 | 2012-04-26 10:48 | 24C | 0.88272 | 3.466e-06 | 5.4777e-05 | 0.88278 | 30.4349 | 0.00013467
2012-19#73 | SOGS-22 | 2012-04-26 10:50 | 24C | 0.7729 | 4.9817e-06 | 5.4777e-05 | 0.77295 | 26.2472 | 0.0001884
2012-19#74 | SOGS-5 | 2012-04-26 10:53 | 24C | 0.88634 | 2.2398e-06 | 5.4717e-05 | 0.88639 | 30.5742 | 8.6157e-05