## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April 2025 | Updated channel names & formats in TOB files. GG |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 21 Aug 2020 | ### Added HPLC Data. S.H. |
| 03-Oct-2017 | Corrected MISSION metadata field in header. R.H. |
| 31 March 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |

## PROCESSING NOTES

Cruise: 2012-25

Agency: OSD

Location: WCVI

Project: La Perouse / WCVI

Party Chief: Yelland D.

Platform: John P. Tully

Date: May 12, 2012 – 22 May, 2012

Processed by: Germaine Gatien

Date of Processing: 20 June 2012 – 14 August 2012

Number of original HEX files: 83	Number of CTD files: 83

Number of bottle files:	52		Number of bottle casts processed: 51 (#141 not needed)

Number of original TSG files:   3	Number of processed TSG files:   3

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1119), a SeaPoint Fluorometer (#2228), a Wet Labs Eco-AFL/FL Fluorometer (#2215), a Biospherical QSP-200L4S PAR sensor (#4615) and an altimeter (#1204).

A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #0603 and a flow meter.

The data logging computer was #3.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, plus details about the rosette and TSG. The logs were mostly clear with warnings about problems. There was some confusion in the rosette file entries and in file names, but problems were resolved with the help of header times and station names.

Files from a test cast in Saanich Inlet were sent to IOS for initial processing; this turned up one error in the configuration file which was corrected before the 2nd event. This is a useful step as it can also ensure reasonable data are being logged on all channels and that correct data are being displayed at sea..

Salinity samples were analyzed within 1 month after the end of the cruise and the comparison with the CTD indicates that the primary salinity is low by 0.005 and the secondary by 0.0015.The standard deviation in the fits were ~0.002.

Two fluorometers were mounted for this cruise, a SeaPoint and a WetLabs ECO fluorometer. Because SeaPoint sensors have been used for many years, there is a need to inter-calibrate with the newer ECO instruments. The problems encountered in 2011 with the ECO sensor response time have been addressed so that only a very small alignment correction was needed. One result is a very noisy signal. On average, the ECO/CHL ratio is 1.4 and the average SeaPoint/CHL ratio is 0.9, with median values lower for both. So the SeaPoint reads a little closer to the CHL overall. The average standard deviation in the data collected in the 10s-windows around bottle firing is 0.2ug/L for the SeaPoint and 0.7ug/L for the ECO.

There is a roughly linear relationship between the 2 sensors when CHL<5ug/L but the slope flattens above that. A polynomial fit can be found but when extrapolated it looks poor, so before using such an approach we need a lot more data. We could use the average ratio and say that WetLabs ECO fluorescence ~ 1.5 times the Seapoint Fluorescence, but that is not valid at extreme values. This may also not be valid when different ranges are chosen for each of the fluorometers.

Close examination of a few casts shows very sharp fluorescence gradients, so outliers are easily explained. To capture samples from fluorescence maxima under such conditions, allowance has to be made for the distance between the bottles and the CTD and for the SeaPoint sensor an allowance for misalignment needs to be made. The ECO, as currently configured, needs only a small alignment correction, so that is not a big issue. Upcasts stops should be about 0.5m deeper than observed downcast peaks in SeaPoint fluorescence and about 1.5m deeper than downcast peaks in the ECO.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.4mL/L from 0 to 300db

±0.1mL/L from 300db to 800db

±0.05mL/L below 800db

The TSG flow was lower than usual and was zero for much of the 3rd file. It is not clear whether this is a problem in the flow or with the meter. The intake temperature has large sections of unbelievably high values at times when the lab temperatures look normal. These bad patches are seen during periods when the ship speed was low, likely due to high winds. There were problems during this cruise and Line P with clogging of filters, so it is possible that something got wrapped around the intake thermometer, or possibly ship vibrations were an issue. Whatever the cause, the data are obviously bad through much of the cruise. To serve as a proxy for the intake temperature, the lab temperature was recalibrated to remove the effects of heating in the loop; the correction was based on comparisons with CTD temperatures and the difference between intake and lab temperatures in sections with believable intake temperature values. The salinity has many one-sided spikes that usually contain only a single value and are associated with temperatures values that are not spiky; those were replaced by interpolated values.

The TSG fluorescence data are raw with volts as units. The calibration parameters provided produced bad values including many significantly negative values. Comparisons with loop samples, CTD fluorescence and rosette samples suggest possible schemes for calibrating these data, but there is too much variability in the results to justify their use.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

After a test cast in Saanich Inlet files were sent to IOS for initial processing and an error in the PAR parameters in the configuration file was found. That was corrected before event #2.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as various analysis logs.

Some errors were found in file names and event numbers. These were corrected in the HEX files. Later, more errors were found – those corrections were made at a later stage.

Extracted chlorophyll, NH4, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. All had been recalibrated shortly before this cruise.

The calibration constants were checked for all instruments. The file from cast #2 after correction of the PAR parameters was saved as 2012-25-ctd.xmlcon.  The pressure offset was changed from -1.37675 to -0.1db based on the results of recent cruises using the same CTD.

The PAR sensor was not always mounted. Based on log notes cast lists were prepared with and without PAR so that it will be easy to remove PAR as appropriate.

##### Initial Rosette File Conversion and DO Calibration Study

In order to study the SBE Dissolved Oxygen sensor calibration, rosette files were converted that included Oxygen Saturation (ml/l) and bottle position. The ROS files were converted to IOS HEADER format. Those files were put through CLEAN to add event numbers (*.BOT). The BOT files were then averaged to enable an ADDSAMP file to be prepared so that sample numbers can be added to the BOT files to produce SAM files. (Since bottles were fired out of order, the file was 1st ordered on bottle position, sample #s added and it was then reordered on bottle number.) Sample numbers were added to the ADDSAMP file based on rosette log records. There were a number of files that had the wrong names: 73, 74, 75, 79 and 80 should be 74, 75, 76, 80 and 82 respectively. There were notes in the log about 2 of these errors. The file saved as 76 has no data; it looks like it was a false start for event #77 so was removed. The notes from cast #32 were confusing – check at SAMAVG stage. There were 2 rosette sheets for cast #132. The first was confusing and had the wrong event # entered. The 2nd one is clear.

The ADDSAMP file was then used to add sample numbers to the BOT files and those files were bin-averaged on bottle numbers to produce SAMAVG files. Those files were then exported to a spreadsheet 2012-25-DO-cal.csv. The titrated DO values were added to that file and lines were removed for which there was no DO sampling.   A calculations was made of Ф using the equation:

Ф = Oxsol (T,S) * (1.0 + A*T + B*T2 + C*T3) * e (E*P/K)

where A, B, C and E are taken from the calibration sheet for the sensor and P,T and K are from the CTD channels – K is temperature in Kelvin degrees.   Then the ratio Titrated DO/ Ф was calculated and plotted against the SBE DO Voltage. This fit provides the M and B for the following equation:

Titrated DO/ Ф = M*(SBE DO Voltage) + B

From M and B the parameters Soc and Voffset that are to be entered in the DO configuration are:

Soc = M

Voffset = B/M

When all values were included the R2 value was 0.9979. Removing values flagged “3” or “4” produced better results. It is difficult and tedious to pick out outliers on the plots and find and eliminate them from the fit. A simpler approach was to use the M and B values from the factory calibration to determine the difference from the fit for each sample, as follows: 
  	Difference = M*Voltage – B – DO/Phi

When the data were sorted on that difference, fits were done excluding DO samples with 3 or 4 flags. Then M and B were updated to the value from that comparison and differences were calculated again. The process stops when obvious outliers have been removed as judged by visual inspection and the R2 value, being careful to stop before the DO range is significantly reduced. (DO values near 0 are not useful, so removing those leads to a better fit.)

For this cruise the changes were fairly small as more data were removed, reflecting the fact that there were few major outliers. The only calibration history for this sensor (since its last factory service) came from CTD data that were of lower quality than usual, so looking at previous results is not terribly useful. The results are shown in the following table:

| Summary of Soc Voffset including the original values in the factory calibration when E=0.036 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | m | b | Soc | Voffset | R2 |
| Bottles used | Original | 0.4348 | -0.2186 | 0.4278 | -0.5109 |  |
| 328 | all | 0.4486 | -0.2229 | 0.4486 | -0.4969 | 0.9979 |
| 303 | all data except 3 and 4 flags | 0.4485 | -0.2232 | 0.4485 | -0.4977 | 0.9981 |
| 281 | excl. outliers diff>0.02 | 0.4482 | -0.2233 | 0.4482 | -0.4982 | 0.9998 |
| 254 | excl. outliers diff>0.01 | 0.4482 | -0.2244 | 0.4482 | -0.5007 | 0.9999 |
| 210 | excl. outliers diff>0.005 | 0.4483 | -0.2246 | 0.4483 | -0.5010 | 1 |
| 165 | excl. outliers diff>0.003 | 0.4485 | -0.2252 | 0.4485 | -0.5021 | 1 |

There is little change in slope as more points are removed but the offset does gradually increase.

Cruise 2012-12, which followed this one, had values for Soc/Voffset of 0.4494/-0.5031.

For this cruise there is unlikely to be any hysteresis because the sampling did not go deeper than 2000m.To be sure, Phi was calculated using E=3.0 instead of E=3.6 and there was no noticeable difference, confirming that hysteresis is not an issue.

There are outliers, but none are severe.

For more details see files 2012-25-do-cal-study.xls.

The configuration files were updated with the new values for parameters Soc. Voffset and E and saved with names 2012-25-ctd-new.con.

##### Hysteresis Study

Hysteresis tests were run in the previous section and as expected there is no evidence of a problem.

##### BOTTLE FILE PREPARATION

The ROS files were recreated with the new DO configuration parameters. They were converted to IOS Header format and then put through CLEAN (output:BOT) to add event numbers.

Header Check was run on the BOT files and no problems were found.

Temperature and salinity were plotted for all BOT files. Cast #141 obviously has no stops and the log indicates that no rosette file is required, so it will not be processed further. The only data that looked odd were a few small spikes (2 or 3 points each) in secondary salinity for cast #58 and in primary salinity for casts #2, 45 and 93. CTDEDIT was used to clean those files and the edited files were copied to BOT.

The addsamp.csv file prepared in the DO calibration step was sorted on Event_Number and Sample_Number and then converted to CST files. The CST files will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2012-25-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2012-25_chl.xls. A simplified version of the spreadsheet was saved as 2012-25_chl.csv which was then converted to individual CHL files. Loop data were removed from that file but were added to file 2012-25-loops.csv.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2012-25oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2012-25oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2012-25SAL.xls. The file was simplified and saved as 2012-25sal.csv. That file was converted to individual SAL files. Loop data were removed, but were first added to file 2012-25-loops.csv. The salinity data were analyzed within 1 month of collection.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2012-25nuts.xls which included a report on precision.  The file was simplified, reordered on sample numbers and saved as 2012-25-nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2012-25_NH4.xls which included a report on precision. The file was simplified and saved as 2012-25-NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

Salinity

Compare was run with pressure as reference channel.

An error in the ADDSAMP file was found, so that was fixed, the merges repeated and COMPARE rerun. This error explains one of the outliers in the DO calibration of section 3.

There was only 1 outlier that was not explained by a large standard deviation in the CTD salinity. That was sample #130 from cast #40. The CTD salinity was lower than the sample by 0.02 and 0.03. The dissolved oxygen sample also looks a little out of line, but there is no evidence that the samples came from another bottle. An examination of the full profile of the CTD shows that there was a lot of vertical motion through the stop and it is quite possible that the Niskin bottle could contain a sample that differs from the CTD by that much. No quality flag is justified.

When data above 150db are excluded the primary salinity was found to be low by 0.005 and the secondary was low by 0.0015; standard deviations were 0.0018 and 0.0019. Both sensors show some variation with time but there is not enough deep data to make this clear. The drift in the primary is a little higher than in the secondary. The difference between the two channels is roughly 0.003 early in the cruise and 0.004 later on, which is consistent with the differences found between the 2 CTD salinity channels as reported later in section 11. The primary also has some suggestion of pressure-dependence, while the secondary sensor looks flat with pressure, but there is only 1 cast with an adequate profile to judge pressure-dependence.

For full details for the COMPARE run see file 2012-25-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

As expected, the fit of differences against DO concentration is quite flat, but there is a lot of scatter within that. Cast #1 was excluded because Saanich Inlet combines a deep anoxic layer and very high gradients near the surface, both of which challenge the CTD DO sensor; moreover, errors in titration are significant for the low DO samples. When differences >0.3mL/L were also excluded the CTD DO was low by an average of 0.008mL/L with a standard deviation of 0.06. The cases with differences >0.3mL/L were investigated in detail. When differences >0.1mL/L were also excluded the CTD DO was low by an average of 0.0034mL/L with a standard deviation of 0.033. Plots of differences versus pressure show no strong trend. A plot against time suggests the differences got slightly smaller with time, but that may be more a factor of DO range variations and steadier descent rates towards the end of the cruise so there is less corruption from shed wakes.

There were many samples flagged by the analyst and a few samples that were not flagged but looked out of line in COMPARE. Spreadsheet 2012-25OXY-quality-flags.xls was prepared with a summary of the analysts comments, comments based on COMPARE and a list of suggested changes to flags or comments.

Two samples (#115 and 116 from cast #40) were found that were only minor outliers in COMPARE, but have values that look too high for their depths, 1500 and 1240m. Nothing in the CTD data suggests that anything odd was going on here. The CTD DO has values that resemble bottle samples from other 2012-25 casts at those depths. The T-S plot looks normal and the upcast and downcast are reasonably close and smooth. The CTD did not touch bottom. Bill Crawford was consulted and he agrees that flags are warranted, but not total rejection of the values.

The analyst was unavailable for consultation. The following changes were made to quality flags:

| Evt # | Sam # | DO Flag | Analyst Comment | Processor comment | Changes Made |
| --- | --- | --- | --- | --- | --- |
| 40 | 115 |  |  | See notes above | Flag 3 |
| 40 | 116 |  |  | See notes above | Flag 3 |
| 57 | 199 | 36 | Replicate outliers. Please check | 2nd value looks better in COMPARE | Use 5.419 with 2 flag |
| 58 | 207 | 36 | Both duplicate values seem low. Possible misfire? | There are local reversals in the CTD DO in both downcast and upcast at this level. Value likely fine | Change flag to 26 |
| 62 | 214 | 3 | Small bubble in flask, not water sealed | Value looks high in COMPARE by ~0.15ml/L | Change flag to 4 |
| 62 | 219 | 3 | Small bubble in flask, not water sealed | Value looks high in COMPARE by ~0.14ml/L | Change flag to 4 |
| 62 | 220 | 3 | Small bubble in flask, not water sealed | Value looks high in COMPARE by ~0.18ml/L | Change flag to 4 |
| 64 | 229 | 3 | Value seems a little high. No evidence as to why. | Values looks ok in COMPARE. CTD profile shows DO reversals in both downcast and upcast. | Change flag to 2 |
| 69 | 244 | 2 | Leaky Niskin (Cap?) | Value high by ~0.08mL/L in COMPARE | Change flag to 3 |
| 111 | 384 |  |  | Bottle low by 0.5mL/L in COMPARE, but Niskin may have been filled with deeper water from a large shed wake | Add comment and 2 flag |

Comments about COMPARE results were added to other samples that did not involve flag changes.

For more details see 2012-25-dox-comp1.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

Fluorescence

COMPARE was run using the Wet Labs ECO CTD fluorescence, the SeaPoint fluorescence and the Extracted Chlorophyll from bottles. The COMPARE fit sheets for ECO vs CHL and SeaPoint vs CHL were combined in a single worksheet. The aim of this study is to see if we can find a reliable way to relate data from the newer WetLabs ECO fluorometers to older data collected with SeaPoint fluorometers. A variety of plots were prepared to look for patterns and these emerged:

- In a plot of one fluorometer versus the other, the ECO reads higher than the SeaPoint. The relationship looks roughly linear for FL<5ug/L with the ECO reading about 1.5 times the SeaPoint. For higher CHL the slope flattens out and by FL~15 the 2 sensors are quite close, though the ECO remains higher for all but 1 point. For the fit with only fluorescence <3.6ug/L, there is an offset at the origin showing that when the SP=0 the ECO~0.1ug/L. This is consistent with the difference in dark values which are 0.037ug/L for the SeaPoint and an average reading of ~0.15ug/L for ECO, whose dark values are not as steady as those of the SeaPoint. Subtracting the dark values will not make a significant difference to the fits.

- A plot of fluorescence from each sensor versus extracted CHL shows the SeaPoint is close to the CHL at low values, but reads too low when CHL>3. The ECO is too high for low CHL. For 1.5>CHL>10 the ECO is noisy but reasonably close to the CHL though the scatter is large. When CHL rises above 10ug/L it mostly reads low, though closer than the SeaPoint..

The problems encountered in 2011 with the ECO sensor response time have been addressed so that only a very small alignment correction was needed. One result is a very noisy signal. On average the ECO/CHL ratio is 1.4 with a range from 0.4 to 5.7 when one outlier is excluded; the median is 1.2.  Excluding the same outlier, the average SeaPoint/CHL ratio is 0.9 with a range of 0.3 to 4.8; the median is 0.8. The standard deviations in these ratios are almost as large as the median values. So one reads high and the other low by approximately the same amount; the SeaPoint may be marginally closer. One notable feature is that the average standard deviation in the data collected in the 10s-windows around bottle firing is 0.2 for the SeaPoint and 0.7 for the ECO.

There is no simple way to relate the results from the 2 sensors except to use the mean ratios and say that WetLabs ECO fluorescence = 1.5 * Seapoint Fluorescence. That looks reasonable except for the lowest and higher CHL values, and those excursions may be at least partly due to the distance between the CTD sensor and the bottles.

The outlier excluded in the average above is interesting. The 2 fluorometers both give high values for the 10m bottle, 15.4 and 15.8ug/L, but the extracted chlorophyll value is only 1.2ug/L. An examination of the full profile shows that during both upcast and downcast there was a very sharp gradient at about that level, so that the 1.5m between bottle and CTD could readily explain the difference. This does show the limitations in such a comparison so that we need many bottles to get a good idea of how these instruments compare.

This also shows the difficulty of sampling the FL maximum. If you judge the level based on the downcast peak, then you have to allow about +1.5m for the distance between the bottle and the CTD. For the SeaPoint there is a further adjustment of about -1m needed to allow for transit time due to pumping, so the offset would be about +0.5m. Judging by the downcast of event #30 the peak appears to be around 14 for the ECO and 15m for the SeaPoint, so stopping the CTD at 15.5m should arrange for the bottle to be in the peak zone if conditions remain the same. The CTD fluorescence would then be lower than the bottles. In fact, the stop was at 10.7m so the bottle was at 9.5m, well above the maximum..

Another cast was examined during which the CTD stopped a little below where 2 peaks had been seen in the downcasts and the extracted CHL is higher than the CTD in both cases, showing that the sample was likely near the peak.

At this point the data from the MRGCLN2 files were exported to a spreadsheet for comparison with the rosette files to ensure no data were misplaced or missing. Errors were found in event #s for the NH4 data. Those were corrected and the merge process was rerun. A few problems were detected:

- Cast #30 – the rosette log indicates a surface salinity but that is not among the data analyzed; a “9” flag was added with a note that it is presumed it was never drawn.

  - Cast #82 – no MRG file. The file had not been included in the cast list due to an error in file names that was corrected after the list was created. The MRGCLN2 file was created – there was no IOS sampling for that cast.

After these corrections the merge process was rerun.

At this stage file 2012-25-00032.MRGCLN2 was examined to ensure that the sample numbers had been added correctly. The rosette log notes were unclear. The file shows that Niskin #20 was fired twice at 251db and 76db; the second firing was not noted on the log sheet – it occurred at the same time as Niskin #13. The oxygen and nutrient samples look fine for Niskin #13. There was no sampling for Niskin #20. Niskin #14 was meant to be fired at 50m but was missed according to the log notes and was fired at 10.5db instead. The temperature values look appropriate for the depths and the nutrient sample values appear ok. The CTD salinity in the bottle file has some reversals, but those are also seen in the full profiles and the reversals are stable in T-S space, so there is no evidence of a problem.

##### Conversion of Full Files from Raw Data

All files were converted using 2012-25-ctd-new.con.

A few casts were examined and all expected channels are present. The descent rate is fairly noisy for many casts with some complete reversals of direction during the descent and obvious shed wake corruption.

The two temperature channels are close during the downcasts though where there is obvious shed wake corruption the two channels differ more with primary temperature generally having larger spikes. The upcast traces differ more and there are odd excursions with sometimes the primary higher and sometimes the secondary. This is likely something to do with how the CTD is mounted. The conductivity channels are similar to temperature.

Altimetry looks useful when the CTD got near the bottom; dissolved oxygen, fluorescence, PAR and transmissivity look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 3 casts to determine the offset between the DO channel and the primary temperature. It is very hard to judge because the temperature is noisy on the upcast, but overall a 4.5s advance looked best, though slightly lower settings looked best at the surface.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

The upcast data are noisy so the usual tests for CELLTM settings were a little hard to interpret. Tests were run on 5 casts using a variety of settings; all settings improved the data and the best choice overall was (α = 0.02, β=9) for both the primary and secondary salinity channels.

CELLTM was run using (α = 0.02, β=9) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are very noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2012-25-0037 | 500 1000 2000 | -0.0008 -0.0003 -0.0001 | +0.00015 +0.00025 +0.00030 | +0.0025 +0.0035 +0.0040 | High, V Noisy |
| 2012-25-0066 | 500 1000 2000 | -0.0004 -0.0004 -0.0001 | +0.00026 +0.00030 +0.00037 | +0.0035 +0.0039 +0.0046 | High, X Noisy |
| 2012-25-0068 | 500 1000 2000 | -0.0001 -0.0001 0.0000 | +0.00026 +0.00029 +0.00036 | +0.0031 +0.0036 +0.0044 | High, X Noisy |
| 2012-25-0096 | 500 1000 2000 | -0.0003 -0.0001 0.0000 | +0.00027 +0.00033 +0.00028 | +0.0035 +0.0040 +0.0034 | High, Mod |

The differences are small and there is no obvious time dependence or pressure dependence.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are some negative values in the ECO fluorescence, as usual. The largest negative number is near the surface and likely to be removed in editing. After editing, this check should be done again and any remaining negative values can be replaced with pad values. No other problems were noted.

There were no negative pressures and neither fluorescence channel went off-scale.

Surface check was run and shows an average surface pressure for the cruise was 2.6db which looks reasonable for the Tully; the offshore casts generally started a little deeper, and Rivers Inlet casts started closer to the surface. During 2012-12 which followed the average surface pressure was 3.1db which again looks reasonable for offshore Tully cruises.

The cross-reference check was compared with the log book and many problems were found in file names, event numbers and station names.

- The file names (and event numbers in the headers) were wrong for 5 events – original names 11, 12, 13, 45 and 46 were changed to 12, 13, 15, 46 and 48. These were changed in the IOS files and then CLEAN was rerun. They were also changed in the only one of these events that included rosette sampling; the analysts’ files already had the correct name for that event.

- The station names were fixed for casts #15, 37, 69 and 99.

- The time and position for cast #133 was a little different from the log entries, but close enough that it must correspond to this site; the headers were assumed to be correct.

- The dates for events #69 to #72 were given as 16 July while the files indicate 17 July; the log dates are impossible. The file was not adjusted.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and SAMAVG files were exported to spreadsheets. Some casts did not get within 15m of the bottom so there are no header entries. A selection of casts were checked and no problems were found in the altimetry headers for the CLN files. The altimetry header was removed from 2 SAMAVG files for which there was an entry but only a surface bottle. It is irrelevant information for these files and it may imply the sample came from close to the bottom. For most such files there was no altimetry entry.

The water depth header entries were checked against log entries; discrepancies were investigated and in all but two cases the log was found to be more accurate. The following changes were made in the CLN and SAMAVG files:

- The water depth entry was changed to match the log for casts #20, 33, 37, 40, 48, 54, 80, 82, 93 and 109. In some cases there was little evidence but the log tends to be more accurate where there is evidence so it was trusted.

- For cast # 31 there is a note in the log that the sounder reading was wrong and the CTD hit bottom. This is confirmed by a sudden drop in transmissivity to 0. The bottom pressure was 929.6db, so the depth is ~920m. That value was entered in the header.

- For cast #74 neither the header nor log entry is possible since they are both lower than the maximum depth for the cast and the altimetry is not useful in assessing how far the CTD was above bottom. So the header entry was removed.

- For cast #76 the header looks a little better than the log entry; the bottom appears to be deeper than the log entry. The file entry is reasonable, so was left unchanged.

##### 13. Shift

Fluorescence

Tests were run on two casts to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence traces look like those of the temperature traces. This task was complicated by extremely noisy upcast temperature and noisy ECO fluorescence. Tests indicate that applying the usual advance of +24 records is effective in bringing the SeaPoint fluorescence into line with the temperature. For the ECO fluorescence the results vary a lot with some casts where the fluorescence traces are closer than the temperature traces to a few spots where an advance of +24 records would be needed to bring the fluorescence into alignment. When last used a setting of +12 records was found appropriate, but that looks a little high, on average. A setting of +6 records looks best overall.

SHIFT was run twice on all casts to advance the ECO fluorescence channel by +6 records and the SeaPoint fluorescence by +24records.

After this step a few plots were made to see if peaks in the 2 fluorescence traces are in reasonable agreement during downcasts and they are.

Conductivity

Tests were run on the conductivity channels using 3 casts and a variety of shifts; the best results were with an advance of -0.6 records for the primary and +0.6 records for the secondary..

Two runs of SHIFT were used to apply those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The results are too complicated to interpret, but there is no evidence to justify further adjustments to the alignment.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The sensors were both recalibrated in late March 2011 and this is the first known use since then. Preliminary data from the cruise that followed indicates that the primary salinity was low by ~0.007 and the secondary salinity was low by ~0.002.

2. Dissolved Oxygen

The DO sensor was recalibrated in April.2011. It was used for 6 other cruises since then. The only reliable calibration sampling was for 2012-04 in April 2012 when setting of 0.4442/-0.4822 for Soc and Voffset were found appropriate. For 2012-12 which followed the settings were .4494/-0.5031.

3. Pressure

The sensor was recalibrated in April 2011 and some drift was noted in cruises earlier this year, so an offset was applied in conversion. There is no evidence of further correction needed.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. There were many cases of near-surface salinity values being below the minimum. These casts were not the ones closest to shore but near the 100m contour to the south in the LB line and seaward of the 1000m contour in the LC and LG lines. It is also low at the bottom at LC4. The salinity is a little high at LBP2 and LBP3 around 45-50db. Temperatures are slightly low at depth in Johnstone Strait and the Strait of Georgia. Given the very wet, cool winter/spring these values are likely real and are not considered evidence of instrumental problems.

Repeat Casts –

The only repeat casts had only shallow data for at least one of the two casts, so are not useful for checking sensor repeatability.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

From this cruise the bottle comparison shows the primary salinity to be low by an average of 0.005 and the secondary by 0.0015 with slight time dependence in both, (more in the primary than the secondary) but this is hard to judge given little deep sampling. There is some hint of pressure dependence in the primary but, again, with only one deep cast with more than 2 salinity samples, this is not clear. It is surprising that the primary salinity is reading this low immediately after factory calibration, but the fact that the salinity analysis was done promptly and the secondary salinity looks ok shows this is likely not a due to salinity analysis problems.

A preliminary look at the salinity samples from the cruise that followed, 2012-12, indicates that the primary is low by 0.007 and the secondary by 0.002. This includes more deep samples so is a more dependable comparison. Neither channel shows much time-dependence, but the primary does show significantly more pressure dependence than the secondary.

So, the secondary looks like a better choice. It was selected for editing and eventual archiving.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records.

For cast #53 there were sections of poor data near the bottom with no obvious explanation. The primary channels also had some problems.

For cast #96 the primary salinity had a section of bad data, so the primary channels were selected for editing.

All EDU files were copied to EDT.

After preliminary metre-averaging T-S plots of all casts were examined on screen and the only unstable features were at the surface of cast #2. That was very close to shore and no obvious instrumental cause was found, so the features were left unedited as they may well be real. There were other odd-looking features in casts #75, 89, 129 and 132 that were re-examined, but they are stable and again there is no obvious instrumental problem, so they were not edited any further. The latter 2 were in Rivers Inlet and look very similar to each other.

##### Initial Recalibration

The pressure looks ok.

The secondary salinity was found to be low by 0.0015 with a standard deviation of 0.002. There was a 4 to 6 week wait for analysis so an error due to evaporation is expected to be small. The preliminary results for 2012-12 which followed show the secondary salinity to be low by 0.002 and that comparison included many deep samples. So recalibration by adding 0.0015 looks reasonable.

For cast #83 the primary salinity was selected so that channel will also be recalibrated by adding 0.005.

The preliminary result for 2012-12 indicates a correction of +0.007, but that may be because there was further drift in that sensor or because there was deeper sampling we may be seeing more depth-dependent error.

There is scatter in the DO comparison with bottles, but there is no evidence of a systematic error and no recalibration is indicated.

No recalibration will be applied to either fluorescence channel at this time, though the WetLabs ECO fluorescence has high dark values ~0.15ug/L. A decision may be made later to do a correction based on studies comparing this fluorometer with extracted CHL and SeaPoint fluorescence. Until further studies have been done no recalibration will be applied to the WetLabs ECO fluorescence.

File 2012-25-recal1.ccf was prepared to add 0.005 to the primary salinity and 0.0015 to the secondary salinity.

CALIBRATE was run on the MRGCLN2, SAM and EDT files using that file.

COMPARE was rerun to check that the salinity corrections were done correctly and they were. The file is 2012-25-sal-comp2.xls. The average difference shows the CTD primary and secondary salinity are high by an average of 0.0002.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. ALIGNCTD corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. When the differences were plotted against DO concentration there was a lot of scatter, partly because some flagged samples are included. When differences >0.1mL/L are excluded the bottle DO is low by an average of 0.005mL/L. If that amount were added to all records, most values below 300db would be high. No further recalibration will be applied.

This comparison is used to make a rough judgment of the accuracy of downcast CTD dissolved oxygen, keeping in mind that the deep outliers were values that have already been flagged.

##### Special Fluorometer Processing

A header check was run to see if the WetLabs ECO fluorescence has negative values. There are some negative values in the EDT files but not in the preliminary bin-averaged files. The negative values occur in isolated spikes only and there are some positive spikes as well. There is no need to run CLEAN to set those values to 0 since they disappear with averaging.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on cast #99 and all casts with no PAR sensor except #96 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

REMOVE was used on cast #96 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

The PAR channel was removed from cast #99 because the cap was left on the sensor.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section, to fix the project name and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, fluorescence and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.4mL/L from 0 to 300db

±0.1mL/L from 300db to 800db

±0.05mL/L below 800db

For details on the processing see processing report: 2012-25-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged widely with most between 105% and 120% but high values in Saanich Inlet (~150%) and low values at LBP3 at ~90%. A few of the outliers were checked by ensuring bottles agreed with the CTD DO data and in every case checked there was good agreement. Saanich Inlet usually does have very high saturation. The lowest value was associated with a well-mixed surface layer.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on cast #99 and casts with no PAR sensor except cast #96 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

Note that cast #99 was included in the above set because the cap was left on.

REMOVE was run on cast #96 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, PAR, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. A few unnecessary records were removed from cast #1 (bottles 13-24 were fired at the surface but not samples) and cast #32 (bottle #20 was fired twice – 2nd firing was removed).

A header check was run and one error was found and corrected; HEAD EDIT was rerun and no further problems were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up errors in station name format; those were fixed.

The track plot was produced on screen and no further errors were found.

##### 25. Thermosalinograph Data

Data were provided in 3 hex files.

Loop data were combined in file 2012-25-loops.xls. These include 8 salinity, extracted chlorophyll and nutrient samples. Time and date were added to the file to enable addition of the TSG data later.

None of the loop samples coincide with rosette casts.

a.) Checking calibrations

The calibrations were checked and the only problem concerns the fluorometry. The configuration file contains a Vblank value of 0.186 found in February 2012. The value from the 2001 check was 0.068 and that value was used at sea during 2012-01 and 2012-14. A few test conversions were done and the value of 0.186 led to many large negative values (as low as -1.5mg/m3), as was also the case for the 2 earlier cruises. The shape of the data looks normal, so the negative values are not likely due to an instrument malfunction. Tests were also run using the old parameters and even assuming the value of 0.186 should be 0.086, but neither of these choices led to data that was consistently close to extracted chlorophyll values. It was decided to use Vblank=0 and scale factor=1 for the conversion as was done for the other 2012 cruises that had the same problem. Later, we can attempt to recalibrate if we have more information, though that did not prove to be the case for 2012-01 or 2012-14.

After that change the CON file was saved as 2012-25-tsg.con. The remote temperature calibration parameters could not be checked. They are not listed in the configuration file or header records. This is because it is connected serially with the new TSG system and details can only be checked through SEATERM when it is connected to a computer directly. This is awkward and we have no way of being sure it was done correctly. This should be checked at the beginning of every cruise to be sure no-one has changed it. If we did find an error at the end, there does not appear to be any easy way to correct it.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

A check was made of fluorescence and the minimum values measured were ~0.1 for the 3 files. This shows that the Vblank must be lower than 0.1.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Time-series plots were produced. There is some noise in the salinity but it looks possible to remove that with editing. The flow rate is unusually low and drops to 0 for much of the 3rd file.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2012-25-ctd-tsg-comp.xls.

All files overlapped with some CTD casts so they were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There was no TSG data to match cast #1 and the flow was off for casts 127-141. There were 76 matches but 7 of them occurred when the flow was zero, so there are 69 useful matches. TSG values were also found for times of underway loop sampling and added to file 2011-01-tsg-loop-comp.xls. The flow for those was ok, but the temperature differences for the TSG are not normal, so these may not be useful.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude are larger than usual with the maximum being 0.005 in latitude and 0.003 in longitude. When 10 outliers were excluded the largest differences are <0.0004° in latitude and <0.0007° in longitude with an average differences both <0.0001° with or without the outliers. The median differences were 0.00000° and 0.00001°. This shows both the times and positions are reliable for both systems. Since the differences are small for most casts it appears that the clock is ok, but what accounts for the outliers? A few were examined and the TSG data showed considerable drift around the time of the CTD casts. The outliers were mostly in Queen Charlotte Sound and near shore; these are likely due to small errors in time combined with significant drift due to tidal currents.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected throughout the cruise, but the data were very noisy at times with values sometimes reading higher than lab temperatures. This noise is not associated with the variable flow rates which did occur during this cruise. The worst sections appear to be when the ship was steaming long distances, in Juan de Fuca Strait and between the LG and LBP lines, though the LBP line itself had noisy temperature as well. In quieter sections the median and average temperature differences (Intake – Lab)were:

| Using quiet sections of TSG records |  |  |
| --- | --- | --- |
|  | median | Average |
| File 1 | -0.256 | -0.255 |
| File 2 | -0.250 | -0.241 |
| File 3 | -0.275 | -0.276 |

The intake temperature does not look reliable enough to archive, but the quiet sections do provide a basis for recalibration of the lab temperature to provide a proxy for the intake temperature. Comparison with the CTD will provide a check on that that. Normally we could consider a temperature-dependent correction, but these data are too noisy to make that possible.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. During stops when there was flow, the intake and lab temperatures differed by an average and median of 0.028Cº.

1. INTAKE TEMP When all data were included except cases where there was no flow, the TSG intake temperature was higher than the CTD by an average of 0.086Cº, a median of 0.02Cº and a standard deviation of 0.020Cº. When differences >0.1Cº were excluded, the average difference was 0.025Cº, the median difference was 0.017Cº and the standard deviation was 0.017Cº. This is more variability than we normally see. Examination of a plot of these differences against the standard deviation in the TSG temperature over 2 minutes shows outliers even when the variability in the TSG is fairly low, but there is a trend towards the TSG being higher by from 0.01Cº to 0.02Cº at the lowest standard deviations, so the median difference is reasonable.

2. LAB TEMP When all data are included, except cases where there was no flow, the TSG lab temperature was higher than the CTD by an average of 0.364Cº, a median of 0.306Cº and a standard deviation of 0.247Cº. When outliers are excluded that have differences >0.1Cº the TSG lab temperature was higher than the CTD by an average of 0.286Cº, a median of 0.289Cº and a standard deviation of 0.051Cº.

3. SALINITY When all data are included, except cases where there was no flow, the TSG salinity is lower than the CTD salinity by an average of 0.164, a median of 0.017 and a standard deviation of 0.863. The comparison is very noisy. When differences were excluded that were <0.1 or >0, the TSG salinity was low by an average of 0.0313 and a median of 0.0177, but the standard deviation is 0.05. There are outliers even when the standard deviation in the salinity is <0.03, but most of those values indicate that the salinity is low by from 0.02 to 0.

4. FLUORESCENCE A plot was made of raw TSG fluorescence against CTD SeaPoint fluorescence. When the data were sorted on standard deviation (over 2 minutes) in the TSG fluorescence, and points gradually removed with the highest standard deviations, the following fit was found using 62 points.

Seapoint FL = 6.9744 * Raw TSG FL – 0.6079   R2 =0.9596

The Seapoint compared reasonably well with extracted CHL samples during this cruise when fluorescence was <3ug/L, but values were somewhat low above that. The ECO fluorometer could be used for the comparison, but it did not do well against CHL at low values which would include most of the samples. Restricting the comparison to Seapoint FL<3ug/L leads to a a slightly different fit, but the same R2

Seapoint FL = 6.3217 * Raw TSG FL – 0.4708   R2 =0.9596

The quality of these fits is questionable and there are questions about how loop samples should compare with Niskin samples. However, if future cruises lead to similar results, this might lead to a method for converting TSG fluorescence.

(See 2012-25-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons A spreadsheet was prepared by combining all loop sample data with TSG salinity and TSG fluorescence (median values over 2 minutes). There were no rosette casts at the same time as loop sampling.  (See 2012-25-tsg-loop-comp.xls.)

The TSG salinity was lower than the loop salinity by an average of 0.038 and a median of 0.034; the standard deviation was 0.043. The first 3 loop samples came from a section in which the external temperature was consistently higher than the lab temperature and the next 4 are in a  section  where the intake temperature is extremely noisy and sometimes higher than the lab temperature. The lab temperature is much smoother, but the scatter in this comparison suggests that it might also be affected adversely The single case in which the TSG intake temperature looks normal, the salinity is lower than the loop sample by 0.023.

A fit of the median TSG fluorescence versus extracted CHL looks quite different from those against CTD fluorescence and against Niskin CHL samples. There are few data points and we know there was something odd about the temperature and salinity data, so it is possible something odd was going with the chlorophyll too. The R2 value is very low.

Loop CHL = 7.6605 * Raw TSG FL + 0.0558   R2 = 0.2364.

The loop should provide better results than comparison with the CTD, but there are grave doubts about these data.

- Surface rosette samples vs TSG fluorescence

A comparison was made between the TSG raw fluorescence and the Niskin bottle CHL data. The fit found was:

Extracted CHL = 10.628 * RAW TSG FL - 0.8787   R2 = 0.8403

The times are not matched but given that the ship was stopped we hope that error is small and random. When the data were restricted to casts with low fluorescence gradient near the surface to reduce errors due to an imperfect match in depths, there were not many points left, but the results were not very different, showing that the surface sampling depth was chosen well for this purpose. The higher slope is reasonable in that we know the Seapoint fluorescence tends to read low for CHL>3. The same pattern appears to be the case for the TSG fluorescence though, and we get a much lower slope if we look only at CHL<3:

Extracted CHL = 6.6942 * RAW TSG FL - 0.2641   R2 = 0.8803

This result is similar to the comparison with the Seapoint fluorescence.

(See 2012-25-ctd-tsg-comp.xls.)

- Calibration History

The TSG primary temperature and conductivity were recalibrated in March 2011 and were used during 2012-01 and 2012-14, in February 2012.

The intake temperature was found to be higher than the CTD by about 0.01Cº during 2012-01 and was equal to the CTD temperature during 2012-14.

The lab temperature was found to be higher than the intake temperature by from 0.25Cº to 0.30Cº and by 0.32Cº while stopped and by ~0.30Cº when moving. We would expect a little less heating in the loop in May due to higher intake temperatures.

The TSG salinity was very noisy during 2012-01; it appeared to be low by 0.035 but the standard deviation was 0.03 in that comparison. During the later cruise, 2012-14, the TSG salinity was highly variable but appeared to reflect real conditions. It was found to be low by a median of 0.007. While the salinity was less noisy for the second cruise, there were fewer points of comparison and no loop samples as were available for 2012-01. Both cruises were recalibrated by adding 0.02.

The TSG fluorescence calibration parameters were clearly bad (producing negative values), so only voltage was archived. When calibration information is updated this should be revisited.

Conclusions

1. The TSG clock appears to have worked well with most casts having very small differences in positions taken from the CTD casts and TSG record when times were matched. There were more cases of slightly larger differences than usual (up to 0.15’ latitude, 0.1’ longitude), but these are in areas where significant ship drift during casts might have exaggerated small clock discrepancies.

2. The flow rate was lower than usual, sometimes noisy and zero for 2/3rds of the third file.

3. The temperature in the loop increases by about 0.26Cº to 0.28Cº. This is based on quiet sections in the temperature records. This looks reasonable for this time of year.

4. The TSG intake temperature appears to be higher than the CTD temperature by ~0.02Cº when the record is fairly smooth, which is close to the result when it was last used. However, many of that data are unreliable.

5. The TSG lab temperature is higher than the CTD by about 0.29Cº, which taken with the previous comparison implies the heating in the loop is about 0.27Cº. This is consistent with the difference between the 2 channels in quiet sections.

5. The differences between the TSG Salinity and that of the CTD are very noisy, but the TSG appears to be low by a median of ~0.02 excluding some outliers; the standard deviation is high at 0.05. The comparison with loop samples is extremely noisy. All the loop samples were taken at a time when the intake temperature was very bad with values higher than the lab temperatures. As far as we can tell the lab temperature is ok, but perhaps there were other factors such as bubbles or biological contamination associated with stormy weather. The one loop sample taken when the intake temperature looks ok shows the TSG salinity to be lower than the loop by 0.023. The recent history of this sensor does not help much as various bits of evidence are inconsistent, suggesting salinity is low by anything between 0.007 and 0.035. As was concluded for the previous 2 cruises, adding 0.02 to the salinity is the best we can do until we get better calibration information.

6. The fluorescence data have been converted with Vblank = 0 and Scale Factor=1, so the units are volts and the data are nominal. The pre-cruise measurement of Vblank as 0.168 appears to be wrong. Several approaches were tried to find a combination of Vblank and Scale Factor that would lead to reasonable fluorescence values. A variety of fits were found against loop CHL, TSG fluorescence and Rosette CHL, but the results are sufficiently different that it is unwise to use any of them to convert the data to chlorophyll units. If more calibration information becomes available this could be revisited.

f.) Editing

Before editing sections of data with intake temperature higher than the lab temperature were examined in detail to determine whether any of these data are useable and how the lab temperature was affected. The sections with the worst data were during the steam from LG8 to LBP8, but any thought that this was due to high speeds is disproven by examining the ship track with hourly ticks; the ship was travelling relatively slowly which is assumed to mean their were adverse sea or wind conditions. When the ship changed direction to run the LBP line the differences are still unbelievable but less extreme. There is a remark in the log about one of the LBP stations requiring extra weight on the rosette due to high winds. There is also a note that the flow meter seemed broken.

Looking at the TSG file in detail you can pick out a sudden increase in the intake temperature about 1 hour after cast #67. The TSG lab temperature stays steady for at least 10 minutes after that sudden rise. If the flow were turned off, or if the intake temperature were correct, we would expect the lab temperature to rise. The salinity looks steady and the flow rate is steady at ~0.7. It appears that the TSG intake temperature is bad. How can it produce bad values, then settle to reasonable ones when conditions improve? Is it possible that seaweed got caught on the temperature sensor allowing it to warm up or friction from unusual vibrations warmed it? Whatever the cause it is unreliable.

The best procedure is to recalibrate the lab temperature by subtracting an amount consistent with the differences between the two sensors when the differences look normal, to produce a proxy for the intake temperature. The intake temperature will be removed throughout as it would be too confusing for users if we archive both the intake and proxy intake temperature channels. The original lab temperature will also be archived.

The last third of file #3 has zero values for flow rate. The data looks believable, so the question arises as to whether it is bad flow meter data or bad TSG data. The change was at the beginning of cast #124 at station Ri3. The data are very noisy before and after the change, but that is not unexpected in Rivers Inlet and Johnstone Strait. The traces are steadier in the Strait of Georgia. Comparisons between the TSG and CTD salinity look bad for the first 2 casts after the flow stopped, ok for the next 4 and bad for the last one. While the variations do not look consistent with no flow, there are enough doubts that they cannot be trusted.

The ATC files were copied to *.EDT.

The ATC files were opened in CTDEDIT. One-sided spikes in salinity (mostly just 1 point) that are not associated with temperature spikes were cleaned in all files.

The temperature, fluorescence and salinity data were removed from the last third of file #3 because the flow rate was zero.

The edited files were copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

ADD CHANNEL was used to add channel Temperature:Lab and to set that equal to Temperature:Primary.

File 2012-25-tsg-recal1.ccf was prepared to adjust salinity by adding 0.02 and subtracting 0.27 from Temperature:Primary so that it may serve as a proxy for Temperature:Intake. A few values were checked to ensure it was applied correctly and it was.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Secondary, Temperature:Difference, Conductivity:Primary, Uploy0, Flag, Position:New and Voltage:0 channels. This routine had to be run twice because Scan_Number was present twice.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

# 1. Par calibration parameters wrong in con file.

# 2. PAR con file fixed.

10. Bottles fired out of order.

18. CHL max @27m – Had to manually fire bottle at 5m.

30. Notes in log about not archiving cast, but bottles were tripped. Data was logged.

31. Bottom sounder wrong. Note in log “ouch” – CTD hit bottom, no upcast bottles.

40. Niskin #22 closed at 250m for bulk water – do not include in CHE file.

48. No data acquired.

54. No latitude so file restarted NMEA window.

64. Position altered before fire range active. Nut Cool are frozen.

68. Pause at 116db for crew change

73. File saved as event 73 corresponds to logs for 74. Renamed

74. File saved as event 74 corresponds to logs for 75. Renamed

75. File saved as event 75 corresponds to logs for 76. Renamed

76. File saved as event 76 contains no data.

79. File saved as event 79 corresponds to logs for 80. Renamed.

80. File saved as event 80 corresponds to logs for 82. Renamed.

93. CHL stored 12 hours before filtering.

99. PAR cap left on.

141. Rosette cast but no sampling so no CHE file to be produced.

PAR on 1-32, 35, 38, 48-64, 78, 82, 99 -141.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 5Apr2011 | Factory |  |  |
| Conductivity | 2280 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2663 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2424 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1396DR | 26Jan2012 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 29Mar2011 | Factory |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Eco-AFL Fluorometer | 2215 | 12May2012 |  |  |  |
| Pressure Sensor | 69698 | 15Apr2011 | Factory |  |  |
| Altimeter | 43281 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2012-25

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 26Mar11 | Factory |  |  |
| Conductivity | 2487 | 26Mar11 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 1Feb12 | IOS |  |  |
| Temperature:Secondary | 0603 | 03Mar11 | “ |  |  |