## REVISION NOTICE TABLE

| DATE | DESCRIPTION OF REVISION |
| --- | --- |

# PROCESSING NOTES

Cruise: 2012-49

Agency: PBS, Salmon and Freshwater Ecosystems, Nanaimo, BC

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia Juvenile Salmon Survey

Party Chief: Neville C.

Platform: W.E. Ricker

Date: 21 June 2012 – 19 July 2012

Processed by: Germaine Gatien

Date of Processing: October 24, 2012 – November 1, 2012

Number of original hex files: 	57

Number of CTD casts processed: 55 (1 shallow test cast and 1 aborted cast not processed)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with an SBE43 Dissolved Oxygen sensor (#997), an SBE pH sensor #0692 and a WetLabs ECO-FL Fluorometer #2216.

# SUMMARY OF QUALITY AND CONCERNS

The log book did not contain an equipment list or personnel list, but the latter was obtained from the cruise report. The log contains a record of the number of bottles taken but not what was sampled. There were notes on problems encountered.

For some casts the CTD was returned to the surface, the Niskin sampled or replaced, and a second lowering done to obtain a near-surface sample. The CTD acquisition stopped at the surface (possibly the CTD was removed), so there are no data available for comparison with those samples and no check on the depth of sampling. Leaving the CTD on would provide potentially useful data.

There were salinity calibration samples taken, but at the time of processing the results were not available. The history of the T/C sensors suggests that the salinity is within ±0.001. Should the salinity analysis results become available these data should be re-examined to see if recalibration is appropriate.

There was no dissolved oxygen calibration sampling for this cruise and the only good recent comparison with bottles was during a cruise in February 2012. Those results were used to convert these data, but a small offset was later applied based on sampling in anoxic waters in Saanich Inlet.

No extracted chlorophyll sampling was available at time of processing for comparison with the ECO fluorometer. (It is unknown if such sampling was done.) In other uses the ECO was found to read higher than extracted chlorophyll when CHL<5ug/L. Inter-comparisons from other cruises suggest ECO fluorescence reads higher than SeaPoint fluorescence by roughly 50% when CHL<5ug/L. At higher CHL values the two types of sensors are generally closer to each other and to extracted CHL.

No field calibration information was available for the pH sensor and there are concerns about how the sensor performs, so the data will not be placed in the IOS Data Library at this time.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained. There were a few problems noted. Those that will affect processing include pumps being off for the downcast section of cast #39, no pH data recorded for cast #54-57 and cast #23 was aborted because the cap had been left on the pH probe.

Bottle data were not available when processing began. There was sampling at the top and bottom of some casts, but there is no note of what was sampled, though it is likely that salinity was included. To date no salinity samples or analysis results have been found. There is also some confusion about the distance between the Niskin bottle and the CTD; the log indicates it was 3m while 5m is the usual offset. There is always some error in determining the depth at which bottles close, but not knowing the distance from CTD to Niskin adds to that. Should bottle results become available this will limit the confidence we would have in the comparison with CTD data.

The first bottle was fired when the CTD was at the bottom. The CTD was then raised to the surface, the bottle replaced and lowered for the near-surface sample. Acquisition stopped before the lowering for the 2nd bottle, so it is impossible to check the depth of the stops. Moreover, any comparison of near-surface salinity would have to be against the initial upcast which is unlikely to be reliable enough for calibration purposes.

The cruise summary sheet was completed.

The histories of the dissolved oxygen, fluorometer, conductivity and pressure sensors were obtained.

The configuration file did not change through the cruise. All entries matched the latest calibrations, but the Soc and Voffset values were updated to match those found during cruise 2012-12 when there was good DO sampling. The file was then saved as 2012-49-ctd-new.xmlcon.

##### Conversion of Raw Data

All data were converted using file 2012-49-ctd-new.xmlcon.

Plots were made for a few casts.

- The temperature and conductivity channels track well on the downcast but the upcasts show larger differences. These differences look as though they are due to alignment or pump problems because when the gradient reverses sign, so do the differences. There is also a significant vertical offset that is seen only in the upcasts. Such observations are common and it is tempting to blame the wake created by the rosette package. But for this cruise there was no rosette, and no apparent difference between casts with Niskin bottles mounted and those without. All sensors have some noisy features; it is not clear which sensor pair is better.

- The descent rate is moderate to high and fairly steady.

- Fluorescence looks noisy so it is hard to judge vertical offset between downcast and upcast. The dark value is about 0.15ug/L.

- Dissolved Oxygen data look ok with the usual vertical offset between downcast and upcast.

- The pH signal looks normal with some hysteresis.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### ALIGNCTD

Various advancements were tested to make DO and temperature traces have the same offset between downcast and upcast data. Since the upcast temperature is very noisy, more weight than usual was put on matching distinctive features in the downcast, but what worked for the latter usually provided the best offset between downcast and upcast traces as well. The best results overall were with an advance of about +4s but +3.5s also looked reasonable. During 2012-14 an advance of +4s was applied while for 2012-01 +3.5s was used.

ALIGNCTD was used to advance the DO signal relative to temperature by 4s.

##### CELLTM

Tests were run on 4 casts using a variety of settings and the best choice for 3 of them was (α=0.03, 1/β=9) for both T/C pairs, while for the other cast the best choice varied from one feature to another, likely due to noisy upcast data. For 2012-01 and 2012-14 the best choice was (α=0.0245, 1/β=9.5) for the primary and (α=0.03, 1/β=9) for the secondary.

CELLTM was run using (0.03, 9) for the primary and the secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration; the tau correction was applied.

on a few casts to look at differences between channel pairs.

##### Test Plots and Channel Check

There were only 3 casts deeper than 500m. The same sensors were used for 2011-11 in November and 2012-01 and 2012-14 in February so data from 1 cast for each of those cruises are included in the comparison to see if there are significant differences. The pressures are different for the earlier results, so some of the variation is likely due to higher gradients for the shallower casts.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-11-0053 | 380 | +0.0009 | +0.0001 | +0.0003 |  |
| 2012-01-0050 | 800 | +0.0005 | -0.0001 | -0.0020 | Mod, noisy |
| 2012-14-0019 | 400 | -0.0001 | +0.00005 | +0.0006 | High, noisy |
| 2012-49-0021 | 575 | +0.0002 | ~0 Noisy | -0.0002 | High, F steady |
| 2012-49-0024 | 575 | +0.0002 | ~0 Noisy | -0.0002 | High, Noisy |
| 2012-49-0057 | 575 | +0.0004 | -0.0001 | -0.0004 | High, F steady |

The differences are all small. The salinity differences were noted to increase with depth during 2012-12, and the comparison with bottle samples showed some pressure-dependence in the secondary sensors. There was no salinity calibration sampling during 2012-14. The results of this cruise do not suggest any significant drift in salinity, but the evidence is weak. None of the other summer Ricker cruises that used this equipment sampled deeper than 500m (only 2012-51 got that deep), so this is the best information available.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. There are some negative values that appear to be at the beginning of files. These will likely disappear in the DELETE or CTDEDIT steps.

A cross-reference list was produced and checked against log records. There were few station names and no bottom depths entered. The station names that were entered look wrong as they all say PATRICIA BAY TEST but are in a variety of locations and only the first is close to PATRICIA BAY. That first cast was not entered in the log book so likely really was a test and should not be processed further. The station names were removed from files 2-6. The times are not listed in the log for the first few casts. The log book positions are generally close to those in the file headers but there are sometimes small differences that may be accounted for by drift during recording in the log and starting the cast.

The track plots were produced and look reasonable.

The surface report indicates that the average surface pressure was 2.25db but the associated salinity values are mostly very low, so these are likely from before the pumps were turned on.

##### SHIFT

Fluorescence

In previous uses of ECO Fluorometers it has been found necessary to advance the fluorescence relative to pressure by as much as 48 records, but in recent uses that adjustment has become smaller due to changes in the set-up of the sensor and an advance of 6 to 12 records has been used on recent cruises. A few casts were examined and a setting of +12 records looks appropriate to make the offset between upcast and downcast fluorescence traces look similar to that of temperature.

SHIFT was run on the fluorescence channel to advance by 12 records.

Conductivity

Tests were run on 3 casts with noisy salinity. For both primary and secondary conductivity the best results in reducing the noise were those runs with the shift parameter set to -0.5 records.

All casts were put through 2 runs of SHIFT with parameter -0.5 records for both the primary and secondary conductivity.

Dissolved Oxygen

The SBE DO data have already been aligned, but a few casts were examined to see if further adjustments are required. There are areas where the vertical offsets between upcast and downcast are slightly larger than in the temperature traces, but in other cases they are smaller. However, there were 3 casts in Puget Sound (#51-53) where near-surface DO traces were very different. These were near-shore casts and there might be real variations in DO that are not seen in temperature or there might have been some contamination of the sensor by biological material. Casts after those 3 appear to be fine.

SHIFT will not be run on this channel.

pH

The pH data are the hardest for which to determine the best alignment. There were few step features to use to check for vertical offset and the noisy upcast temperature makes it difficult to compare the vertical offset of the two traces. In previous uses a wide variety of alignment values have been found appropriate; the differences may be due to different arrangements of sensors on the CTD and in some cases the pH sensor is pumped and sometimes not.. For several casts the unshifted data looked best, for a few others a shift of ~2s looked appropriate. Given the uncertainty and the fact that the data will not be placed in the IOS data archive at this time, no shift was applied. The SHFC1 data will be backed up so that these can be processed further in future if and when more information becomes available about the sensor.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for 1 cast and was in the surface section before the pumps were turned on, so will not affect data to be archived.

##### COMPARISON WITH BOTTLES

There was no calibration sampling available at the time of processing.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

The pumps were off for the downcast of cast #39, so the file was put through REVERSE and then DELETE to produce an upcast file for editing.

The decision about which sensor pair to use was based on the results of 2012-01since there was no calibration sampling for this cruise or other cruises using the same equipment during July and August 2012. During 2012-01 the primary salinity was closer to bottles and there was evidence of pressure dependence in the secondary sensors. So the primary will be used for this cruise as well.

CTDEDIT was used to clean the primary temperature and salinity data.

For most casts some near-surface records were removed due to pumps being off or descent rate being low and noisy. Some records corrupted by shed wakes were removed, mostly near the bottom of casts. Most of these casts come from areas where small unstable features are expected, so editing was light.

Editing of salinity was applied where it was clear that spikes were due to a mismatch between temperature and conductivity.

All casts required some editing.

On-screen plots of descent rate and pump status were used to guide editing.

Note was made of the editing details in the headers of the relevant files.

The edited files were copied to *.EDT files.

##### Other Comparisons

Previous experience with these sensors –

- Both temperature and conductivity sensors have been used on many other cruises since their last factory service, but for most the CTD data were accidentally averaged over 24 scans limiting the value of the results. They were used for cruise 2012-01 in February with good sampling and quick salinity analysis. The primary salinity was found to be low by 0.0004 while the secondary was low by 0.0028. There was some pressure-dependence in the secondary salinity.

- The pressure sensor has been used 8 times since it was last recalibrated and no evidence suggests that the offset needs adjustment.

- The dissolved oxygen sensor was used for 2012-01 and there was good sampling so that calibration tests could be done to produce updated Soc/Offset values.

Historic ranges – Data were plotted with the 3-standard deviation plots of temperature and salinity ranges where local climatology was available. The near-surface salinity was lower than the range minimum for casts on the northern side of the central part of the Strait of Georgia and the temperature was low between 80 and 140db at the northern end of the Strait. The temperature excursions look like those seen during 2012-05 in the same area but a few weeks earlier. Since 2012-05 used different sensors, these excursions are not indicative of instrumental problems, but rather, reflect real conditions. Low near-surface salinity values were also seen during 2012-05 in the same general area but they were not as deep and not as extensive as seen in 2012-49. That is likely due to the fact that 2012-49 sampled closer to shore than 2012-05 and possibly due to changing conditions due to freshwater inflow. Again, these excursions are likely real, not instrumental.

##### Initial Recalibration

The pressure values look ok.

Salinity samples were not available. Given that the samples came from a Niskin on the wire and there are some uncertainties about the distance between the CTD and bottles, it did not seem worth delaying the preparation of files for the IOS Data Archive. If the data have not yet been analyzed the results are unlikely to be trustworthy as samples degrade after 2 months with scatter increasing significantly. If the data do become available, they will be examined. If there is sufficient evidence to justify recalibration, the archived data will be corrected.

There was no dissolved oxygen calibration sampling, but there was a cast in Saanich Inlet where anoxic conditions are expected. The dissolved oxygen concentration sampling has negative values at the bottom of cast #42 in Saanich Inlet. Negative values could be due to spikes or a small error in the calibration. The full profile was examined to see which is most likely. The associated temperature gradient is not high, and there is no evidence of an “overshoot” in the oxygen concentration. The values go down gradually during the downcast and the beginning of the upcast until the DO reaches equilibrium at a value of ~0.048mL/L. Sampling in anoxic waters does not usually work well for calibration of the instrument in general because of limitations in the Winkler analysis. But in this case it does provide useful information about the offset. The Voffset applied in the configuration file is to the voltage, while this correction is applied to the concentration. The error could arise from either a faulty slope or offset in the original calibration. The correction is small given the uncertainties in the calibration for this cruise, but will ensure that there are no negative values in the data. The CTD did not fully reach equilibrium during the downcast so there will be small positive values in the CTD file for cast #42.

CALIBRATE was run to apply a correction of +0.048mL/L using calibration control file 2012-49-recal.ccf.

##### Fluorescence Processing

The COR1 files were put through a median filter, size 11, applied to the fluorescence channel only.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the filtered files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and look fine.

Profile and T-S plots were made to check all channels and no problems were found.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE,  SBE:pH, Status:Pump, Descent_Rate and Flag. (*.REM)

A second set of files were prepared with the pH data included except for casts 54-57 during which the pH sensor was not connected. (*.REMph)

These two sets of files were both put through the final steps.

A second CTD dissolved oxygen channel was derived with units umol/kg.

Dissolved Oxygen saturation was calculated and surface values were plotted. Most values were between 90% and 110%, with a few higher values in Puget Sound, Saanich Inlet and one northern inlet. Lower values were found in Satellite Channel, Haro Strait, the southern side of Juan de Fuca Strait and a few places on the south-western side of the Strait of Georgia. The DO profiles associated with the low saturations show low DO gradients, indicative of active or recent mixing which would explain low surface saturation. The high saturation values are found in regions with high DO surface gradients. There is no suggestion of poor DO sensor calibration.

In order to add the station names and positions to the files spreadsheet 2012-49-hdr_merge.csv was prepared with that information included, where available. Program MERGE CSV file to Headers was run and a few casts checked to ensure the information was entered correctly and it was. There were no station names in the log for some of the casts.

REORDER was used to put the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence data are nominal and unedited, except that some

records were removed in editing temperature and salinity.

There was no calibration sampling for dissolved oxygen, but the

results of cruise 2012-01 were used to derive DO concentration.

A small correction was later added based on negative DO values

found in the anoxic layer of Saanich Inlet.

There was no calibration sampling for salinity, but the results

of cruise 2012-01 in February 2012 suggest calibration drift

is <0.001.

The SBE pH sensor data has been removed from this file since there

was no field calibration information  available and there are concerns

about how the sensor performs.

For details on the processing see processing report: 2012-49-proc.doc.

For the CTDpH files the section concerning the SBE pH sensor was replaced with::

THIS FILE WAS PREPARED FOR THE USE OF THE CHIEF SCIENTIST. IT CONTAINS

SBE ph DATA WHICH WAS REMOVED FROM THE FILE TO BE PLACED IN THE IOS

DATA LIBRARY, BECAUSE THERE WAS NO FIELD CALIBRATION INFORMATION AND

THERE ARE SOME CONCERNS ABOUT HOW THE SENSOR PERFORMS.

For the last 4 casts a different note was added because the pH sensor was not connected.

##### Producing final files

The Standards Check routine was run and no problems were found. The final files were named CTD and CTDph.

A cross-reference listing was produced for CTD files.

HEADER CHECK was run and no problems were found.

A cruise track was plotted and no errors found.

The sensor history was updated for the CTD sensors.

Particulars: (notes from log book)

3. Pump on but read “110”

15. Surface Niskin deployed deeper than intended at depth of 7m.

23. pH capsule left on. CTD cancelled returned to surface at 180m. File not processed.

24. Rerun of 23.

39. O2 feature may be incorrect. Ship had to reposition, winch stopped, pump off on trip down, on for up.

54. pH bad or not recorded.

55-57. pH not recorded

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | No | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2710 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2128 | 29Mar2011 | Factory |  |  |
| SBE 43 DO sensor | 0997 | 23Apr11 | Factory |  |  |
| SBE18 pH sensor | 692 | Dec 2010 | Factory |  |  |
| Eco-AFL Fluorometer | 2216 | June2012 | IOS? |  |  |
| Pressure Sensor | 63507 | 12Apr2011 | Factory |  |  |