## REVISION NOTICE TABLE

| DATE | DESCRIPTION OF REVISION |
| --- | --- |

# PROCESSING NOTES

Cruise: 2012-51

Agency: PBS, Marine Ecosystems and Aquaculture Division, Nanaimo, BC

Location: West Coast Vancouver Island

Project: Sardine Multi-species WCVI Night Trawl Survey

Party Chief: McFarlane S.

Platform: W.E. Ricker

Date: 17 July 2012 – 02 August 2012

Processed by: Germaine Gatien

Date of Processing: October 31, 2012 – 2 November 2012

Number of original hex files: 40

Number of CTD casts processed: 40

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with an SBE43 Dissolved Oxygen sensor (#997), an SBE pH sensor #0852 and a WetLabs ECO-FL Fluorometer #2216.

# SUMMARY OF QUALITY AND CONCERNS

Log sheets were provided and a personnel list was available from the cruise report. The log contains a record of the number of bottles taken. There were notes on problems encountered.

The event numbers should be used to name the CTD files. For this cruise the numbers correspond to consecutive CTD casts which were recorded in the Station Name column. The file names were changed to match the log entries of event numbers.

Acquisition started late for 3 casts (events #3, 5 and #39) so that the downcast profiles were incomplete. So upcast data were processed and will be used for the IOS Data Archive. Upcast data quality is lower than usual with noisier data, and temperature is systematically lower and salinity higher than would be the case for downcast data at the same pressure, because the CTD travels through the wake created by the wire and CTD. However, upcast and downcast T-S curves are similar.

There were salinity calibration samples taken, but at the time of processing the results were not available. The history of the T/C sensors suggests that the salinity is within ±0.001. Should the salinity analysis results become available these data should be re-examined to see if recalibration is appropriate.

There was no dissolved oxygen calibration sampling for this cruise and the only good recent comparison with bottles was during a cruise in February 2012. Those results were used to convert these data, but a small offset was later applied based on 2012-49 sampling in anoxic waters in Saanich Inlet.

No extracted chlorophyll sampling was available at time of processing for comparison with the ECO fluorometer. (It is unknown if such sampling was done.) In other uses the ECO was found to read higher than extracted chlorophyll when CHL<5ug/L. Inter-comparisons from other cruises suggest ECO fluorescence reads higher than SeaPoint fluorescence by roughly 50% when CHL<5ug/L. At higher CHL values the two types of sensors are generally closer to each other and to extracted CHL.

No field calibration information was available for the pH sensor and there are concerns about how the sensor performs, so the data will not be placed in the IOS Data Library at this time.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

Photocopies of the log sheets were obtained. The only problem noted was that the CTD hit bottom during cast #51. It was believed to function properly after that incident.

Bottle data were not available when processing began. There is a note in the log about salinity samples, but not the depth – likely they were fired at the bottom, but there is some confusion about the distance between the Niskin bottle and the CTD; the log on the previous cruise indicates it was 3m while 5m is the usual offset on the Ricker. There is always some error in determining the depth at which bottles close, but not knowing the distance from CTD to Niskin adds to that. Should bottle results become available this will limit the confidence we would have in the comparison with CTD data.

The cruise summary sheet was completed.

The histories of the dissolved oxygen, fluorometer, conductivity and pressure sensors were obtained.

The configuration file did not change through the cruise. All entries matched the latest factory calibrations, but the Soc and Voffset values were updated to match those found during cruise 2012-12 when there was good DO sampling. The file was then saved as 2012-51-ctd-new.xmlcon.

##### Conversion of Raw Data

All data were converted using file 2012-51-ctd-new.xmlcon.

Plots were made for a few casts. The results were similar to 2012-49 when the same configuration was used (but a different pH sensor).

- The temperature and conductivity channels track well on the downcast but the upcasts show larger differences. As noted for cruise 2012-49 when the same configuration was used, these differences look as though they are due to alignment or pump problems because when the gradient reverses sign, so do the differences. There is also a significant vertical offset that is seen only in the upcasts. Such observations are common and it is tempting to blame the wake created by the rosette package. But for this cruise, there was no rosette, and no apparent difference between casts with Niskin bottles mounted and those without. All sensors have some noisy features; it is not clear which sensor pair is better.

- The descent rate is moderate to high and generally noisy.

- Fluorescence is noisy but the vertical offset between downcast and upcast is slightly larger than that between temperature traces, suggesting fluorescence should be advanced by ~0.5s.

- Dissolved Oxygen data look ok with the usual vertical offset between downcast and upcast.

- The pH signal looks normal with some hysteresis.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels.  Parameters used were: Pass 1: Std Dev = 2,  Pass 2: Std Dev = 5, Points per block = 50

##### ALIGNCTD

Tests done for cruise 2012-49, which had the same configuration, found an advance of +4s was best.

During 2012-14 an advance of +4s was applied while for 2012-01 +3.5s was used.

ALIGNCTD was used to advance the DO signal relative to temperature by 4s. A few casts were checked before and after this step and the results were satisfactory.

##### CELLTM

The results found for cruise 2012-49 were tested on a few casts and the results were satisfactory.

CELLTM was run using (0.03, 9) for the primary and the secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration; the tau correction was applied.

on a few casts to look at differences between channel pairs.

##### Test Plots and Channel Check

There were few casts deeper than 500m and the descent rate was very noisy for those. The same sensors were used for 2011-11 in November and 2012-01 and 2012-14 in February and 2012-49 in June so data from 1 cast for each of those cruises are included in the comparison to see if there are significant differences. The test levels differ, so some of the variation is likely due to higher gradients for the shallower casts. These are very rough estimates.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-11-0053 | 380 | +0.0009 | +0.0001 | +0.0003 |  |
| 2012-01-0050 | 800 | +0.0005 | -0.0001 | -0.0020 | Mod, noisy |
| “ | 2500 | +0.0014 | -0.00014 | -0.0034 |  |
| 2012-14-0019 | 400 | -0.0001 | +0.00005 | +0.0006 | High, noisy |
| 2012-49-0057 | 575 | +0.0004 | -0.0001 | -0.0004 | High, F steady |
| 2012-51-0017 | 480 | +0.0003XN | -0.00003 XN | -0.0005 XN | High, V Noisy |

The differences are all small and similar to 2012-49. The temperature and salinity differences were noted to increase with depth during 2012-12, and the comparison with bottle samples showed some pressure-dependence in the secondary sensors. There was no salinity calibration sampling during 2012-14. The results of this cruise do not suggest any significant drift in salinity, but the evidence is weak. The other 2 summer Ricker cruises that used this equipment did not sample as deep as 500m.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

The file names correspond to consecutive CTD casts not the event numbers found in the log; the consecutive numbers are entered in the station name column. The file names were changed to match the event numbers and the consecutive numbers were kept as station names..

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. There are some negative values in fluorescence for cast #24 but these are at the bottom of the cast in a very spiky patch that should disappear when DELETE is run. There is a spike in primary temperature in event #6 but this is in the upcast, so should not be a problem.

A cross-reference list was produced and checked against log records. There were 2 incorrect station names which were fixed.

The track plots were produced and look reasonable.

The surface report indicates that the average surface pressure was 3.7db with in-water associated salinity values; this is a typical value for the Ricker when it is operating off-shore.

##### SHIFT

Fluorescence

In previous uses of ECO Fluorometers it has been found necessary to advance the fluorescence relative to pressure by as much as 48 records, but in recent uses that adjustment has become smaller due to changes in the set-up of the sensor and an advance of 6 to 12 records has been used on recent cruises. A few casts were examined and, as for 2012-49, a setting of +12 records looks appropriate to make the offset between upcast and downcast fluorescence traces look similar to that of temperature, though both channels are very noisy on the upcast.

SHIFT was run on the fluorescence channel to advance by +12 records.

Conductivity

In processing 2012-49 tests were run on 3 casts with noisy salinity. For both primary and secondary conductivity the best results in reducing the noise were those runs with the shift parameter set to -0.5 records.

All casts were put through 2 runs of SHIFT with parameter -0.5 records for both the primary and secondary conductivity. A few casts were checked to see if those settings worked well for this cruise and they did.

Dissolved Oxygen

The SBE DO data have already been aligned, but a few casts were examined to see if further adjustments are required. There are areas where the vertical offsets between upcast and downcast are slightly larger than in the temperature traces, but in other cases they are smaller.

SHIFT will not be run on this channel.

pH

The pH data are the hardest for which to determine the best alignment. Tests were run on a few casts with the aim of finding a setting that makes the vertical offset between distinctive features about the same for temperature and pH. An advance of +48 records looked best overall. (Note: This was a different pH sensor from the one used for the previous Ricker cruise, 2012-49.)

SHIFT was run using +48 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### COMPARISON WITH BOTTLES

There was no calibration sampling available at the time of processing.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

The decision about which sensor pair to use was based on the results of 2012-01since there was no calibration sampling for this cruise or other cruises using the same equipment during July and August 2012. During 2012-01 the primary salinity was closer to bottles and there was evidence of pressure dependence in the secondary sensors. So the primary will be used for this cruise as well.

CTDEDIT was used to clean the primary temperature and salinity data.

All casts required some editing.

On-screen plots of descent rate and pump status were used to guide editing.

Note was made of the editing details in the headers of the relevant files.

The edited files were copied to *.EDT files.

For casts #3, 5 and 39 acquisition did not start until the CTD was at 17db, 103db & 17db, respectively. Those files were put through REVERSE and then DELETE so that the upcast section could be edited. The output files were saved with names such as 2012-51-9003.edu so comparisons could be made with the downcast files.

Plots were made to compare the upcast and downcast where they overlap and the upcast temperatures are lower and salinity higher at the same pressure. On T-S plots the upcast data are noisier and there are some discrepancies, but mostly they look alike. So the upcast data is of some use, but definitely of lower quality. The upcast data will be selected for the IOS Data Archive, so the edited upcast files were saved as 2012-51-0003.EDT, 2012-51-0005.EDT and 2012-51-0039.EDT thus replacing the partial downcast files.

##### Other Comparisons

Previous experience with these sensors –

- Both temperature and conductivity sensors have been used on many other cruises since their last factory service, but for most the CTD data were accidentally averaged over 24 scans limiting the value of the results. They have also been used for some other cruises with no calibration sampling. They were used for cruise 2012-01 in February with good sampling and quick salinity analysis. The primary salinity was found to be low by 0.0004 while the secondary was low by 0.0028. There was evidence of pressure-dependence in the secondary temperature and salinity.

- The pressure sensor has been used 8 times since it was last recalibrated and no evidence suggests that the offset needs adjustment.

- The dissolved oxygen sensor was used for 2012-01 and there was good sampling so that calibration tests could be done to produce updated Soc/Offset values.

Historic ranges – Data were plotted with the 3-standard deviation plots of temperature and salinity ranges where local climatology was available. The only excursion from those ranges was between 10 and 30db of cast #49 where salinity was low. That was a near-shore cast where the 3-standard deviation standard is too severe. There is no evidence of poor calibration.

##### Initial Recalibration

The pressure values look ok.

Salinity samples were not available at the time of processing. Given that the samples came from a Niskin on the wire and there are some uncertainties about the distance between the CTD and bottles, it did not seem worth delaying the preparation of files for the IOS Data Archive. If the data have not yet been analyzed the results are unlikely to be trustworthy as samples degrade after 2 months with scatter increasing significantly. If the data do become available, they will be examined. If there is sufficient evidence to justify recalibration, the archived data will be corrected.

There was no dissolved oxygen calibration sampling. Evidence found during 2012-49 in anoxic waters suggested that the dissolved oxygen data should have a small offset applied, so this will also be applied to these data as well.

CALIBRATE was run to apply a correction of +0.048mL/L using calibration control file 2012-51-recal.ccf.

##### Fluorescence Processing

The recalibrated files were put through a median filter, size 11, applied to the fluorescence channel only.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the filtered files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and look fine.

Profile and T-S plots were made to check all channels. The only problem noted was an unstable feature between 8 and 10db in event #57, but examination of the upcast suggests the feature is real. No further editing was applied.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE,  SBE:pH, Status:Pump, Descent_Rate and Flag. (*.REM)

A second set of files were prepared with the pH data included. (*.REMph)

These two sets of files (with and without pH) were both put through the final steps.

A second CTD dissolved oxygen channel was derived with units umol/kg.

Dissolved Oxygen saturation was calculated and surface values were plotted. There was a lot of variability with most casts between 95% and 115%, but a few as high as 145%. None were lower than 95%. The highest saturations were late in the cruise, near the 100m contour and south of 49.5° latitude except for one cast near Brooks Peninsula. This is likely real variability rather than evidence of calibration problems, but in the absence of DO calibration sampling this is not certain. When the same sensor was used in the Strait of Georgia during 2012-49 most values were between 90% and 110%

REORDER was run to rearrange channel names to get the 2 DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence data are nominal and unedited, except that some

records were removed in editing temperature and salinity.

There was no calibration sampling for dissolved oxygen, but the

results of cruise 2012-01 were used to derive DO concentration.

A small correction was later added based on negative DO values

found during 2012-49  in the anoxic layer of Saanich Inlet.

There was no calibration sampling for salinity, but the results

of cruise 2012-01 in February 2012 suggest calibration drift

is <0.001.

The SBE pH sensor data has been removed from this file since there

was no field calibration information  available and there are concerns

about how the sensor performs.

For details on the processing see processing report: 2012-51-proc.doc.

For the files that include pH the section concerning the SBE pH sensor was replaced with:

THIS FILE WAS PREPARED FOR THE USE OF THE CHIEF SCIENTIST. IT CONTAINS

SBE ph DATA WHICH WAS REMOVED FROM THE FILE TO BE PLACED IN THE IOS

DATA LIBRARY, BECAUSE THERE WAS NO FIELD CALIBRATION INFORMATION AND

THERE ARE SOME CONCERNS ABOUT HOW THE SENSOR PERFORMS.

##### Producing final files

The Standards Check routine was run and no problems were found. The final files were named CTD and CTDpH.

A cross-reference listing was produced for CTD files.

HEADER CHECK was run and no problems were found.

A cruise track was plotted and no errors found.

The sensor history was updated for the CTD sensors.

Particulars: (notes from log book)

3. Acquisition started at about 17db. Upcast processed.

5. Acquisition started at about 103db. Upcast processed.

39. Acquisition started at about 17db. Upcast processed.

51. CTD hit bottom – worked ok afterwards.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | No | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2710 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2128 | 29Mar2011 | Factory |  |  |
| SBE 43 DO sensor | 0997 | 23Apr11 | Factory |  |  |
| SBE18 pH sensor | 852 | Dec 2010 | Factory |  |  |
| Eco-AFL Fluorometer | 2216 | June2012 | IOS? |  |  |
| Pressure Sensor | 63507 | 12Apr2011 | Factory |  |  |