## REVISION NOTICE TABLE

| DATE | DESCRIPTION OF REVISION |
| --- | --- |

# PROCESSING NOTES

Cruise: 2012-54

Agency: PBS, Marine Ecosystems and Aquaculture Division, Nanaimo, BC

Location: West Coast Vancouver Island

Project: Conservation Biology Trawl Survey

Party Chief: Boldt J.

Platform: W.E. Ricker

Date: 2 August 2012 – 13 August 2012

Processed by: Germaine Gatien

Date of Processing: 3 November 2012 – 5 November 2012

Number of original hex files: 9

Number of CTD casts processed:  9

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with an SBE43 Dissolved Oxygen sensor (#997), an SBE pH sensor #0852 and a WetLabs ECO-FL Fluorometer #2216.

# SUMMARY OF QUALITY AND CONCERNS

Log sheets were provided and a personnel list was available from the cruise report. The log contains a record of the number of bottles taken. There were notes on problems encountered.

This cruise was one of 4 Ricker cruises using the same equipment (except a change in pH sensor) between June and August 2012. The settings used for various processing steps were based on the first two of these cruises, but before and after tests were run on these data to ensure that the results were satisfactory, and they were.

There were salinity calibration samples taken, but at the time of processing the results were not available. The history of the T/C sensors suggests that the salinity is within ±0.001. Should the salinity analysis results become available these data should be re-examined to see if recalibration is appropriate, though the samples are from shallow waters where the comparison is not likely to be very useful.

There was no dissolved oxygen calibration sampling for this cruise and the only good recent comparison with bottles was during a cruise in February 2012. Those results were used to convert these data, but a small offset was later applied based on 2012-49 sampling in anoxic waters in Saanich Inlet.

No extracted chlorophyll sampling was available at time of processing for comparison with the ECO fluorometer. (It is unknown if such sampling was done.) In other uses the ECO was found to read higher than extracted chlorophyll when CHL<5ug/L. Inter-comparisons from other cruises suggest ECO fluorescence reads higher than SeaPoint fluorescence by roughly 50% when CHL<5ug/L. At higher CHL values the two types of sensors are generally closer to each other and to extracted CHL.

No field calibration information was available for the pH sensor and there are concerns about how the sensor performs, so the data will not be placed in the IOS Data Library at this time.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

Photocopies of the log sheets were obtained. The only problem noted was that the CTD hit bottom during cast #5. It was believed to have functioned properly after that incident.

Bottle data were not available when processing began. There is a note in the log about salinity samples, but not the depth – likely they were fired at the bottom, but there is some confusion about the distance between the Niskin bottle and the CTD; the log on the previous cruise indicates it was 3m while 5m is the usual offset on the Ricker. There is always some error in determining the depth at which bottles close, but not knowing the distance from CTD to Niskin adds to that. Should bottle results become available this will limit the confidence we would have in the comparison with CTD data. Moreover, these casts were all shallow and not likely to provide reliable information.

The cruise summary sheet was completed.

The histories of the dissolved oxygen, fluorometer, conductivity and pressure sensors were obtained.

The configuration file did not change through the cruise. All entries matched the latest factory calibrations, but the Soc and Voffset values were updated to match those found during cruise 2012-12 when there was good DO sampling. The file was then saved as 2012-54-ctd-new.xmlcon.

##### Conversion of Raw Data

All data were converted using file 2012-54-ctd-new.xmlcon.

Plots were made for a few casts. The results were similar to 2012-51 when the same configuration was used.

- The temperature and conductivity channels track well on the downcast but the upcasts show larger differences. As noted for cruises 2012-49 and 2012-51 when the same configuration was used, these differences look as though they are due to alignment or pump problems because when the gradient reverses sign, so do the differences. There is also a significant vertical offset that is seen only in the upcasts. Such observations are common and it is tempting to blame the wake created by the rosette package. But for this cruise, there was no rosette, and no apparent difference between casts with Niskin bottles mounted and those without. All sensors have some noisy features; it is not clear which sensor pair is better.

- The descent rate is moderate to high and generally noisy.

- Fluorescence is noisy but the vertical offset between downcast and upcast is slightly larger than that between temperature traces, suggesting fluorescence should be advanced by ~0.5s.

- Dissolved Oxygen data look ok with the usual vertical offset between downcast and upcast.

- The pH signal looks normal with some hysteresis.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels.  Parameters used were: Pass 1: Std Dev = 2,  Pass 2: Std Dev = 5, Points per block = 50

##### ALIGNCTD

Tests done for cruise 2012-49, which had the same configuration, found an advance of +4s was best.

During 2012-14 an advance of +4s was applied while for 2012-01 +3.5s was used.

ALIGNCTD was used to advance the DO signal relative to temperature by 4s. A few casts were checked before and after this step and the results were satisfactory.

##### CELLTM

The results found for cruise 2012-49 were tested on a few casts and the results were satisfactory.

CELLTM was run using (0.03, 9) for the primary and the secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration; the tau correction was applied.

on a few casts to look at differences between channel pairs.

##### Test Plots and Channel Check

The deepest cast was to 192db so the usual tests are not reliable, but that cast was examined to see if it was similar to tests done on other cruises that used this equipment. The test levels differ, so some of the variation is likely due to higher gradients for the shallower casts. These are very rough estimates.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-11-0053 | 380 | +0.0009 | +0.0001 | +0.0003 |  |
| 2012-01-0050 | 800 | +0.0005 | -0.0001 | -0.0020 | Mod, noisy |
| “ | 2500 | +0.0014 | -0.00014 | -0.0034 |  |
| 2012-14-0019 | 400 | -0.0001 | +0.00005 | +0.0006 | High, noisy |
| 2012-49-0057 | 575 | +0.0004 | -0.0001 | -0.0004 | High, F steady |
| 2012-51-0017 | 480 | +0.0003XN | -0.00003 XN | -0.0005 XN | High, V Noisy |
| 2012-54-0017 | 175 | +0.0003 | +0.00008 | +0.001 |  |

The differences are sufficiently close enough to the other cruises so as not to suggest any problem. The temperature and salinity differences were noted to increase with depth during 2012-12, and the comparison with bottle samples showed some pressure-dependence in the secondary sensors at that time.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. There are some negative values in fluorescence that will likely disappear in processing.

A cross-reference list was produced and checked against log records. No problems were found.

The track plots were produced and look reasonable.

The surface report indicates that the average surface pressure was 4.0db with in-water associated salinity values; this is a typical value for the Ricker when it is operating off-shore.

##### SHIFT

Fluorescence

In 2011 uses of ECO Fluorometers it was found necessary to advance the fluorescence relative to pressure by as much as 48 records, but in recent uses that adjustment has become smaller due to changes in the set-up of the sensor and an advance of 6 to 12 records has been used on recent cruises. A few casts were examined and, for 2012-49 and 2012-51, a setting of +12 records looks appropriate to make the offset between upcast and downcast fluorescence traces look similar to that of temperature, though both channels are very noisy on the upcast.

SHIFT was run on the fluorescence channel to advance by +12 records and the results look reasonable.

Conductivity

In processing 2012-49 tests were run on 3 casts with noisy salinity. For both primary and secondary conductivity the best results in reducing the noise were those runs with the shift parameter set to -0.5 records.

All casts were put through 2 runs of SHIFT with parameter -0.5 records for both the primary and secondary conductivity. A few casts were checked to see if those settings worked well for this cruise and they did.

Dissolved Oxygen

The SBE DO data have already been aligned, but a few casts were examined to see if further adjustments are required. There are areas where the vertical offsets between upcast and downcast are slightly larger than in the temperature traces, but in other cases they are smaller.

SHIFT will not be run on this channel.

pH

The pH data are the hardest for which to determine the best alignment. For 2012-51 tests were run on a few casts with the aim of finding a setting that makes the vertical offset between distinctive features about the same for temperature and pH. An advance of +48 records looked best overall. (Note: This was a different pH sensor from the one used for Ricker cruise, 2012-49.)

SHIFT was run using +48 records. A few casts were checked before and after this step and the results were good.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning concerned a spike in the upcast data.

##### COMPARISON WITH BOTTLES

There was no calibration sampling available at the time of processing.

##### DETAILED EDITING

All DEL files were copied to *.EDT.

The decision about which sensor pair to use was based on the results of 2012-01 since there was no calibration sampling for this cruise or other cruises using the same equipment during June - August 2012. During 2012-01 the primary salinity was closer to bottles and there was evidence of pressure dependence in the secondary sensors. So the primary will be used for this cruise as well.

CTDEDIT was used to clean the primary temperature and salinity data.

All casts required some editing.

On-screen plots of descent rate and pump status were used to guide editing.

Note was made of the editing details in the headers of the relevant files.

The edited files were copied to *.EDT files.

##### Other Comparisons

Previous experience with these sensors –

- Both temperature and conductivity sensors have been used on many other cruises since their last factory service, but for some the CTD data were accidentally averaged over 24 scans limiting the value of the results. They have also been used for some other cruises with no calibration sampling. They were used for cruise 2012-01 in February with good sampling and quick salinity analysis. The primary salinity was found to be low by 0.0004 while the secondary was low by 0.0028. There was evidence of pressure-dependence in the secondary temperature and salinity.

- The pressure sensor has been used 8 times since it was last recalibrated and no evidence suggests that the offset needs adjustment.

- The dissolved oxygen sensor was used for 2012-01 and there was good sampling so that calibration tests could be done to produce updated Soc/Offset values.

Historic ranges – Data were plotted with the 3-standard deviation plots of temperature and salinity ranges where local climatology was available. The only excursion from those ranges was between 80 and 100db of cast #13 where salinity was low. That was a near-shore cast where the 3-standard deviation standard is too severe. There is no evidence of poor calibration.

##### Initial Recalibration

The pressure values look ok.

Salinity samples were not available at the time of processing for this or any of the June-August cruises using this equipment. Given that the samples were gathered from a Niskin on the wire and there are some uncertainties about the distance between the CTD and bottles, it did not seem worth delaying the preparation of files for the IOS Data Archive. If the data have not yet been analyzed the results are unlikely to be trustworthy as samples degrade after 2 months with scatter increasing significantly. If the data do become available, they will be examined. If there is sufficient evidence to justify recalibration, the archived data will be corrected.

There was no dissolved oxygen calibration sampling. Evidence found during 2012-49 in anoxic waters suggested that the dissolved oxygen data should have a small offset applied, so this will also be applied to these data as well.

CALIBRATE was run to apply a correction of +0.048mL/L using calibration control file 2012-54-recal.ccf.

##### Fluorescence Processing

The recalibrated files were put through a median filter, size 11, applied to the fluorescence channel only. A few plots were examined before and after this step to ensure the filter was applied properly and it was.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the filtered files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and look fine.

Profile and T-S plots were made to check all channels and no problems were found.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE,  SBE:pH, Status:Pump, Descent_Rate and Flag. (*.REM)

A second set of files were prepared with the pH data included. (*.REMph)

These two sets of files (with and without pH) were both put through the final steps.

A second CTD dissolved oxygen channel was derived with units umol/kg.

Dissolved Oxygen saturation was calculated and surface values were plotted. There was a lot of variability with values between 80% and 145%. The highest saturations were are associated with higher temperature gradients in the top 15m, suggesting that the lower values are due to vertical mixing. The lowest is at the mouth of Juan de Fuca Strait where strong mixing would be expected. The 2 casts with the highest DO saturation are also the 2 with the highest fluorescence values, >10ug/L, so a bloom may account for the variability. Similar variability was noted during cruise 2012-51 and was thought to be real rather than evidence of calibration problems. In the absence of DO calibration sampling this is not certain. When the same sensor was used in the Strait of Georgia during 2012-49 most values were between 90% and 110%, which does suggest the sensor was performing reasonably well.

REORDER was run to rearrange channel names to get the 2 DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence data are nominal and unedited, except that some

records were removed in editing temperature and salinity.

There was no calibration sampling for dissolved oxygen, but the

results of cruise 2012-01 were used to derive DO concentration.

A small correction was later added based on negative DO values

found during 2012-49  in the anoxic layer of Saanich Inlet.

There was no calibration sampling for salinity, but the results

of cruise 2012-01 in February 2012 suggest calibration drift

is <0.001.

The SBE pH sensor data has been removed from this file since there

was no field calibration information  available and there are concerns

about how the sensor performs.

For details on the processing see processing report: 2012-54-proc.doc.

For the files that include pH the section concerning the SBE pH sensor was replaced with:

THIS FILE WAS PREPARED FOR THE USE OF THE CHIEF SCIENTIST. IT CONTAINS

SBE ph DATA WHICH WAS REMOVED FROM THE FILE TO BE PLACED IN THE IOS

DATA LIBRARY, BECAUSE THERE WAS NO FIELD CALIBRATION INFORMATION AND

THERE ARE SOME CONCERNS ABOUT HOW THE SENSOR PERFORMS.

##### Producing final files

The Standards Check routine was run and no problems were found. There were no negative fluorescence or dissolved oxygen values. The final files were named CTD and CTDpH.

A cross-reference listing was produced for CTD files.

HEADER CHECK was run and no problems were found.

A cruise track was plotted and no errors found.

The sensor history was updated for the CTD sensors.

Particulars: (notes from log book)

5. CTD touched bottom – hose detached where syringe attaches – reconnected. Cast looked fine.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | No | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2710 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2128 | 29Mar2011 | Factory |  |  |
| SBE 43 DO sensor | 0997 | 23Apr11 | Factory |  |  |
| SBE18 pH sensor | 852 | Dec 2010 | Factory |  |  |
| Eco-AFL Fluorometer | 2216 | June2012 | IOS? |  |  |
| Pressure Sensor | 63507 | 12Apr2011 | Factory |  |  |