## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 May 2018 | Added DIC and Alkalinity data to 2 CHE casts. For details see document Carbon_Data_Addition.docx. G.G. |
| 31-Mar-2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2012-57

Agency: OSD

Location: Strait of Georgia

Project: SoG Moorings

Party Chief: Johannessen S.

Platform: Vector

Date: July 14, 2012 – July 16, 2012

Processed by: Germaine Gatien

Date of Processing: 5 October 2012 – 15 January 2013

Number of original HEX files: 12 (10 with surface data only for bottles)

Number of CTD files: 2	Number of bottle casts:	12

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1176), a SeaPoint Fluorometer (#2356) with a 3X cable, a Wet Labs Eco-AFL/FL Fluorometer (#2215), an SBE18 pH sensor (#0691) and an altimeter (no serial # available).

The deck unit was a Seabird model 11, serial #0425. The logging computer was #2.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, but there was no mention of the surface PAR. Since this is often in the equipment list but not actually mounted, it is helpful to have it listed in the log. The CTD and rosette logs were in good order with useful notes about problems encountered.

The CTD data in rosette files are usually prepared by averaging over a 10s window centered on firing time, but that was impossible for most of the surface-only casts. Decreasing the window to 4s worked, but there is considerable overlap in the data selected for the 2 bottles as they were typically fired within about 1s of each other. A wait of 10s before firing would be helpful in assessing the stability of the data even if it can’t be included in the averages. It is understood that there is a desire for the 2 bottles to be fired very quickly. For this data set the averages are based on 4s (96 records) rather than the usual 10s window.

CTD and CHE files were prepared for casts #11 and #13; all other casts were restricted to surface sampling so were used to create CHE (rosette) files only.

The pumps were not turned on during data acquisition for events #2 and #3 and were on for only a few seconds during event #4. No surface samples are included in the comparison between bottle and CTD salinity, so this does affect decisions on recalibration. All pumped channels were removed from the CHE files for the affected casts.

Salinity bottles from the surface to 26m were flagged 3 because they were significantly out of line in the comparison with CTD data. It is possible that there was a linearity problem in the Autosal so that only low salinity samples were notably affected. The possibility that the problem is with the CTD salinity was investigated, but the evidence is not consistent with that explanation. A study was made of other cruises using this CTD in June and July 2011 and even worse problems were found for 2012-16 and a possible problem in a few bottles for 2012-05. A later cruise, 2012-71, shows similar problems. A full report on this issue can be found in document CTD550-June-July2012-sal-study.doc.

Two fluorometers were used for this cruise but there was no extracted chlorophyll sampling. Comparisons of the two fluorometers were made during bottle stops and during downcasts and both show the WetLabs ECO fluorometer reading much higher than the SeaPoint at the low end of the range, but the two are fairly close when fluorescence is between 0.5 and 3ug/L. There were no fluorescence values >3.5ug/L

Duplicate dissolved oxygen samples were taken to enable a comparison of analysis at sea with post-cruise analysis at IOS. The differences were slight, and as likely to be due to differences in kit standardization as to where they were analyzed.

A test was run to see if the new method of recalibration of DO (which forces a zero offset in fits of dissolved oxygen versus titrated DO samples) has a negative impact on how well downcast CTD data compare with upcast bottle data. The difference was negligible.

The Oxygen:Dissolved:SBE data are considered to be, very roughly (due to limited calibration sampling):.

±0.5    mL/L	 above 10db

±0.05   mL/L	 from 10 to 100db

±0.03 mL/L 	below 100db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen and salinity. The only problem noted was that the buffer was left on the pH probe during cast #7.

Nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. The same sensors were used in June 2012 during 2012-05 and the pressure was found to be low by 0.5db. A few casts were examined from this cruise and the results are consistent with 2012-05.

The calibration constants were checked for all instruments and the only error found was a date of calibration. That was fixed and the pressure offset changed to 0.74db. The adjusted file was saved as 2012-57-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using configuration file 2012-57-ctd.xmlcon.

For 6 of the 10 shallow casts, the ROS files contain only 1 bottle, Niskin #2. Since there are samples from the missing bottles, the bottles obviously closed. The BL file indicates that 2 bottles were fired. Attempts were made to convert by skipping the first records, and by using “Scans marked with bottle confirm bit” instead of using the BL file, without success. Using a shorter window did work, and it was found that a 4s window (usual is 10s) would work for all files.

The ROS files were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files.

The BOT files were bin-averaged based on bottle numbers. A preliminary header check turned up no problems with negative or off-scale fluorescence values.

Those files were used to create file ADDSAMP.csv. Sample numbers were added to that file based on information in the rosette log sheets. The addsamp.csv file was sorted on Event_Number and Sample_Number and then converted to CST files. Those will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

It was later discovered that the pumps were not on for events #2 and 3 and for only a few seconds during event #4. So the pumped channels will be removed later.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2012-57-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

- DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2012-57oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2012-57oxy.csv. That file was converted into individual *.OXY files.

- SALINITY

Salinity analysis was obtained in QF2012-57SAL.xls. The analysis was done 2.5 months after collection. The files were simplified and saved as 2012-57sal.csv. That file was then converted to individual SAL files.

- NUTRIENTS

The nutrient data were obtained in spreadsheet QF2012-57nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers and the file was saved as 2012-57-nuts.csv. The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

Salinity

Compare was run with pressure as reference channel. There are a lot of near-surface outliers. When data from above 50db are excluded plus samples where the CTD data were noisy, the CTD primary salinity was found to be low by an average of 0.0095 and the secondary low by 0.007, with standard deviations of 0.002 for both. The results are sensitive to how outliers are identified. There were no duplicates. There is insufficient information to determine if there was any time-dependence. The salinity samples waited 2.5 months before analysis which tends to lead to high bottle values, so the CTD salinity is probably more accurate than this comparison would imply. The surface samples tend to show the CTD being too high, which is likely explained by the 1.5m distance between the CTD and bottles being very significant. This is especially so for the southern casts near the Fraser River outflow where the surface gradients are higher than at SOGN.

While there were no duplicates, there were 2 Niskin bottles sampled at each level, so we do have some measure of variability. For the bottles at ~2m the differences are all >0.1 except for events #8 and 11. For event #7 the differences are 2.3. For event #6 there was a flag on one value, so the fairly large difference may not be significant. The cast to the north, #11, has a difference of 0.45 at 6m and 0.027 at 10m, but the rest are fairly small. Cast #13 in the south has differences >0.1 from 1 to 15m. The bottom bottles at cast #13 are further apart than expected at that depth, but the CTD salinity looks unusually noisy as well, so temporal variability is likely real there. The fact that the bottles are lower than the CTD is likely due to this being a near-bottom bottle. Similar results are often found for near-bottom sampling, likely due to a bottom boundary current. So the samples and CTD are likely both fine, just not measuring the same conditions.

So outliers with CTD values higher than bottles are explained by a combination of high salinity gradients near the surface and local variability near the bottom of cast #13. For the cases in which the CTD looks lower than the bottles, the explanation is not as clear. We expect some problems when samples are stored more than 2 months with some (or maybe all) bottle values being too high, with a lot of scatter in the results. These data do show bottle values being too high, but the most significantly high values (bottles higher than CTD by from 0.08 to 0.19) are all between 5 and 25m which suggests this is not due to random bottle degradation. A possible explanation might be that there is something in the near-surface samples that could not be seen by the analyst but affects the results. Some of these correspond to repeat bottles that are quite close in value, while others have large differences. The input of large river discharge may be accompanied by oil slicks or particles too small to be seen, but significant. Those could also be present in the 2m samples but be masked by the effect of the large near-surface gradient effect. It is also possible that the CTD conductivity cell was contaminated which is likely to lead to low CTD values, but we would not expect that to affect both sensor pairs so closely. Another possibility is that there is some non-linearity in the Autosal leading to errors in analysis at low salinity values. COMPARE hints that such a salinity-dependence exists but the complications of analysis delays and high salinity gradients make it impossible to say this with confidence.

Similar results were found during 2012-16 for bottles in the top 25m in Baynes Sound. The differences from CTD salinity were even higher for that cruise, but that is partly due to longer storage before analysis. Salinity from 2012-71 in October was analyzed quickly and also looks out of line. For a report on this issue including evidence from 5 cruises that used the same equipment see “CTD550-June-July2012-sal-study.doc.” Autosal non-linearity may be explain the results as there is evidence of salinity dependence in the fits.

Flag “3” was added to the samples from 0 to 26m with the comment:

SAL: Problem suspected in all samples from 0 to 26db based on comparison with CTD salinity;                possibly due to Autosal non-linearity.

Note that this comparison included a few surface bottles for which the pumps were not turned on. Excluding them does not change the comparison conclusions.

For full details for the COMPARE run see file 2012-57-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. Following recent recommendations from SeaBird recalibration will be based on a fit of differences between bottles and CTD DO with the offset forced to 0. The manufacturer warns that this will not fit well for values <2mL/L due to limitations in Winkler titrations at low values, but there are no values in that range from this cruise.

The fit of differences against DO concentration is noisy. Near-surface samples were removed because the 1.5m between CTD and bottles is likely to be significant in this region. The following fit was found:

Bottle DO = 1.055 * CTD DO

There are few bottles and the range is very small.

The slope is lower than the 1.0704 found with the same equipment in June. When an offset is allowed, the slopes look similar, but the offsets are quite different. What would account for these differences? An increase in slope would be expected if there were significant calibration drift, but this is a decrease. There is a lot of structure in the DO profiles, so that even after the usual 30s wait the temperature shows significant variability. So this comparison may not do a good job of capturing the calibration drift. The comparison is thought to measure calibration drift plus some small contribution due to response time delays. Response time becomes a bigger issue because of the variability, but we would expect some randomness in that error. But we have only 2 casts and 24 bottles available, so there is more room for error than usual. Cruise 2012-05 included casts from the same area with similar variability, but also included other areas with smoother profiles, so it is perhaps a better measure of calibration drift. However, that cruise required a second recalibration when downcast data were compared with upcast bottles, so it may have overcorrected in the first place. On balance, it seems best to use the 1.055 slope correction, as it may suit this DO range best.

For full details of the comparison see 2012-57-dox-comp1.xls.

The only significant outlier was associated with a high standard deviation in the CTD DO and a very high DO gradient in both downcast and upcast.

Fluorescence

There was no extracted CHL sampling, but there were 2 fluorometers, so a comparison was made of how those compared during bottle stops. The ratio of the WetLabs ECO fluorescence to the SeaPoint Fluorescence was plotted against the SeaPoint FL. The ECO was about 3 times the SeaPoint at low fluorescence. The ratio then fell steeply and was ~1 for fluorescence between 0.7 and 3ug/L the two values were very close. There were no fluorescence values >3ug/L. This pattern looks like that found for 2012-58 and 2012-16. A second comparison was done using thinned downcast files and the pattern is similar, though the ratio is <1 between 1 and 2.5ug/L. It is hard to determine dark values in this region but it is clear that the ECO dark value is higher than that for SeaPoint. If dark values were subtracted, then the ECO would still read high at the low end, but not by as much. However, the match in the middle range for the downcast files would be poorer.

For more detail see 2012-58-dual-fl-comp1.xls and 2012-58-dula-fl-comp2.xls.

The MRG files and SAM files were recalibrated using file 2012-57-recal1.ccf to multiply all DO data by 1.055. COMPARE was then rerun to determine if the correction looked appropriate.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no problems were detected.

At this point the data from the MRGCLN2 files were exported to a spreadsheet for comparison with the rosette files to ensure no data were misplaced or missing. No problems were found.

##### Conversion of Full Files from Raw Data

All files were converted using 2012-57-ctd.con.

A few casts were examined and all expected channels are present. The descent rate was steady and high for the two casts that sampled below the surface.

The two temperature and conductivity channels are reasonably close during the downcasts but they differ much more during upcasts. The conductivity channels have some spikes in upcasts.

Altimetry looks spiky at the bottom, so the values will be checked later. The SeaPoint fluorescence channel is lower than the WetLabs ECO fluorescence below 15m, but close or higher above that. The pH, DO and transmissivity look normal, though there is a notable decrease in transmissivity in the bottom 50m of cast #13 (SOGS) that is accompanied by reversals in DO and pH.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on a few casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge from offsets between upcasts and downcasts because the temperature is so noisy on the upcast. But looking at just downcasts and matching features in T and DO suggests an offset of 3s. For 2011 cruises when this DO sensor was used a setting of +4s was selected, but the DO range is relatively small for this cruise which may minimize the response errors. For 2012-05 which preceded this cruise a setting of 3s was found appropriate.

ALIGNCTD was used to advance the DO Voltage by 3.0s relative to the pressure.

##### CELLTM

The upcast data are extremely noisy so the usual tests for CELLTM settings are not helpful. The same equipment was used during 6 2011 cruises. Many had the same problem but the best choice overall was found to be (α = 0.02, β=7) for the primary and (α = 0.03, β=9). For 2012-05 which preceded this cruise looked best with a choice of (α = 0.03, β=9) for the primary and (α = 0.02, β=7) for the secondary. It is odd to see the reverse of the earlier results, but perhaps there was an exchange of pumps or the way the equipment was mounted. For this cruise there are only 2 casts available and they are very noisy. The secondary does look best with after a run with (α = 0.02, β=7) but the primary was hard to judge.

CELLTM was run using (α = 0.03, β=9) for the primary and (α = 0.02, β=7) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

- on a the only 2 casts that sampled to 300db to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. These are shallow casts, but a few were examined to ensure there were no obvious problems. For comparison one of the last casts from a September 2011 cruise is shown; the same T and C sensors were used then.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-17-0121 | 800 | -0.0004 | -0.00008 | -0.0005 | High, V Noisy |
| 2012-05-0028 | 315 | -0.0001 | +0.00015 | +0.0017 | Steady, high |
| 2012-05-0032 | 250 | 0.0001 | +0.00015 | +0.0018 | Steady, high |
| 2012-05-0039 | 350 | 0.0002 | +0.00014 | +0.0013 | Steady, high |
| 2012-57-0011 | 310 | -0.0001 | +0.00025 | +0.0026 | Steady, high |
| 2012-57-0013 | 270 | -0.0002 | +0.00025 | +0.0027 | Steady, high |

The differences are reasonably small, but there may some drift in conductivity.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are still some negative values in pressure despite increasing the offset by 0.5db, but they are at the end of casts when the pumps were off, so the CTD might have just left the water. For the bottle files the lowest pressure is 1.1db, when one spike is excluded.

Surface check was run and shows an average surface pressure for the cruise was 2.05db which looks about right for the beginning of acquisition for the Vector. A few casts with negative pressure readings were examined and these occur at the end of casts after or very shortly before the pumps were turned off.

The cross-reference check was compared with the log book. There was a small difference in position for one cast, but the time is also slightly different, so a little drift would explain the discrepancy.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and MRGCLN2 files were exported to a spreadsheet. For the 2 deep CLN files the entries look reasonable in plots despite spikes near the bottom. The CLN files for events #1-10 will not be processed further, so they were not corrected.

For the bottle files (MRGCLN2 and SAMAVG) the altimetry values for files #2-4 and 6-10 are wrong and were removed from the headers; for most the CTD was never within 15m of the bottom if the bottom depth recorded is correct. For cast #2 the altimetry reading is higher than the bottom depth; this could be due to drifting into deeper water, but given it is uncertain it was removed.

The Water Depth headers were also examined. For cast #13 the water depth was corrected in the log but not in the header, so the header of the CLN, IOS, MRGCLN2 and SAMAVG files were adjusted to match the log.

For bottle file #1 the water depth entry was missing but was in the log book, so it was added to the MRGCLN2 and SAMAVG file. For files #3, 4 and 8 there was no header depth entry so those were entered as listed in the log.

##### Shift

Note: From this point onwards only casts 11 and 13 will be processed for CTD files.

Fluorescence

Tests were run on the two deep casts to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence trace look like that of the temperature trace. The many bottle stops and noisiness of the upcast temperature traces makes this a difficult judgment, but the value used on other recent uses produce reasonable results. SHIFT was run on all casts to advance the ECO fluorescence channel by +12 records and to advance the SeaPoint fluorescence by +24records.

Conductivity

Tests were run on the two conductivity channels using a variety of shifts on the 2 casts and then examining the results on a T-S plot to see what setting best minimizes unstable features without oversmoothing. The results looked best overall when a shift of -0.5s was applied to the primary and a shift of -1.0s to the secondary conductivity. The shifts are the same as found during 2012-05.

SHIFT was run twice on all casts using those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. The 2 casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary. Comparing distinctive downcasts features in DO and temperature shows good depth correspondence.

pH

This sensor was not pumped so it probably does not need alignment. Plots are inconclusive because of the noisy upcast temperature so that vertical offsets are very hard to judge. When distinctive downcast features are compared with those in temperature profiles, there is no significant difference. The SHFC1 files should be saved so that this step can be reconsidered at some future date.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

Header Check was repeated on the DEL files and there remain no negative pressure or fluorescence values.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late March 2011 and were used for 2011-44, 2011-16, 2011-26, 2011-09, 2011-27, 2011-17 and 2012-05. For the first 4 of these cruises there were few calibration samples and/or the bottle calibration was not trusted. No corrections were applied. For 2011-27 both sensors were low by about 0.0015 but the samples had been stored for 2.5 months which can lead to evaporation, making the CTD look low compared to bottles. For 2011-17 which was analyzed within a month the primary was found to be high by just 0.0001 and the secondary low by 0.0004. This supports the notion that evaporation affected the earlier analysis. For cruise 2012-05, run shortly before this cruise, all samples were analyzed within 10 days of collection and the primary salinity was found to be low by 0.0012 and the secondary high by 0.0004. The analyst expressed some concerns about Autosal stability. The difference between those two results was in good agreement with the differences found between sensor pairs in tests reported for 2012-05 in section 10 of this report.

Dissolved Oxygen

The method for calibrating the DO sensors changed recently, so there was only one other cruise, 2012-05 with which a clear comparison is possible.

The DO sensor was repaired and recalibrated in March 2011. It was used for 2011-44, 2011-16, 2011-26, 2011-09, 2011-27 and 2011-17. There appeared to be some time dependence in the 2011-26 data, but not in any of the others. 2011-27 required a 2nd calibration to further correct hysteresis. For 2012-05 two corrections were applied, a first slope correction of 1.0704*DO, followed by a linear correction to profile data only which was 0.9883*DO+0.037.

Pressure

The sensor was recalibrated in April 2011 and was used for 2011-44, 2011-16, 2011-26, 2011-09 and 2011-27 and 2011-17. No further offset was applied to any of those cruises. During 2012-05 it was found that there were many negative pressure values corresponding to “in-water” CTD values. A correction of +0.5db was applied to the pressure. This produced reasonable values for near-surface bottle stops.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The temperature was a little low around 80-100db at station SOGN and the salinity was mostly low down to 30db at station SOGS. This likely reflects real conditions and similar results were found for 2012-05 in the same region and a month earlier.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration - There were no post-cruise calibrations available.

##### DETAILED EDITING

The bottle comparison is not trusted but does show the secondary salinity closer to bottles than the primary. The secondary sensors were chosen for archiving for all cruises in the past year that used this equipment. So the secondary T and S channels were chosen for archiving, and so, editing.

CTDEDIT was used to remove or clean small spikes in salinity that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface and bottom records. Editing was light.

All edited files were copied to EDT.

##### Initial Recalibration

The pressure data appear to be ok with the updated configuration file, so further recalibration is unnecessary.

The CTD salinity data appear to be low but this is likely due to the delay in analysis and a possible Autosal problem. Given excellent results with the same equipment during 2012-05, no recalibration is justified.

The dissolved oxygen data were recalibrated based on the results of COMPARE described in section 4:

Corrected DO = 1.055 * CTD DO

CALIBRATE was run using file 2012-57-recal1.ccf to apply the dissolved oxygen correction to the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. ALIGNCTD corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against DO concentration and a few outliers were excluded, there was clear pressure dependence:

Bottle DO = DO:SBE +0.0006*Pressure – 0.1716 (*)

Since this correction is larger than we usually see, a test was done to see if this was a result of forcing a zero-offset fit in the original comparison. The following fit was found when an offset was allowed:

Corrected DO = -0.0191 + 1.0598 * CTD DO

Recalibrating the SBE DO file in the thinned files using that equation and running COMPARE again produced the following fit:

Bottle DO = DO:SBE +0.0006*Pressure – 0.1749

So there is very little difference between the two fit methods. The need for a 2nd re-calibration may be due to response time errors not addressed by the alignment of DO. These were unusual profiles. A similar result was found for 2012-05 in the same region. Further use of this new calibration procedure in other regions may lead to a better understanding, but for now it does appear necessary to apply the 2nd re-calibration

CALIBRATE was run using file 2012-57-recal2.ccf to apply equation * above. Note that this recalibration is not appropriate for the bottle files.

COMPARE was then rerun using thinned files after the 2nd recalibration. The results were good with average differences in the fit <0.001 mL/L. However, the fits are all sensitive to the choice of bottles included.

(See 2012-57-dox-comp3.xls and 2012-57-dox-comp4.xls for details.)

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

The two fluorometer channels were put through a filter, size 11 to reduce spikiness.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

From this point on 2 sets of CTD files were prepared; one is intended for the OSD_Archive and the other which will include pH data will be provided to the chief scientist. The same steps will be taken with the exception of what channels are removed. The special files will have extensions that end in pH.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, pH:SBE, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity and Fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

NOte #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.50 mL/L from 0 to 10db

±0.05 mL/L from 10 to 100db

±0.03 mL/L below 100db

For details on the processing see processing report: 2012-57-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The surface values were 97% and 105%, reasonable values for this time and place.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

From this point on 2 sets of files were produced, one for the OSD Archive without pH and one for the chief scientist with pH. The pH files have extensions that end in pH.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For casts #2-4 the pumped CTD channels were also removed. Temperature:Secondary, Salinity:T1:C1, Oxygen:Dissolved:SBE and Fluorescence:Seapoint.

For the special pH files, the SBE:pH channel was also removed from casts #1 and 7 because the buffer was left on the sensor for the #7 and the values are bad for #1 so the same problem probably occurred.

Change Units was run twice to second versions of Oxygen:Dissolved:SBE and Oxygen:Dissolved channels in mass units. REORDER was run to get the 2 sets of DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and a few errors were found and fixed.

Header Check was run on the final files until all problems were resolved.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

# 1, Very low pH values – buffer probably left on sensor.

# 2 & 3. Pumps off

4. Pumps off for most of the cast.

# 7. Buffer left on pH problem.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 1Apr2011 | Factory |  |  |
| Conductivity | 3396 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2668 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2754 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | Jan2012 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 1Apr11 | Factory |  |  |
| pH | 391 | 29Dec10 |  |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| SBE Fluorometer | 2356 |  |  |  |  |
| Eco-AFL Fluorometer | 2215 | Aug2012 |  |  |  |
| Pressure Sensor | 75636 | 13Apr2011 | Factory |  |  |
| Altimeter | ? |  |  |  |  |