## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 13 Feb 2014 | Error in CHE file for event #55 – samples had been misplaced. GG |
| 27 May 2013 | Corrections made to CHL data, flags and comments based on comparisons to HPLC chl a and SeaPoint fluorescence data. |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |
| 21 Nov 2013 | Redid Nitrate/Phosphate correction – in some cases it had been applied to the wrong files so that the May 2013 corrections to CHL were missed. Updated processing report summary of TSG performance. |

## PROCESSING NOTES

Cruise: 2012-59

Agency: OSD

Location: WCVI

Project: La Perouse / WCVI

Party Chief: Robert M. (Scientist responsible for program: Yelland D.)

Platform: John P. Tully

Date: August 31, 2012 – 11 September, 2012

Processed by: Germaine Gatien

Date of Processing: 17 January 2013 – 26 February 2013

Number of original HEX files: 102	Number of CTD files: 101 (1 file had upcast data only)

Number of bottle files: 50		Number of bottle casts processed: 47 (3 not sampled)

Number of original TSG files:   2	Number of processed TSG files:

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1119), a SeaPoint Fluorometer (#2228), a Wet Labs Eco-AFL/FL Fluorometer (#2214), a Biospherical QSP-200L4S PAR sensor (#4601) and an altimeter.

A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #0603 and a flow meter.

The data logging computer was #3.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

A IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, plus details about the rosette and TSG. The logs were clear with warnings about problems.

Files from the first cast in Saanich Inlet were sent to IOS for initial processing. This is a useful step as it can ensure reasonable data are being logged on all channels and that correct data are being displayed at sea.

The pumps were not turned on until 45m of the first cast, so the upcast data were used for archiving; data quality is lower than usual.

The usual protocol of turning on acquisition only after a return to the surface after the 10m soak was not used for many casts. This makes processing more complex. The casts involved a particular watch, so is clearly a matter of training.

The primary salinity was found to be low by an average of 0.0047 and the secondary by 0.0015. Salinity samples were analyzed about 2 months and some evaporation of samples is expected by then leading to higher bottle salinity values. The CTD salinity is probably more accurate than suggested by the comparison. Results of the previous cruise which had samples analyzed a little more promptly showed the two salinity channels to be low by averages of 0.0025 and 0.0002.

Two fluorometers were mounted for this cruise, a SeaPoint and a WetLabs ECO fluorometer and there was good extracted chlorophyll sampling enabling a good comparison. The ECO read higher than the SeaPoint by a factor that averaged about 2 for very low extracted CHL and that factor decreased slightly as CHL values rose reaching 1.7 for CHL>5ug/L. The SeaPoint fluorometer read much higher than extracted CHL for very low CHL values but was reasonably close for 2<CHL<5. For CHL>5ug/L the SeaPoint fluorometer read about 20% lower and the ECO 20% higher than CHL. These results differ from those found for other ECO fluorometers; this is the first good comparison for this particular sensor.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.8mL/L from 0 to 50db

±0.3mL/L from 50db to 100db

±0.1mL/L below 100db

The recent history of the TSG system suggests caution before recalibrating any data. There have been some issues with clogged filters, poor flow and poor performance of the intake temperature. The intake temperature appears to have performed well for this cruise but the flow problems are not fully resolved and the filters got clogged again. How poor flow might affect salinity is unknown. The CTD salinity calibration is thought to be ok but there are some issues concerning storage of samples that make that judgment less confident than usual. The fact that intake temperatures are a little high might suggest that the water samples came from a little higher in the water column, but having salinity also reading too high argues against that. There are too many unknowns to consider recalibration.

TSG fluorescence data are given in both calibrated (ug/L) and raw (volts) channels. The calibration was based on comparison with a SeaPoint fluorometer when both instruments were mounted on a CTD during 1 cast of cruise 2012-13.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The files from the initial cast in Saanich Inlet files were sent to IOS for initial processing and the only potential problem found was with the offset to the pressure calibration. The value used at sea was from the last factory calibration, but an updated value would lower pressures by about 1db. This was not changed at sea, but the CTD operators were probably aware the pressures were reading low.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as various analysis logs. A few problems were noted in the log. There was one split cast, file #83 contains only downcast data and cast #84 has only upcast data, so the only remedial action required is to rename the bottle file as #83. Cast #79 has an error in the cruise # in the header. Many casts include data from an initial drop to 100m followed by a return to the surface, then a full cast. The initial drop will need to be removed before DELETE is run.

Extracted chlorophyll, NH4, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments.

The pressure offset was changed from -1.28 to -0.1 based on observations during previous cruises.

Hysteresis parameter E in the dissolved oxygen calibration was changed to 0.0385 based on tests during 2012-13.

The PAR sensor is identified as #4601 in the log but the calibrations and sensor # in the configuration file is #4615, The cruise that preceded this one has #4601 in both the log and the configuration file and has the correct parameters for #4601. The CTD technician who prepared the equipment says there was no change between cruises and that the log entry was prepared by visual check of the instrument so that should be trusted. Tests were run using the 2 different configurations to see if the results make it obvious which is correct. Both have reasonable shapes but different values.

The configuration file was changed by entering the parameters for #4601.

The PAR sensor was not always mounted. Based on log notes cast lists were prepared with and without PAR so that it will be easy to remove PAR as appropriate.

##### Hysteresis Study

Hysteresis tests were not done as the deepest casts were to 2000db and tests had been run on the previous cruise, 2012-13.

##### BOTTLE FILE PREPARATION

The ROS files were converted using file 2012-59-ctd.xmlcon.

Those files were put through CLEAN to add event numbers (*.BOT).

Header Check was run on the BOT files and the only potential problem noted is that the WetLabs fluorometer channel has negative values in at least cast #1.

Temperature and salinity were plotted for all BOT files. Casts #40 and 55 had small spikes near the surface in the primary salinity channel. CTDEDIT was used to remove the spiky points and the edited files were copied to BOT.

The BOT files were then averaged to enable an ADDSAMP file to be prepared. Sample numbers were added to the ADDSAMP file based on rosette log records. A few problems were encountered:

- Cast #1 – a bottle was fired but no sample number was assigned. Do not process.

- Cast #62 – complex – check after bottle files assembled.

- Cast #103 – bottles fired but no sample numbers or sampling recorded. Do not process.

- Cast #143 – bottles fired to get bulk water. No sample #s. Do not process.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was sorted on Event_Number and Sample_Number and then converted to CST files. The CST files will form the framework for the bottle files. For cast #55 the CST file had to be reordered into sample # order.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2012-59-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2012-59_chl.xls. A simplified version of the spreadsheet was saved as 2012-59_chl.csv which was then converted to individual CHL files. Loop data were removed from that file but were added to file 2012-59-loops.csv.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2012-59oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2012-59oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2012-59SAL.xls. The file was simplified and saved as 2012-59sal.csv. That file was converted to individual SAL files. Loop data were removed, but were first added to file 2012-59-loops.csv. The salinity data were analyzed 2 months after collection.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2012-59nuts.xls which included a report on precision.  The file was simplified, reordered on sample numbers and saved as 2012-59-nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2012-59_NH4.xls which included a report on precision. The file was simplified and saved as 2012-59-NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

Data from the SAMAVG files were exported to a spreadsheet and a few casts were checked against rosette sheets to ensure all expected bottle data were present. Cast #62 was checked to see if the assignment of sample numbers to bottle numbers was correct. The CTD and bottle DO profiles are close, so this is likely fine. No missing data were noted.

##### Compare

Salinity

Compare was run with pressure as reference channel.

An error in the ADDSAMP file was found, so that was fixed, the merges repeated and COMPARE rerun.

There was only 1 major outlier that was not explained by a large standard deviation in the CTD salinity. That was sample #390 which had been flagged “36” with replicates that difference by >0.05. One of the replicates differs from the CTD by about 0.05 while the other lies between the primary and secondary CTD salinity. So one value was rejected and the value 34.5095 was entered with flags “26” and amended comment. Other outliers are mostly associated with high standard deviation in the CTD salinity. There is more scatter than expected. Pressure dependence is slight below 500m especially for the secondary salinity, but there is a lot of scatter making this a weak conclusion.

When bottles above 500db are excluded the primary salinity was found to be low by an average of 0.0047 and the secondary by 0.0015. Standard deviations were ~0.002 for both channels. Salinity samples were analyzed about 2 months after collection and some evaporation of samples is expected by then, leading to higher bottle salinity values. The CTD salinity is probably more accurate than suggested by the comparison. Results of the previous cruise which had samples analyzed a little more promptly showed the two salinity channels to be low by averages of 0.0025 and 0.0002.

There is some hint of time-dependence but if bottles above 500m are excluded it disappears. This is because cast #102 is the only cast with sampling above 500db. This cast has a complex temperature structure in the top 300m which may account for why the CTD data are noisier than usual. The primary salinity is much noisier than the secondary.

Tests of non-linearity are inconclusive. There is too much noise to judge this, particularly in the primary channel. There is a slight hint of it in the secondary with the CTD reading lower relative to bottles at lower salinity values. This could be due to non-linearity in the Autosal or poor flushing of Niskin bottles which becomes more obvious in higher salinity gradient regions.

The only significant outlier that is not associated with a high standard deviation in the CTD salinity is sample #401 at 127db during event #102. While the CTD data are quiet during the 10s captured in the BOT file, a wider examination shows a lot of variability and vertical motion during the stop. The temperature structure is complex in this area with a local temperature minimum around 105m and a maximum around 130m. The mismatch between bottle and CTD is likely due to variability rather than a problem with the bottle. No flag was assigned.

For full details for the COMPARE run see file 2012-59-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

As expected, the fit of differences against DO concentration is quite linear if an offset is allowed. SeaBird suggest setting the offset to 0, but as found during 2012-13 that leads to a poor fit. The reason given for not allowing an offset is that the titrations are not reliable enough at low values. But even if we exclude all bottle values <1mL/L from the fit, the result does not change significantly, so this offset seems essential for a good fit. When outliers are excluded based on residuals being <-0.1 or >0.1, the fit with the offset set to zero is:

DOX_BOT  = 1.0540

The fit with free offset including values <1mL/L is:

DOX_BOT  = 1.0463 * DOX_CTD +0.0331

The fit with free offset excluding values <1mL/L is:

DOX_BOT  = 1.0470 * DOX_CTD +0.0302

The fit with free offset excluding values <0.5mL/L is:

DOX_BOT  = 1.0465 * DOX_CTD +0.0329

The result used for 2012-13 when the same equipment was used was:

DOX_BOT  = 1.0435 * DOX_CTD + 0.0159

The largest outlier comes from a surface sample for which there were 2 bottles sampled at about 6db. The CTD values are close but there is a larger difference in the bottle samples. The shallower bottle (6.4db) compares well with the CTD DO but the DO from the lower bottle (6.6db) is lower. The nutrients also look quite different with the shallower bottles having much lower values. It is interesting that for the lower of the two the CTD temperature and salinity channel pairs are also in poor agreement, so there was likely some local variability that the DO sensor did not capture due to its slow response. So there is no evidence of a problem with the bottles.

Flagged bottle samples were examined in light of COMPARE results and no evidence was found to suggest any changes to flags. Some were minor outliers but the CTD data were noisy or local gradients high enough that the 1.5m offset between bottle and CTD could explain the difference. Others may be ok but again the local variability makes it impossible to determine that flags should be changed.

One comment suggested that sample #220 from cast #58 had a value that seemed very high. The CTD DO is a little lower than the bottle but close enough that a flag is not justified. This was close to the surface where the vertical offset plus high local gradients can easily explain the difference.

Other outliers that had not been flagged by the analyst were investigated. Most had high standard deviation in the CTD data or were easily explained by local gradients. Several were near reversals in the DO gradient complicating interpretation greatly. No clear case of a bad bottle sample was found.

There is no hint of time dependence in the fits.

A test for hysteresis was run. There are few deep bottles, with 3 at 2000db being the deepest. When a fit was done against CTD DO and excluding the bottles below 900db, one of the 2000db bottles looks a little out of line but the other two are close to the fit for the shallow bottles. So, it is possible that there is some hysteresis, but it is not clear and there are insufficient data available to fine-tune the setting obtained during 2012-13.

For more details see 2012-59-dox-comp1.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

Fluorescence

COMPARE was run using the Wet Labs ECO CTD fluorescence, the SeaPoint fluorescence and the Extracted Chlorophyll from bottles. The COMPARE fit sheets for ECO vs CHL and SeaPoint vs CHL were combined in a single worksheet. The aim of this study is to see if we can find a reliable way to relate data from the newer WetLabs ECO fluorometers to older data collected with SeaPoint fluorometers. A variety of plots were prepared to look for patterns and these emerged:

- In a plot of one fluorometer versus the other, the ECO reads higher than the SeaPoint. The relationship looks roughly linear for FL<5ug/L:

WetLabs FL = 1.56 * SeaPoint FL +0.46

However, a polynomial fit would be better as the ratio of WetLabs to SeaPoint starts at roughly 2 and gradually reduces to about 1.5 as extracted chlorophyll values rise. There is a lot of scatter so these are rough conclusions.

- The linear fit of WetLabs fluorescence against extracted CHL has a slope of 1.12, but it looks like a bad fit. If an offset is allowed the fit is WetLabs = 0.8907* CHL + 2.0939. For the SeaPoint the linear fit has a slope of 0.68 but again looks like a poor fit. With an offset allowed we get SeaPoint = 0.5638*CHL + 1.0692. These large offsets make very little sense, and likely derive from the fact that so much of the source data come from near the origin.

- If instead of looking at a fit we look at averages the picture is different. Breaking those averages into groups based on the CHL bottle values is most interesting.

| # of bottles | CHL range ug/L | CHL |  | SeaPoint/CHL | WetLabs/CHL | WetLabs/SeaPoint |
| --- | --- | --- | --- | --- | --- | --- |
| 102 | 0-28 | 2.36 | 3.5 |  | 6.7 | 1.9 |
| 32 | 0-.05 | 0.114 | 7.3 |  | 14.2 | 2.0 |
| 27 | 0.5-1 | 0.506 | 3.0 |  | 5.5 | 1.9 |
| 10 | 1-2 | 1.45 | 1.6 |  | 2.9 | 1.9 |
| 15 | 2-5 | 3.20 | 1.1 |  | 1.9 | 1.9 |
| 18 | >5 | 8.93 | 0.8 |  | 1.2 | 1.7 |

Using all data the ratio of WetLabs to SeaPoint is fairly steady only dropping off noticeably when CHL is >5ug/L.

The ratio of the WetLabs and SeaPoint fluorometers to extracted CHL are 3.5 and 6.7 using all data. The SeaPoint does a better job of matching CHL between 0 and 5ug/L while above that the Seapoint is low by 20% and the WetLabs high by 20%. Neither fluorometer matches CHL well below 2ug/L and especially not below 0.5ug/L. While comparing two noisy data sets is bound to be problematic, there is a trend that fluorometers read much higher at the low end. Is this due to what they are measuring or errors in one or both data sources?

This is the first comparison of a SeaPoint and ECO #2214 that includes extracted CHL sampling. It has been suspected that the different ECO sensors behave differently and this certainly looks different from the results of 2012-16 with a different ECO sensor. Both of those cruises included good CHL sampling.  While the details differ, both studies found better overall results from the SeaPoint when compared to extracted CHL.

Studies to date suggest that the SeaPoint is more likely to be close to extracted CHL, but that there are clearly many possible sources of variations from CHL range, individual sensors, gain on the SeaPoints and health of the biota and other seasonal effects. While this particular study might suggest that the ECO could be recalibrated by divided by 1.9, that does not look suitable for other data sets that have been analyzed. So advice continues to be that for long-standing programs it is wise to use the SeaPoint fluorometers.

##### Conversion of Full Files from Raw Data

All files were converted using 2012-59-ctd-new.con.

A few casts were examined and all expected channels are present. The descent rate was kept high, on average, but is frequently noisy enough that there is likely to be shed wake corruption..

The two temperature and conductivity channels are close during the downcasts but, as usual, noisy on the upcasts. There are some spikes and odd excursions in the primary channels, mostly during upcasts.

Altimetry looks useful; dissolved oxygen, fluorescence, PAR and transmissivity look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 3 casts to determine the offset between the DO channel and the primary temperature. It is very hard to judge because the temperature is noisy on the upcast, but overall a 4.5s advance looked best, though slightly higher and lower settings looked best for some features.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

The upcast data are noisy so the usual tests for CELLTM settings were a little hard to interpret. Tests were run on a few casts using a variety of settings. The differences among settings are slight.

CELLTM was run using (α = 0.02, β=7) for the primary and (α = 0.03, β=7) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are very noisy so these are very rough estimates and if there was a spike at the given depth, nearby values were chosen. Results from a few casts during 2012-13 are included for comparison.

There is some increase in the differences in conductivity and salinity. The results are at least roughly consistent with the bottle comparison which implies an average difference of 0.0032.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2012-13-0052 | 500 | +0.0002 | +0.0001 | +0.001 | Noisy, high |
| “ | 1000 | -0.0004 VN | +0.00009 | +0.0014 | “ |
| “ | 1800 | -0.0006 | +0.0001 | +0.0018 | “ |
| “ | 3500 | -0.0012 | +0.00008 | +0.0023 | “ |
| 2012-13-0066 | <1800 | Too Noisy | Too Noisy | Too Noisy | X Noisy, high |
| “ | 1800 | -0.0006 | +0.0001 | +0.002 | V Noisy, high |
| “ | 3500 | -0.0012 | +0.00012 | +0.0027 | “ |
| “ | 4000 | -0.0013 | +0.00012 | +0.0028 | “ |
| 2012-59-0033 | 500 | +0.0004 | +0.0003 | +0.0032 N | V Noisy, high |
|  | 1000 | 0 | +0.0028 | +0.0033 |  |
|  | 1800 | -0.0005 | +0.0027 | +0.0039 |  |
| 2012-59-0102 | 500 | +0.0004 VN | +0.0003 | +0.0032 N | X Noisy, high |
|  | 1000 | -0.0002 VN | +0.0003 | +0.0036 |  |
|  | 1800 | -0.0006 XN | +0.0003 | +0.0042 |  |

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are some negative values in the ECO fluorescence, as usual. After editing, this check should be done again and any remaining negative values can be replaced with pad values. No other problems were noted.

Only cast #65 had negative pressures. Those occurred at the end of the cast with pumps off. While the salinity looks like the data came from below the surface this can’t be trusted with pumps off. The transmissivity and SeaPoint fluorescence suggest that the surface is between -0.1 and -0.2db. There is insufficient evidence to justify changing the offset and any error looks to be at about the level of the initial accuracy and resolution of the instrument.

Surface check was run and shows an average surface pressure for the cruise was 2.9db which looks reasonable for the Tully. Very few casts have well-mixed surface waters.

The cross-reference check was compared with the log book and no problems were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and MRGCLN2 files were exported to spreadsheets. Some casts did not get within 15m of the bottom so there are no header entries. A selection of casts were checked and no problems were found in the altimetry headers for the CLN files. The altimetry header was removed from some SAMAVG files (casts 112, 129, 131, 135, 138, 140 & 142) for which there were entries but no sampling within 15m of the bottom. It is irrelevant information for these files and it may imply the sample came from close to the bottom.

The water depth header entries were checked against log entries; discrepancies were investigated and in all but two cases the header looks more accurate. The water depth header was changed to match the log entries in the CLN files for casts #72 and 151 and the SAMAVG file for cast 72.

##### 13. Shift

Fluorescence

Tests were run on two casts to see what SHIFT value should be used to make the offset between the downcast and upcast fluorescence traces look like those of the temperature traces. This task was complicated by extremely noisy upcast temperature and noisy ECO fluorescence.

Tests indicate that applying the usual advance of +24 records is effective in bringing the SeaPoint fluorescence into line with the temperature. For the ECO fluorescence the s setting of +6 records looks best overall.

SHIFT was run twice on all casts to advance the ECO fluorescence channel by +6 records and the SeaPoint fluorescence by +24records.

After this step a few plots were made to see if peaks in the 2 fluorescence traces are in reasonable agreement during downcasts and they are.

Conductivity

During the previous Tully cruise, 2012-13, when the same sensors were used the best results for this step were with an advance of -1 records for the primary conductivity and +0.4 records for the secondary. Tests were run on 2 casts to see if these settings worked well for these data and they did.

Two runs of SHIFT were used to apply those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The results are too complicated to interpret, but there is no evidence to justify further adjustments to the alignment.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings were from upcast portions of files, so of no concern.

Some files required special treatment.

Cast #1: The pumps were not turned on until the CTD was at 47db so upcast data were chosen for archiving. The SHFC1 file was put through REVERSE and then DELETE.

Many files included data from an initial drop to about 12m followed by a return to the surface, then a full cast. Records from the initial drop were removed from the SHFC1 file and then DELETE was rerun for that cast only. The affected casts were: 20, 24, 25, 26, 29, 43, 45, 46, 48, 50, 68, 70, 81, 83, 85, 96, 102, 114, 116, 119, 120, 122, 142, 143, 144, 146, 153, 157, 160.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The conductivity sensors were both recalibrated in late March 2011 and were used for many cruises since then. There have not been many good comparisons with bottles; the best available was from 2012-05 and 2012-13. The first, analyzed after 1 week, indicated that primary salinity was low by about 0.001 and the secondary was high by about 0.0004.  The latter, analyzed about 6 weeks after collection, showed the primary to be low by 0.002 and the secondary low by 0.0005. The difference between those results may be due to storage time or drift.

2. Dissolved Oxygen

The DO sensor was recalibrated in April.2011 and has been used for 9 cruises since then. Only 2012-13 used the bottle comparison method now recommended by SeaBird. Recalibration used a linear correction with a slope = 1.0435 and offset = 0.0159

3. Pressure

The pressure sensor has been used 7 times in 2012 and the offset was increased from -1.4 to -0.1 in April 2012.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. These are not appropriate for the casts closest to shore. The temperature data were all within the climatology except for some low temperatures for 2 casts in the northern Strait of Georgia at about 140db. The offshore salinity profiles were entirely within the climatology, but close to shore there were many excursions to the low side. These were mostly in the top 10 to 20m during the parts of the SS and CS lines closest to shore. There were a few deeper excursions in salinity in Rivers Inlet and the northern Strait of Georgia. These are likely due to real variations rather than instrument calibration problems.

Repeat Casts –

There were no repeat casts deep enough tobe useful for checking sensor repeatability.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

From this cruise the bottle comparison shows the primary salinity to be low by an average of 0.0047 and the secondary by 0.0015. The analysis was run about 2 months after collection. For 2012-13 when the analysis was a little quicker, the primary was found to be low by 0.002 and the secondary low by 0.0005. For 2012-13 the secondary sensors were selected for archiving. That looks like the best choice for this cruise as well because the CTD data are closer to bottles and there is a little more noise in the primary salinity than the secondary.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure looks ok.

The secondary salinity was found to be lower than bottles by an average of ~0.0015 below 500m. Given there was a 2 month delay in analysis no recalibration will be applied since the error due to that delay is expected to be on the same order.

No recalibration will be applied to either fluorescence channel.

SBE Dissolved Oxygen data will be recalibrated using the fit based on data excluding outliers and cases where SBE:DO<0.5mL/L:

DOX_CTD Corrected = 1.0465 * DOX_CTD +0.0329

File 2012-59-recal1.ccf was prepared to apply the DO correction.

CALIBRATE was run on the MRGCLN2, SAM and EDT files using that file.

COMPARE was rerun to check that the DO correction was done correctly and it was. The file is 2012-59-sal-comp2.xls. The average difference (after CTD DO<0.5mL/L and outliers are excluded) shows the CTD DO to be within 0.0002 of the bottles.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. ALIGNCTD corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. When the differences were plotted against DO concentration there was a lot of scatter, partly because some flagged samples are included. The near-surface values are always highly variable but are expected to be especially so for this cruise because there are many casts in areas with active mixing that will lead to large differences between upcast and downcast values especially near the surface. Some flagged values were removed as well as bottles from above 13db and with DO<0.1mL/L (that picks up cases with no CTD data in the bins) and cases with differences >0.6mL/L. The average difference of the remaining cases indicates that the CTD DO is low by 0.013 but the median difference indicates it is high by 0.002. Given the scatter no further recalibration will be applied.

This comparison is used to make a rough judgment of the accuracy of downcast CTD dissolved oxygen.

##### Special Fluorometer Processing

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channels in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

A header check was run to see if the WetLabs ECO fluorescence has negative values. There are some negative values in the EDT files in isolated spikes. but not in the bin-averaged files.

##### Final CTD File Steps (REMOVE and HEADEDIT)

A check was made of the cast lists with and without PAR sensor by plotting PAR. Two errors were found. Casts #77 and 78 have a PAR signal though the log does not indicate that the sensor was mounted. The lists were updated to be used in the REMOVE step.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

Profile and T-S plots were produced at this point to check for errors. A few problems were investigated:

Cast #31 – There is a spike in WetLabs fluorescence at 1315db but no such spike in the SeaPoint fluorescence. This was present before filtering but probably was spread by the filter. There are some very large values and some small negative values. The bad data were replaced with pad values in the AVG file.

Cast #86 – ECO fluorescence is very high around 1112db while SeaPoint fluorescence does not vary. This looks unreasonable so the value was replaced with a pad value in the AVG file.

Cast #88 – There is a patch of near-zero transmissivity data. This is not just a spike, so was left unchanged.

Cast #135 - There was a spike in SeaPoint fluorescence during cast #135 at about 25m. This is not seen in the ECO fluorescence, but there are some features in temperature and salinity with local reversals and suggestions in the T-S plot of local mixing. The reversals might be due to problems with plumbing but they are seen in both channel pairs and don’t really look like the usual sort of response to such problems. So no editing was applied in case this feature is real

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section, to fix the project name and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, fluorescence and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

SBE DO calibration was done using the method described in the SeaBird

Application Note #64-2.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.4mL/L from 0 to 300db

±0.1mL/L from 300db to 800db

±0.05mL/L below 800db

For details on the processing see processing report: 2012-25-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged widely with most between 90% and 110% but higher values were found in the middle of the LC and LB lines (mostly between 110% and 130% but occasionally >130%) and in the middle of Queen Charlotte Sound. The saturation was >130% close to shore on the LD line. Values between 70% and 90% were found in Rivers Inlet area and close to shore near the north end of Vancouver Island. Values <70% were found at one cast near Rivers Inlet and in Saanich Inlet. A few of the outliers were checked by ensuring bottles agreed with the CTD DO data and in every case checked there was good agreement.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. A few problems were found with a few samples in casts #2 and 3 misplaced and the nutrients from cast #46 missing due to wrong event number in spreadsheet. The bottle processing steps were repeated for those casts.

A header check was run and one error was found and corrected; HEAD EDIT was rerun and no further problems were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up errors in station name format; those were fixed.

The track plot was produced on screen and no further errors were found.

##### 25. Thermosalinograph Data

Data were provided in 2 hex files.

There was no loop sampling.

a.) Checking calibrations

The calibrations were checked and the only problem concerns the fluorometry. There is no record of the calibration information entered. There is an old calibration that is known to be inadequate, producing very high values. In January 2012 there was a cleaning and recalibration, but it produced very bad results.

During 2012-13 calibration information was obtained by mounting the TSG fluorometer on the CTD for one cast. This led to an estimate of the scale factor as 7.211 and blank output as +0.073.

Those parameters were entered into the configuration file for this cruise. Voltage will also be archived so users may choose whether to use either raw voltage or concentration units based on the 2012-13 calibration.

After that change the CON file was saved as 2012-59-tsg.con.

b.) The files were converted to CNV files using configuration file 2012-59-tsg.con. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Time-series plots were produced. The results look ok except for one very noisy patch in the 2nd file.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2012-59-ctd-tsg-comp.xls.

60 CTD casts with data from 4m overlapped with TSG data. The 2 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all <0.002° and the median differences in both are 0.0000°. This shows both the times and positions are reliable for both systems. (The differences are smaller if single values are chosen for the TSG rather than the medians over 2 minutes, presumably due to significant ship drift in some areas.)

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected throughout the cruise and the data appear to be much better than from other recent uses when the intake temperature was generally higher than the lab temperature. The outliers in a plot of differences against flow rate show some to be related to low flow rate but others are not. A plot against event numbers indicates the outliers are common in areas near shore where local gradients are high. It is likely that the temporal mismatch between the lab and intake records is significant there.

For file #1 the lab temperature was higher than the intake temperature by an average of 0.23C° and a median of 0.20C° and a standard deviation of 0.31C°. The average flow rate was 1.0.

For file #2 the average, median and standard deviations in the differences were 0.22C°, 0.21C° and 0.38C°. The average flow rate was 1.05.

It is surprising to see the difference being lower in the more northerly part of the cruise where intake temperatures are lower. Heating in the loop is expected to be rise as difference from ambient ship temperature increases. However, the first file was one where the flow rate was lower, so that may account for the difference; moreover the 2nd file includes considerable time in the Strait of Georgia where temperatures were higher.

During stops for CTD casts the Lab Temperature was higher than the intake by an average of 0.199C° and a median of 0.204C° with a standard deviation of 0.071C°.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

1. Intake Temperature The difference between the TSG intake temperature and the CTD temperature was noisy. Plots show that while many outlier differences are associated with low flow rate, others are not. Some outliers are better explained by geographic variations coming from near shore where higher near-surface gradients are likely, especially in the Rivers Inlet area. In some cases both low flow rate and high surface gradients were both present. Outliers are not well removed by using standard deviation in the TSG temperature alone, perhaps because of the flow rate issue. So outliers were identified by first removing cases where flow rate was <0.8 or >1.2, then cases where CTD and intake TSG temperatures differed by >0.4. The remaining cases were ordered on the standard deviation in the intake temperature and a plot then made it clear that the differences were fairly steady until the standard deviation was >0.03. Using those 17 records, the TSG intake temperature is higher than the CTD by an average of 0.008C° and a median of 0.009C° and a standard deviation of 0.024 C°. When only the 11 casts that were well-mixed to 10db (as indicated by salinity changing by <0.005 between 4m and 10m) the temperature looks to be high by a median of 0.009C° with a standard deviation is >0.013C°. We can conclude that the intake temperature may be a little high but likely by no more than 0.01C°.

2. LAB TEMP Using the same data as in the analysis of the intake temperature the lab temperature is found to be higher than the intake temperature by 0.23C° with a standard deviation of 0.05C°. This is a little higher than found in part 1 of this section, possibly because the average intake temperature may be higher when underway then when stopped for CTD casts because water is drawn from higher in the water column when the ship is moving. Another explanation may be that there were few CTD casts in the Strait of Georgia section where the water was warmer than average for the cruise. The warming in the loop is larger when the intake temperatures are further from the ambient ship temperature.

3. SALINITY Using the same data as used in the analysis of the intake temperature, the TSG salinity is higher than the CTD salinity by an average of 0.040 and a median value of 0.0.036 and a standard deviation of 0.024. It is impossible to determine if this is a measure of error in the TSG, since the variable flow rate will affect the comparison as will any systematic mismatch in depth of sampling from the CTD and the level from which the water in the TSG actually comes.

4. FLUORESCENCE Again using the data from part 1 of this section, the TSG fluorescence is found to be an average of 1.02 times the SeaPoint fluorescence from the CTD with a median of 0.94 times and a standard deviation of 0.13. It is an average of 0.63 times the WetLabs fluorescence from the CTD with a median of 0.61 and a median of 0.05.

5. Intake Temperature during Low Flow While it seems that low flow rates should not affect the intake temperature, the few casts during which flow was low were examined to see if there is any evidence to suggest there is a problem. A plot of differences between intake temperature and CTD temperature during well-mixed casts has a lot of scatter with no obvious trend. The one outlier is from a cast with the not unusual Pacific character of being well-mixed in salinity but not so much in temperature.

(See 2012-59-ctd-tsg-comp.xls.)

Loop Bottle - TSG Comparisons There were no loop samples

- Surface rosette samples vs TSG fluorescence

A comparison was made between the TSG raw fluorescence and the Niskin bottle CHL data. This fit is not expected to be great because the times for the TSG data are from the start of casts, not the time of the rosette sampling and the TSG samples are likely from higher in the water column than the rosette samples. There is a lot of noise in the fit. When samples flagged “3” and “4” are excluded and a trendline is forced through the origin, the TSG fluorescence is found to equal 0.69 times the CHL. The ratio of TSG fluorescence / Extracted CHL is high when CHL <0.3ug/L, look close at about 1ug/L and appears to be approaching 0.5 for values >4ug/L. However, there is a lot of noise around that trend.

(See 2012-59-ctd-tsg-comp.xls.)

- Calibration History

The TSG primary temperature and conductivity were recalibrated in March 2011 and have been used during 2012-01, 2012-14, 2012-25, 2012-12, 2012-31 and 2012-13. There was no intake temperature available on the first 2 of those cruises, and it worked badly for some of the others. There was no calibration sampling on 2012-31.

During 2012-25 and 2012-12 there were problems with flow, bad intake temperature and apparent shifts in relationship of CTD and TSG salinity. For 2012-13 many of those problems continued but there were a few files with useful intake temperature and from those the heating in the loop was estimated to be about 0.18Cº. Looking further back to see how much heating is expected in the loop in late summer in this region, values found were 0.14 and 0.17C°.

The fluorometer calibration parameters were found to be wrong for all 2012 cruises (producing many negative values), so only fluorescence voltage was archived for the first 5 of those cruises. During 2012-13 the TSG fluorometer was mounted on the CTD for one cast so it could be compared with the SeaPoint fluorometer used for all the CTD casts. The comparison was used to derive slope and offset to apply to the TSG fluorometer. The results looked reasonable for 2012-13 but based on comparisons with loop and with CTD fluorescence, the TSG fluorometer tends to read too high for low CHL values (<0.9ug/L) and too low for high values (>5ug/L).

The TSG salinity was very noisy during 2012-01; it appeared to be low by 0.035 but the standard deviation was 0.03 in that comparison. During the later cruise, 2012-14, the TSG salinity was highly variable but appeared to reflect real conditions. It was found to be low by a median of 0.007. While the salinity was less noisy for the second cruise, there were fewer points of comparison and no loop samples as were available for 2012-01. Both cruises were recalibrated by adding 0.02. for 2012-13 rhe salinity was very close to the CTD salinity so no correction was applied.

Conclusions

1. The TSG clock appears to have worked well with most casts having very small differences in positions taken from the CTD casts and TSG record when times were matched.

2. The flow rate was noisy but mostly varied between 0.8 and 1.2. There were a few sections in each file where the flow was significantly reduced. A few of these occurred during CTD casts.

3. The temperature in the loop increases by about 0.23Cº based on comparisons during stops for CTD casts. Based on the underway records the increase is an average of 0.22.5Cº or a median of 20.5Cº. This looks reasonable for this time of year. During 2012-13 which preceded this cruise the value was 0.19Cº; the average temperature near the surface was higher during the earlier cruise, so this is consistent.

4. The TSG intake temperature appears to be higher than the CTD temperature by ~0.009Cº when outliers are excluded. During 2012-13 it was found to be high by ~0.005Cº. Given some mismatch in time and depth of sampling between the two sources, recalibration is not justified. The TSG intake temperature is probably a little high is unlikely to be >0.01Cº. It is also possible that the CTD temperature is a little low, but that is less likely.

5. The differences between the TSG Salinity and that of the CTD are very noisy, but the TSG appears to be high by a median of ~0.04 excluding outliers; the standard deviation is 0.02. During 2012-13 the salinity was found to be very close to CTD salinity. Loop samples were available for that cruise but there was some delay in analysis and results were too variable to help. In other recent uses the salinity was though to be low by about 0.2, but there was little information on which to base this. No recalibration is justified, but if the sensors are recalibrated, this can be revisited.

6. The fluorescence data compare well with the SeaPoint fluorometer on the CTD and as well against extracted CHL as a TSG fluorometer ever does.  The raw voltage will be archived as well as the converted values in case a better method is developed for producing reliable data in concentration units.

7. The recent history of the TSG system suggests caution before recalibrating any data. There have been some issues with clogged filters, poor flow and poor performance of the intake temperature. The intake temperature appears to have performed well for this cruise but the flow problems are not fully resolved and the filters got clogged again. How poor flow might affect salinity is unknown. The CTD salinity calibration is thought to be ok but there are some issues concerning storage of samples that make that judgment less confident than usual. The fact that intake temperatures are a little high might suggest that the water samples came from a little higher in the water column, but having salinity also reading too high argues against it. There are too many unknowns to consider recalibration.

f.) Editing

The ATC files were copied to *.EDT.

There were some interruptions in the flow in the loop. These were deliberate and very short, ~2 minutes, for the purpose of cleaning the filter. Examination of the files using CTDEDIT shows little obvious effect on salinity, fluorescence and temperature channels. Fluorescence does peak before and during the stops but it is not clear what is cause and what is effect (high fluorescence leading to clogged filter, or clogged filter leading to high fluorescence). Temperature and salinity are likely affected, but it is not obvious which data points are affected and the effects appear small. So no data were removed due to low flow. The flow rate will be left in the final files in case users want to remove data based on that variable.

The only editing applied was to clean a few single-point spikes in salinity in file #2.

The edited file was copied to *.EDT.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

PAR on: 1-28, 35, 41-63, 79, 112 -160

PAR off: 29-33, 37-40, 65 -78?, 81- 110

1. Pumps did not come on until CTD was at 47m.

12. Pause at 110m due to odd altimeter readings.

40. Event #40 run before #39 due to winch problem.

46. May have hit bottom.

62. Bottle 8 fired at 200m by accident.

65. Pressure on deck -0.4 before, -1.4 after.

72. Touched bottom.

79. UVIC cast, wrong cruise # in header, computer froze at start of cast.

83/84 split cast – 84 is upcast.

102. Went to 95m, then to surface and back down.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 1Apr2011 | Factory |  |  |
| Conductivity | 3396 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2668 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2754 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1396DR | 26Jan2012 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 29Mar2011 | Factory |  |  |
| SeaBird Fluorometer | 2228 |  |  |  |  |
| Eco-AFL Fluorometer | 2214 | ? |  |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Pressure Sensor | 69698 | 15Apr2011 | Factory |  |  |
| Altimeter | 43281 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2012-59

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 26Mar11 | Factory |  |  |
| Conductivity | 2487 | 26Mar11 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 1Feb12 | IOS |  |  |
| Temperature:Secondary | 0603 | 03Mar11 | “ |  |  |