## File List
Number of Files= |  | 8
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2012-59-0030.mrg | 2012-59-0030.sam | 12 | 14 | 2012-59-0030.mrg & 2012-59-0030.sam
2 | 2012-59-0037.mrg | 2012-59-0037.sam | 16 | 19 | 2012-59-0037.mrg & 2012-59-0037.sam
3 | 2012-59-0062.mrg | 2012-59-0062.sam | 21 | 22 | 2012-59-0062.mrg & 2012-59-0062.sam
4 | 2012-59-0067.mrg | 2012-59-0067.sam | 24 | 27 | 2012-59-0067.mrg & 2012-59-0067.sam
5 | 2012-59-0083.mrg | 2012-59-0083.sam | 29 | 34 | 2012-59-0083.mrg & 2012-59-0083.sam
6 | 2012-59-0086.mrg | 2012-59-0086.sam | 36 | 39 | 2012-59-0086.mrg & 2012-59-0086.sam
7 | 2012-59-0102.mrg | 2012-59-0102.sam | 41 | 60 | 2012-59-0102.mrg & 2012-59-0102.sam
8 | 2012-59-0104.mrg | 2012-59-0104.sam | 62 | 66 | 2012-59-0104.mrg & 2012-59-0104.sam

## (Sal0 - Sal_BOT)
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.046727 |  | X | CALC Y
m= | -1.153e-06 | 9.2061e-07 |  | Std Dev of Resid |  | 0.0021847 |  | 500.443 | -0.0040765
b= | -0.0034995 | 0.00099526 |  | Resid Sum of Squares |  | 0.00015274 |  | 2002.51 | -0.0058084
# of Values Extracted = |  |  | 48
# of Values Used in Fit = |  |  | 34
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1000.75 | 34.3925 | 0.00023103 | -0.0046349 | -0.0046533 | -1.8472e-05 |  | 1000.75 | -0.0046349
1 | YES | 751.727 | 34.2641 | 0.00023262 | -0.0080948 | -0.0043662 | 0.0037286 |  | 751.727 | -0.0080948
1 | YES | 500.528 | 34.1188 | 0.00020272 | -0.0049553 | -0.0040766 | 0.00087871 |  | 500.528 | -0.0049553
2 | YES | 1251.35 | 34.4652 | 0.00039884 | -0.0050316 | -0.0049423 | 8.9318e-05 |  | 1251.35 | -0.0050316
2 | YES | 1001.56 | 34.3783 | 0.00022226 | -0.0070763 | -0.0046543 | 0.002422 |  | 1001.56 | -0.0070763
2 | YES | 751.43 | 34.2631 | 0.0002153 | -0.0052452 | -0.0043659 | 0.00087934 |  | 751.43 | -0.0052452
2 | YES | 501.308 | 34.1365 | 0.00021472 | -0.00634 | -0.0040775 | 0.0022625 |  | 501.308 | -0.00634
3 | YES | 751.39 | 34.2669 | 0.00025409 | -0.0055275 | -0.0043658 | 0.0011617 |  | 751.39 | -0.0055275
3 | YES | 502.188 | 34.1011 | 0.00022382 | -0.0018272 | -0.0040785 | -0.0022513 |  | 502.188 | -0.0018272
4 | YES | 1502.71 | 34.5156 | 0.00024087 | -0.0054436 | -0.0052321 | 0.00021148 |  | 1502.71 | -0.0054436
4 | YES | 1251.86 | 34.465 | 0.00021171 | -0.0069961 | -0.0049429 | 0.0020533 |  | 1251.86 | -0.0069961
4 | YES | 1001.84 | 34.3929 | 0.00011736 | -0.002285 | -0.0046546 | -0.0023696 |  | 1001.84 | -0.002285
4 | YES | 750.972 | 34.2919 | 0.00033367 | -0.0005455 | -0.0043653 | -0.0038198 |  | 750.972 | -0.0005455
5 | YES | 2001.2 | 34.5927 | 0.00020148 | -0.0059357 | -0.0058068 | 0.00012883 |  | 2001.2 | -0.0059357
5 | YES | 1501.01 | 34.5027 | 0.00023623 | -0.0030136 | -0.0052301 | -0.0022165 |  | 1501.01 | -0.0030136
5 | YES | 1253.32 | 34.4528 | 0.00026304 | -0.0025635 | -0.0049445 | -0.0023811 |  | 1253.32 | -0.0025635
5 | YES | 1000.51 | 34.3728 | 0.00036008 | -0.0046654 | -0.0046531 | 1.2314e-05 |  | 1000.51 | -0.0046654
5 | YES | 750.946 | 34.2701 | 0.00024109 | -0.0025711 | -0.0043653 | -0.0017942 |  | 750.946 | -0.0025711
5 | YES | 501.478 | 34.0856 | 0.00043872 | -0.0032692 | -0.0040777 | -0.00080848 |  | 501.478 | -0.0032692
6 | YES | 1252.01 | 34.4514 | 0.00027136 | -0.0082664 | -0.004943 | 0.0033234 |  | 1252.01 | -0.0082664
6 | YES | 1002.27 | 34.3665 | 0.00017757 | -0.0088654 | -0.0046551 | 0.0042103 |  | 1002.27 | -0.0088654
6 | YES | 751.596 | 34.2595 | 0.000434 | -0.0023041 | -0.0043661 | -0.002062 |  | 751.596 | -0.0023041
6 | YES | 501.376 | 34.0935 | 0.00030313 | -0.0032654 | -0.0040776 | -0.00081218 |  | 501.376 | -0.0032654
7 | YES | 2002.51 | 34.5736 | 0.00021434 | -0.0065537 | -0.0058084 | 0.0007453 |  | 2002.51 | -0.0065537
7 | YES | 1502 | 34.5079 | 0.00022555 | -0.0015831 | -0.0052313 | -0.0036482 |  | 1502 | -0.0015831
7 | YES | 1250.7 | 34.4478 | 0.00023666 | -0.0045662 | -0.0049415 | -0.00037533 |  | 1250.7 | -0.0045662
7 | YES | 1001.09 | 34.3636 | 0.00021434 | -0.0045242 | -0.0046537 | -0.00012949 |  | 1001.09 | -0.0045242
7 | YES | 752.41 | 34.2535 | 0.00030025 | 0.00067902 | -0.004367 | -0.005046 |  | 752.41 | 0.00067902
7 | YES | 500.443 | 34.0687 | 0.00018568 | -0.0052948 | -0.0040765 | 0.0012183 |  | 500.443 | -0.0052948
7 | NO | 402.603 | 33.961 | 0.00031477 | -0.0017013 | -0.0039637 | -0.0022623 |  | 1503.09 | -0.0054627
7 | NO | 301.557 | 33.8733 | 0.00021117 | -0.0035438 | -0.0038472 | -0.00030332 |  | 1251.75 | -0.0069389
7 | NO | 251.091 | 33.8716 | 0.00032512 | -0.00098419 | -0.003789 | -0.0028048 |  | 1000.78 | -0.0048637
7 | NO | 201.69 | 33.8704 | 0.00021908 | -0.0040016 | -0.003732 | 0.0002696 |  | 751.343 | -0.0049858
7 | NO | 176.343 | 33.7784 | 0.00039839 | 0.0026169 | -0.0037028 | -0.0063197 |  | 501.632 | -0.0054283 | -0.0046543 | 0.0022035
7 | NO | 151.291 | 33.5269 | 0.0033372 | 0.0005188 | -0.0036739 | -0.0041927
7 | NO | 127.105 | 33.1768 | 0.0033041 | 0.012611 | -0.003646 | -0.016257
7 | NO | 101.798 | 32.7449 | 0.0014918 | -0.0041771 | -0.0036168 | 0.00056024
7 | NO | 76.4527 | 32.676 | 0.0010909 | -0.0078888 | -0.0035876 | 0.0043012
7 | NO | 51.1422 | 32.6318 | 0.0048955 | -0.0037117 | -0.0035584 | 0.00015326
7 | NO | 31.3669 | 32.3784 | 0.047587 | -0.0084038 | -0.0035356 | 0.0048681
7 | NO | 22.3599 | 32.2865 | 0.029018 | 0.024521 | -0.0035253 | -0.028046
7 | NO | 11.5943 | 31.9455 | 0.0011439 | -0.0064793 | -0.0035128 | 0.0029664
7 | NO | 5.4989 | 31.9441 | 0.00046435 | -0.0078411 | -0.0035058 | 0.0043353
8 | YES | 1503.09 | 34.5073 | 0.00022292 | -0.0054627 | -0.0052325 | 0.00023012
8 | YES | 1251.75 | 34.4487 | 0.000226 | -0.0069389 | -0.0049427 | 0.0019962
8 | YES | 1000.78 | 34.3733 | 0.0008876 | -0.0048637 | -0.0046534 | 0.00021037
8 | YES | 751.343 | 34.2687 | 0.00027653 | -0.0049858 | -0.0043658 | 0.00062004
8 | YES | 501.632 | 34.0442 | 0.00027649 | -0.0054283 | -0.0040779 | 0.0013505
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (Sal1 - Sal_BOT)
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0033944 |  | X | CALC Y
m= | -3.0018e-07 | 9.0925e-07 |  | Std Dev of Resid |  | 0.0021578 |  | 500.443 | -0.0016342
b= | -0.001484 | 0.00098299 |  | Resid Sum of Squares |  | 0.00014899 |  | 2002.51 | -0.0020851
# of Values Extracted = |  |  | 48
# of Values Used in Fit = |  |  | 34
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1000.75 | 34.3954 | 0.00028053 | -0.0016708 | -0.0017844 | -0.00011352 |  | 1000.75 | -0.0016708
1 | YES | 751.727 | 34.2667 | 0.00029969 | -0.0055275 | -0.0017096 | 0.0038179 |  | 751.727 | -0.0055275
1 | YES | 500.528 | 34.1213 | 0.00030618 | -0.0025139 | -0.0016342 | 0.00087968 |  | 500.528 | -0.0025139
2 | YES | 1251.35 | 34.4685 | 0.00036068 | -0.0017052 | -0.0018596 | -0.00015442 |  | 1251.35 | -0.0017052
2 | YES | 1001.56 | 34.3813 | 0.00032259 | -0.0040817 | -0.0017846 | 0.0022971 |  | 1001.56 | -0.0040817
2 | YES | 751.43 | 34.2657 | 0.00028666 | -0.0026169 | -0.0017095 | 0.00090736 |  | 751.43 | -0.0026169
2 | YES | 501.308 | 34.1391 | 0.00033749 | -0.0037155 | -0.0016344 | 0.0020811 |  | 501.308 | -0.0037155
3 | YES | 751.39 | 34.2693 | 0.00040168 | -0.0030823 | -0.0017095 | 0.0013728 |  | 751.39 | -0.0030823
3 | YES | 502.188 | 34.1031 | 0.0002911 | 0.00018692 | -0.0016347 | -0.0018216 |  | 502.188 | 0.00018692
4 | YES | 1502.71 | 34.5185 | 0.00031663 | -0.0025291 | -0.001935 | 0.0005941 |  | 1502.71 | -0.0025291
4 | YES | 1251.86 | 34.4675 | 0.00026751 | -0.0045204 | -0.0018597 | 0.0026607 |  | 1251.86 | -0.0045204
4 | YES | 1001.84 | 34.3954 | 0.00030039 | 0.00017166 | -0.0017847 | -0.0019564 |  | 1001.84 | 0.00017166
4 | YES | 750.972 | 34.2946 | 0.0005572 | 0.0021744 | -0.0017094 | -0.0038838 |  | 750.972 | 0.0021744
5 | YES | 2001.2 | 34.5965 | 0.00035946 | -0.0021172 | -0.0020847 | 3.2485e-05 |  | 2001.2 | -0.0021172
5 | YES | 1501.01 | 34.5057 | 0.0002724 | 1.1444e-05 | -0.0019345 | -0.001946 |  | 1501.01 | 1.1444e-05
5 | YES | 1253.32 | 34.4558 | 0.00033627 | 0.00044632 | -0.0018602 | -0.0023065 |  | 1253.32 | 0.00044632
5 | YES | 1000.51 | 34.3755 | 0.00056957 | -0.0020142 | -0.0017843 | 0.00022987 |  | 1000.51 | -0.0020142
5 | YES | 750.946 | 34.2727 | 0.00027779 | -1.9073e-05 | -0.0017094 | -0.0016903 |  | 750.946 | -1.9073e-05
5 | YES | 501.478 | 34.088 | 0.00041069 | -0.00085831 | -0.0016345 | -0.00077619 |  | 501.478 | -0.00085831
6 | YES | 1252.01 | 34.4546 | 0.00040448 | -0.0050888 | -0.0018598 | 0.003229 |  | 1252.01 | -0.0050888
6 | YES | 1002.27 | 34.3694 | 0.00032593 | -0.0059776 | -0.0017848 | 0.0041928 |  | 1002.27 | -0.0059776
6 | YES | 751.596 | 34.2621 | 0.00052682 | 0.00027084 | -0.0017096 | -0.0019804 |  | 751.596 | 0.00027084
6 | YES | 501.376 | 34.096 | 0.00042877 | -0.00078583 | -0.0016345 | -0.00084863 |  | 501.376 | -0.00078583
7 | YES | 2002.51 | 34.5777 | 0.00030487 | -0.0025444 | -0.0020851 | 0.00045933 |  | 2002.51 | -0.0025444
7 | YES | 1502 | 34.5114 | 0.00030696 | 0.0018616 | -0.0019348 | -0.0037964 |  | 1502 | 0.0018616
7 | YES | 1250.7 | 34.4509 | 0.00034966 | -0.0015068 | -0.0018594 | -0.00035258 |  | 1250.7 | -0.0015068
7 | YES | 1001.09 | 34.3667 | 0.00031889 | -0.0014229 | -0.0017845 | -0.00036158 |  | 1001.09 | -0.0014229
7 | YES | 752.41 | 34.2561 | 0.00033789 | 0.003334 | -0.0017098 | -0.0050439 |  | 752.41 | 0.003334
7 | YES | 500.443 | 34.0716 | 0.00042574 | -0.0024223 | -0.0016342 | 0.00078815 |  | 500.443 | -0.0024223
7 | NO | 402.603 | 33.9636 | 0.00034916 | 0.00085068 | -0.0016048 | -0.0024555 |  | 1503.09 | -0.001873
7 | NO | 301.557 | 33.8758 | 0.00033755 | -0.0010109 | -0.0015745 | -0.00056359 |  | 1251.75 | -0.0038795
7 | NO | 251.091 | 33.8742 | 0.00039369 | 0.0015869 | -0.0015593 | -0.0031462 |  | 1000.78 | -0.0018959
7 | NO | 201.69 | 33.873 | 0.00036811 | -0.0014267 | -0.0015445 | -0.0001178 |  | 751.343 | -0.0020218
7 | NO | 176.343 | 33.7807 | 0.0007215 | 0.0048714 | -0.0015369 | -0.0064083 |  | 501.632 | -0.0027428 | -0.0017846 | 0.0021285
7 | NO | 151.291 | 33.5278 | 0.0030761 | 0.0013618 | -0.0015294 | -0.0028912
7 | NO | 127.105 | 33.18 | 0.0040478 | 0.015778 | -0.0015221 | -0.0173
7 | NO | 101.798 | 32.7496 | 0.00071428 | 0.00049591 | -0.0015145 | -0.0020104
7 | NO | 76.4527 | 32.6785 | 0.00057357 | -0.0054207 | -0.0015069 | 0.0039138
7 | NO | 51.1422 | 32.6355 | 0.0024449 | -4.9591e-05 | -0.0014993 | -0.0014497
7 | NO | 31.3669 | 32.4088 | 0.0037977 | 0.021996 | -0.0014934 | -0.023489
7 | NO | 22.3599 | 32.2818 | 0.0082293 | 0.01984 | -0.0014907 | -0.021331
7 | NO | 11.5943 | 31.9471 | 0.00078157 | -0.0049191 | -0.0014874 | 0.0034316
7 | NO | 5.4989 | 31.9463 | 0.00028327 | -0.0055523 | -0.0014856 | 0.0040667
8 | YES | 1503.09 | 34.5109 | 0.00033733 | -0.001873 | -0.0019352 | -6.2134e-05
8 | YES | 1251.75 | 34.4517 | 0.000357 | -0.0038795 | -0.0018597 | 0.0020198
8 | YES | 1000.78 | 34.3763 | 0.00099448 | -0.0018959 | -0.0017844 | 0.00011153
8 | YES | 751.343 | 34.2717 | 0.00039048 | -0.0020218 | -0.0017095 | 0.00031229
8 | YES | 501.632 | 34.0469 | 0.00041638 | -0.0027428 | -0.0016345 | 0.0011082
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (Sal0 - Sal_BOT) (2)
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_BOT)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_BOT), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.011597 |  | X | CALC Y
m= | 0.00010486 | 0.00017114 |  | Std Dev of Resid |  | 0.0022246 |  | 1 | -0.0050706
b= | -0.0051755 | 0.00093231 |  | Resid Sum of Squares |  | 0.00015837 |  | 8 | -0.0043366
# of Values Extracted = |  |  | 48
# of Values Used in Fit = |  |  | 34
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_BOT) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1000.75 | 34.3925 | 0.00023103 | -0.0046349 | -0.0050706 | -0.00043578 |  | 1 | -0.0046349
1 | YES | 751.727 | 34.2641 | 0.00023262 | -0.0080948 | -0.0050706 | 0.0030241 |  | 1 | -0.0080948
1 | YES | 500.528 | 34.1188 | 0.00020272 | -0.0049553 | -0.0050706 | -0.00011535 |  | 1 | -0.0049553
2 | YES | 1251.35 | 34.4652 | 0.00039884 | -0.0050316 | -0.0049658 | 6.5814e-05 |  | 2 | -0.0050316
2 | YES | 1001.56 | 34.3783 | 0.00022226 | -0.0070763 | -0.0049658 | 0.0021105 |  | 2 | -0.0070763
2 | YES | 751.43 | 34.2631 | 0.0002153 | -0.0052452 | -0.0049658 | 0.00027943 |  | 2 | -0.0052452
2 | YES | 501.308 | 34.1365 | 0.00021472 | -0.00634 | -0.0049658 | 0.0013743 |  | 2 | -0.00634
3 | YES | 751.39 | 34.2669 | 0.00025409 | -0.0055275 | -0.0048609 | 0.00066659 |  | 3 | -0.0055275
3 | YES | 502.188 | 34.1011 | 0.00022382 | -0.0018272 | -0.0048609 | -0.0030337 |  | 3 | -0.0018272
4 | YES | 1502.71 | 34.5156 | 0.00024087 | -0.0054436 | -0.004756 | 0.00068752 |  | 4 | -0.0054436
4 | YES | 1251.86 | 34.465 | 0.00021171 | -0.0069961 | -0.004756 | 0.0022401 |  | 4 | -0.0069961
4 | YES | 1001.84 | 34.3929 | 0.00011736 | -0.002285 | -0.004756 | -0.002471 |  | 4 | -0.002285
4 | YES | 750.972 | 34.2919 | 0.00033367 | -0.0005455 | -0.004756 | -0.0042105 |  | 4 | -0.0005455
5 | YES | 2001.2 | 34.5927 | 0.00020148 | -0.0059357 | -0.0046512 | 0.0012845 |  | 5 | -0.0059357
5 | YES | 1501.01 | 34.5027 | 0.00023623 | -0.0030136 | -0.0046512 | -0.0016376 |  | 5 | -0.0030136
5 | YES | 1253.32 | 34.4528 | 0.00026304 | -0.0025635 | -0.0046512 | -0.0020877 |  | 5 | -0.0025635
5 | YES | 1000.51 | 34.3728 | 0.00036008 | -0.0046654 | -0.0046512 | 1.4187e-05 |  | 5 | -0.0046654
5 | YES | 750.946 | 34.2701 | 0.00024109 | -0.0025711 | -0.0046512 | -0.0020801 |  | 5 | -0.0025711
5 | YES | 501.478 | 34.0856 | 0.00043872 | -0.0032692 | -0.0046512 | -0.001382 |  | 5 | -0.0032692
6 | YES | 1252.01 | 34.4514 | 0.00027136 | -0.0082664 | -0.0045463 | 0.0037201 |  | 6 | -0.0082664
6 | YES | 1002.27 | 34.3665 | 0.00017757 | -0.0088654 | -0.0045463 | 0.004319 |  | 6 | -0.0088654
6 | YES | 751.596 | 34.2595 | 0.000434 | -0.0023041 | -0.0045463 | -0.0022422 |  | 6 | -0.0023041
6 | YES | 501.376 | 34.0935 | 0.00030313 | -0.0032654 | -0.0045463 | -0.0012809 |  | 6 | -0.0032654
7 | YES | 2002.51 | 34.5736 | 0.00021434 | -0.0065537 | -0.0044415 | 0.0021122 |  | 7 | -0.0065537
7 | YES | 1502 | 34.5079 | 0.00022555 | -0.0015831 | -0.0044415 | -0.0028584 |  | 7 | -0.0015831
7 | YES | 1250.7 | 34.4478 | 0.00023666 | -0.0045662 | -0.0044415 | 0.00012474 |  | 7 | -0.0045662
7 | YES | 1001.09 | 34.3636 | 0.00021434 | -0.0045242 | -0.0044415 | 8.2775e-05 |  | 7 | -0.0045242
7 | YES | 752.41 | 34.2535 | 0.00030025 | 0.00067902 | -0.0044415 | -0.0051205 |  | 7 | 0.00067902
7 | YES | 500.443 | 34.0687 | 0.00018568 | -0.0052948 | -0.0044415 | 0.00085335 |  | 7 | -0.0052948
7 | NO | 402.603 | 33.961 | 0.00031477 | -0.0017013 | -0.0044415 | -0.0027401 |  | 8 | -0.0054627
7 | NO | 301.557 | 33.8733 | 0.00021117 | -0.0035438 | -0.0044415 | -0.0008976 |  | 8 | -0.0069389
7 | NO | 251.091 | 33.8716 | 0.00032512 | -0.00098419 | -0.0044415 | -0.0034573 |  | 8 | -0.0048637
7 | NO | 201.69 | 33.8704 | 0.00021908 | -0.0040016 | -0.0044415 | -0.00043983 |  | 8 | -0.0049858
7 | NO | 176.343 | 33.7784 | 0.00039839 | 0.0026169 | -0.0044415 | -0.0070583 |  | 8 | -0.0054283
7 | NO | 151.291 | 33.5269 | 0.0033372 | 0.0005188 | -0.0044415 | -0.0049603
7 | NO | 127.105 | 33.1768 | 0.0033041 | 0.012611 | -0.0044415 | -0.017053
7 | NO | 101.798 | 32.7449 | 0.0014918 | -0.0041771 | -0.0044415 | -0.00026436
7 | NO | 76.4527 | 32.676 | 0.0010909 | -0.0078888 | -0.0044415 | 0.0034473
7 | NO | 51.1422 | 32.6318 | 0.0048955 | -0.0037117 | -0.0044415 | -0.00072975
7 | NO | 31.3669 | 32.3784 | 0.047587 | -0.0084038 | -0.0044415 | 0.0039623
7 | NO | 22.3599 | 32.2865 | 0.029018 | 0.024521 | -0.0044415 | -0.028962
7 | NO | 11.5943 | 31.9455 | 0.0011439 | -0.0064793 | -0.0044415 | 0.0020378
7 | NO | 5.4989 | 31.9441 | 0.00046435 | -0.0078411 | -0.0044415 | 0.0033997
8 | YES | 1503.09 | 34.5073 | 0.00022292 | -0.0054627 | -0.0043366 | 0.0011261
8 | YES | 1251.75 | 34.4487 | 0.000226 | -0.0069389 | -0.0043366 | 0.0026023
8 | YES | 1000.78 | 34.3733 | 0.0008876 | -0.0048637 | -0.0043366 | 0.00052715
8 | YES | 751.343 | 34.2687 | 0.00027653 | -0.0049858 | -0.0043366 | 0.00064922
8 | YES | 501.632 | 34.0442 | 0.00027649 | -0.0054283 | -0.0043366 | 0.0010917
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (Sal1 - Sal_BOT) (2)
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_BOT)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_BOT), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.034905 |  | X | CALC Y
m= | 0.00017574 | 0.00016335 |  | Std Dev of Resid |  | 0.0021234 |  | 1 | -0.0024824
b= | -0.0026581 | 0.00088989 |  | Resid Sum of Squares |  | 0.00014428 |  | 8 | -0.0012522
# of Values Extracted = |  |  | 48
# of Values Used in Fit = |  |  | 34
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_BOT) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 1000.75 | 34.3954 | 0.00028053 | -0.0016708 | -0.0024824 | -0.00081154 |  | 1 | -0.0016708
1 | YES | 751.727 | 34.2667 | 0.00029969 | -0.0055275 | -0.0024824 | 0.0030451 |  | 1 | -0.0055275
1 | YES | 500.528 | 34.1213 | 0.00030618 | -0.0025139 | -0.0024824 | 3.1506e-05 |  | 1 | -0.0025139
2 | YES | 1251.35 | 34.4685 | 0.00036068 | -0.0017052 | -0.0023066 | -0.00060148 |  | 2 | -0.0017052
2 | YES | 1001.56 | 34.3813 | 0.00032259 | -0.0040817 | -0.0023066 | 0.0017751 |  | 2 | -0.0040817
2 | YES | 751.43 | 34.2657 | 0.00028666 | -0.0026169 | -0.0023066 | 0.00031023 |  | 2 | -0.0026169
2 | YES | 501.308 | 34.1391 | 0.00033749 | -0.0037155 | -0.0023066 | 0.0014089 |  | 2 | -0.0037155
3 | YES | 751.39 | 34.2693 | 0.00040168 | -0.0030823 | -0.0021309 | 0.00095137 |  | 3 | -0.0030823
3 | YES | 502.188 | 34.1031 | 0.0002911 | 0.00018692 | -0.0021309 | -0.0023178 |  | 3 | 0.00018692
4 | YES | 1502.71 | 34.5185 | 0.00031663 | -0.0025291 | -0.0019552 | 0.00057397 |  | 4 | -0.0025291
4 | YES | 1251.86 | 34.4675 | 0.00026751 | -0.0045204 | -0.0019552 | 0.0025652 |  | 4 | -0.0045204
4 | YES | 1001.84 | 34.3954 | 0.00030039 | 0.00017166 | -0.0019552 | -0.0021268 |  | 4 | 0.00017166
4 | YES | 750.972 | 34.2946 | 0.0005572 | 0.0021744 | -0.0019552 | -0.0041296 |  | 4 | 0.0021744
5 | YES | 2001.2 | 34.5965 | 0.00035946 | -0.0021172 | -0.0017794 | 0.00033772 |  | 5 | -0.0021172
5 | YES | 1501.01 | 34.5057 | 0.0002724 | 1.1444e-05 | -0.0017794 | -0.0017909 |  | 5 | 1.1444e-05
5 | YES | 1253.32 | 34.4558 | 0.00033627 | 0.00044632 | -0.0017794 | -0.0022258 |  | 5 | 0.00044632
5 | YES | 1000.51 | 34.3755 | 0.00056957 | -0.0020142 | -0.0017794 | 0.00023472 |  | 5 | -0.0020142
5 | YES | 750.946 | 34.2727 | 0.00027779 | -1.9073e-05 | -0.0017794 | -0.0017604 |  | 5 | -1.9073e-05
5 | YES | 501.478 | 34.088 | 0.00041069 | -0.00085831 | -0.0017794 | -0.00092113 |  | 5 | -0.00085831
6 | YES | 1252.01 | 34.4546 | 0.00040448 | -0.0050888 | -0.0016037 | 0.0034851 |  | 6 | -0.0050888
6 | YES | 1002.27 | 34.3694 | 0.00032593 | -0.0059776 | -0.0016037 | 0.0043739 |  | 6 | -0.0059776
6 | YES | 751.596 | 34.2621 | 0.00052682 | 0.00027084 | -0.0016037 | -0.0018745 |  | 6 | 0.00027084
6 | YES | 501.376 | 34.096 | 0.00042877 | -0.00078583 | -0.0016037 | -0.00081787 |  | 6 | -0.00078583
7 | YES | 2002.51 | 34.5777 | 0.00030487 | -0.0025444 | -0.001428 | 0.0011164 |  | 7 | -0.0025444
7 | YES | 1502 | 34.5114 | 0.00030696 | 0.0018616 | -0.001428 | -0.0032895 |  | 7 | 0.0018616
7 | YES | 1250.7 | 34.4509 | 0.00034966 | -0.0015068 | -0.001428 | 7.8846e-05 |  | 7 | -0.0015068
7 | YES | 1001.09 | 34.3667 | 0.00031889 | -0.0014229 | -0.001428 | -5.0837e-06 |  | 7 | -0.0014229
7 | YES | 752.41 | 34.2561 | 0.00033789 | 0.003334 | -0.001428 | -0.004762 |  | 7 | 0.003334
7 | YES | 500.443 | 34.0716 | 0.00042574 | -0.0024223 | -0.001428 | 0.00099437 |  | 7 | -0.0024223
7 | NO | 402.603 | 33.9636 | 0.00034916 | 0.00085068 | -0.001428 | -0.0022786 |  | 8 | -0.001873
7 | NO | 301.557 | 33.8758 | 0.00033755 | -0.0010109 | -0.001428 | -0.00041707 |  | 8 | -0.0038795
7 | NO | 251.091 | 33.8742 | 0.00039369 | 0.0015869 | -0.001428 | -0.0030149 |  | 8 | -0.0018959
7 | NO | 201.69 | 33.873 | 0.00036811 | -0.0014267 | -0.001428 | -1.2637e-06 |  | 8 | -0.0020218
7 | NO | 176.343 | 33.7807 | 0.0007215 | 0.0048714 | -0.001428 | -0.0062993 |  | 8 | -0.0027428
7 | NO | 151.291 | 33.5278 | 0.0030761 | 0.0013618 | -0.001428 | -0.0027898
7 | NO | 127.105 | 33.18 | 0.0040478 | 0.015778 | -0.001428 | -0.017206
7 | NO | 101.798 | 32.7496 | 0.00071428 | 0.00049591 | -0.001428 | -0.0019239
7 | NO | 76.4527 | 32.6785 | 0.00057357 | -0.0054207 | -0.001428 | 0.0039927
7 | NO | 51.1422 | 32.6355 | 0.0024449 | -4.9591e-05 | -0.001428 | -0.0013784
7 | NO | 31.3669 | 32.4088 | 0.0037977 | 0.021996 | -0.001428 | -0.023423
7 | NO | 22.3599 | 32.2818 | 0.0082293 | 0.01984 | -0.001428 | -0.021268
7 | NO | 11.5943 | 31.9471 | 0.00078157 | -0.0049191 | -0.001428 | 0.0034911
7 | NO | 5.4989 | 31.9463 | 0.00028327 | -0.0055523 | -0.001428 | 0.0041243
8 | YES | 1503.09 | 34.5109 | 0.00033733 | -0.001873 | -0.0012522 | 0.00062079
8 | YES | 1251.75 | 34.4517 | 0.000357 | -0.0038795 | -0.0012522 | 0.0026273
8 | YES | 1000.78 | 34.3763 | 0.00099448 | -0.0018959 | -0.0012522 | 0.00064367
8 | YES | 751.343 | 34.2717 | 0.00039048 | -0.0020218 | -0.0012522 | 0.00076956
8 | YES | 501.632 | 34.0469 | 0.00041638 | -0.0027428 | -0.0012522 | 0.0014905
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (Sal0 - Sal_BOT) (3)
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_BOT), Prop=Sal0
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.085829 |  | X | CALC Y
m= | 0.0012738 | 0.0013146 |  | Std Dev of Resid |  | 0.0030252 |  | 31.9441 | -0.0056753
b= | -0.046365 | 0.044651 |  | Resid Sum of Squares |  | 9.1516e-05 |  | 34.5736 | -0.002326
# of Values Extracted = |  |  | 20
# of Values Used in Fit = |  |  | 12
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_BOT) |  |  |  | Fit vs. | Sal0
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
7 | YES | 2002.51 | 34.5736 | 0.00021434 | -0.0065537 | -0.002326 | 0.0042277 |  | 34.5736 | -0.0065537
7 | YES | 1502 | 34.5079 | 0.00022555 | -0.0015831 | -0.0024096 | -0.00082654 |  | 34.5079 | -0.0015831
7 | YES | 1250.7 | 34.4478 | 0.00023666 | -0.0045662 | -0.0024862 | 0.00208 |  | 34.4478 | -0.0045662
7 | YES | 1001.09 | 34.3636 | 0.00021434 | -0.0045242 | -0.0025934 | 0.0019308 |  | 34.3636 | -0.0045242
7 | YES | 752.41 | 34.2535 | 0.00030025 | 0.00067902 | -0.0027337 | -0.0034127 |  | 34.2535 | 0.00067902
7 | YES | 500.443 | 34.0687 | 0.00018568 | -0.0052948 | -0.0029691 | 0.0023257 |  | 34.0687 | -0.0052948
7 | YES | 402.603 | 33.961 | 0.00031477 | -0.0017013 | -0.0031063 | -0.0014049 |  | 33.961 | -0.0017013
7 | YES | 301.557 | 33.8733 | 0.00021117 | -0.0035438 | -0.003218 | 0.00032588 |  | 33.8733 | -0.0035438
7 | YES | 251.091 | 33.8716 | 0.00032512 | -0.00098419 | -0.0032201 | -0.0022359 |  | 33.8716 | -0.00098419
7 | YES | 201.69 | 33.8704 | 0.00021908 | -0.0040016 | -0.0032217 | 0.00077995 |  | 33.8704 | -0.0040016
7 | YES | 176.343 | 33.7784 | 0.00039839 | 0.0026169 | -0.0033389 | -0.0059557 |  | 33.7784 | 0.0026169
7 | NO | 151.291 | 33.5269 | 0.0033372 | 0.0005188 | -0.0036592 | -0.004178 |  | 31.9441 | -0.0078411
7 | NO | 127.105 | 33.1768 | 0.0033041 | 0.012611 | -0.0041052 | -0.016717
7 | NO | 101.798 | 32.7449 | 0.0014918 | -0.0041771 | -0.0046553 | -0.0004782
7 | NO | 76.4527 | 32.676 | 0.0010909 | -0.0078888 | -0.0047431 | 0.0031457
7 | NO | 51.1422 | 32.6318 | 0.0048955 | -0.0037117 | -0.0047994 | -0.0010877
7 | NO | 31.3669 | 32.3784 | 0.047587 | -0.0084038 | -0.0051221 | 0.0032817
7 | NO | 22.3599 | 32.2865 | 0.029018 | 0.024521 | -0.0052392 | -0.02976
7 | NO | 11.5943 | 31.9455 | 0.0011439 | -0.0064793 | -0.0056735 | 0.00080572
7 | YES | 5.4989 | 31.9441 | 0.00046435 | -0.0078411 | -0.0056753 | 0.0021658
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (Sal1 - Sal_BOT) (3)
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_BOT), Prop=Sal1
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.31062 |  | X | CALC Y
m= | 0.0019361 | 0.00079996 |  | Std Dev of Resid |  | 0.0026681 |  | 31.9463 | -0.0042025
b= | -0.066053 | 0.026937 |  | Resid Sum of Squares |  | 9.2543e-05 |  | 34.5777 | 0.00089204
# of Values Extracted = |  |  | 20
# of Values Used in Fit = |  |  | 15
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_BOT) |  |  |  | Fit vs. | Sal1
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
7 | YES | 2002.51 | 34.5777 | 0.00030487 | -0.0025444 | 0.00089204 | 0.0034364 |  | 34.5777 | -0.0025444
7 | YES | 1502 | 34.5114 | 0.00030696 | 0.0018616 | 0.00076368 | -0.0010979 |  | 34.5114 | 0.0018616
7 | YES | 1250.7 | 34.4509 | 0.00034966 | -0.0015068 | 0.00064654 | 0.0021534 |  | 34.4509 | -0.0015068
7 | YES | 1001.09 | 34.3667 | 0.00031889 | -0.0014229 | 0.00048353 | 0.0019064 |  | 34.3667 | -0.0014229
7 | YES | 752.41 | 34.2561 | 0.00033789 | 0.003334 | 0.0002694 | -0.0030647 |  | 34.2561 | 0.003334
7 | YES | 500.443 | 34.0716 | 0.00042574 | -0.0024223 | -8.7809e-05 | 0.0023345 |  | 34.0716 | -0.0024223
7 | YES | 402.603 | 33.9636 | 0.00034916 | 0.00085068 | -0.00029691 | -0.0011476 |  | 33.9636 | 0.00085068
7 | YES | 301.557 | 33.8758 | 0.00033755 | -0.0010109 | -0.00046689 | 0.000544 |  | 33.8758 | -0.0010109
7 | YES | 251.091 | 33.8742 | 0.00039369 | 0.0015869 | -0.00046999 | -0.0020569 |  | 33.8742 | 0.0015869
7 | YES | 201.69 | 33.873 | 0.00036811 | -0.0014267 | -0.00047231 | 0.00095439 |  | 33.873 | -0.0014267
7 | YES | 176.343 | 33.7807 | 0.0007215 | 0.0048714 | -0.00065101 | -0.0055224 |  | 33.7807 | 0.0048714
7 | NO | 151.291 | 33.5278 | 0.0030761 | 0.0013618 | -0.0011406 | -0.0025025 |  | 32.7496 | 0.00049591
7 | NO | 127.105 | 33.18 | 0.0040478 | 0.015778 | -0.001814 | -0.017592 |  | 32.6785 | -0.0054207
7 | YES | 101.798 | 32.7496 | 0.00071428 | 0.00049591 | -0.0026473 | -0.0031432 |  | 31.9471 | -0.0049191
7 | YES | 76.4527 | 32.6785 | 0.00057357 | -0.0054207 | -0.002785 | 0.0026357 |  | 31.9463 | -0.0055523
7 | NO | 51.1422 | 32.6355 | 0.0024449 | -4.9591e-05 | -0.0028682 | -0.0028186
7 | NO | 31.3669 | 32.4088 | 0.0037977 | 0.021996 | -0.0033071 | -0.025303
7 | NO | 22.3599 | 32.2818 | 0.0082293 | 0.01984 | -0.003553 | -0.023393
7 | YES | 11.5943 | 31.9471 | 0.00078157 | -0.0049191 | -0.004201 | 0.00071805
7 | YES | 5.4989 | 31.9463 | 0.00028327 | -0.0055523 | -0.0042025 | 0.0013497
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
NO
NO
NO
YES
YES
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.3 | Run at 2013/02/12 11:49:39
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_BOT | Sal0 |  | Sal0 - Sal_BOT
Comp 2 | Sal_BOT | Sal1 |  | Sal1 - Sal_BOT
Average Size: | 0.4
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_BOT | Sal0 | Sal0 | (S-P) | Sal_BOT | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2012-59-0030.mrg | 2012-59-0030.sam | 121 | 1000.75 | 1000.75 | 0.088559 | -0.00042725 | 34.3971 | 34.3925 | 0.00023103 | -0.0046349 | 34.3971 | 34.3954 | 0.00028053 | -0.0016708
2012-59-0030.mrg | 2012-59-0030.sam | 121 | 751.727 | 751.727 | 0.15391 | 0.00024414 | 34.2722 | 34.2641 | 0.00023262 | -0.0080948 | 34.2722 | 34.2667 | 0.00029969 | -0.0055275
2012-59-0030.mrg | 2012-59-0030.sam | 121 | 500.528 | 500.528 | 0.16694 | -0.00021362 | 34.1238 | 34.1188 | 0.00020272 | -0.0049553 | 34.1238 | 34.1213 | 0.00030618 | -0.0025139
2012-59-0037.mrg | 2012-59-0037.sam | 53 | 1251.43 | 1251.35 | 0.25423 | -0.084351 | 34.4702 | 34.4652 | 0.00039884 | -0.0050316 | 34.4702 | 34.4685 | 0.00036068 | -0.0017052
2012-59-0037.mrg | 2012-59-0037.sam | 102 | 1001.51 | 1001.56 | 0.24477 | 0.054443 | 34.3854 | 34.3783 | 0.00022226 | -0.0070763 | 34.3854 | 34.3813 | 0.00032259 | -0.0040817
2012-59-0037.mrg | 2012-59-0037.sam | 121 | 751.43 | 751.43 | 0.15689 | 0.00042725 | 34.2683 | 34.2631 | 0.0002153 | -0.0052452 | 34.2683 | 34.2657 | 0.00028666 | -0.0026169
2012-59-0037.mrg | 2012-59-0037.sam | 121 | 501.308 | 501.308 | 0.12787 | -0.00030518 | 34.1428 | 34.1365 | 0.00021472 | -0.00634 | 34.1428 | 34.1391 | 0.00033749 | -0.0037155
2012-59-0062.mrg | 2012-59-0062.sam | 86 | 751.314 | 751.39 | 0.26451 | 0.075867 | 34.2724 | 34.2669 | 0.00025409 | -0.0055275 | 34.2724 | 34.2693 | 0.00040168 | -0.0030823
2012-59-0062.mrg | 2012-59-0062.sam | 62 | 502.058 | 502.188 | 0.24304 | 0.12982 | 34.1029 | 34.1011 | 0.00022382 | -0.0018272 | 34.1029 | 34.1031 | 0.0002911 | 0.00018692
2012-59-0067.mrg | 2012-59-0067.sam | 118 | 1502.72 | 1502.71 | 0.26166 | -0.010864 | 34.521 | 34.5156 | 0.00024087 | -0.0054436 | 34.521 | 34.5185 | 0.00031663 | -0.0025291
2012-59-0067.mrg | 2012-59-0067.sam | 121 | 1251.86 | 1251.86 | 0.23226 | 0.00024414 | 34.472 | 34.465 | 0.00021171 | -0.0069961 | 34.472 | 34.4675 | 0.00026751 | -0.0045204
2012-59-0067.mrg | 2012-59-0067.sam | 14 | 1001.83 | 1001.84 | 0.24294 | 0.010925 | 34.3952 | 34.3929 | 0.00011736 | -0.002285 | 34.3952 | 34.3954 | 0.00030039 | 0.00017166
2012-59-0067.mrg | 2012-59-0067.sam | 15 | 751 | 750.972 | 0.2425 | -0.028442 | 34.2924 | 34.2919 | 0.00033367 | -0.0005455 | 34.2924 | 34.2946 | 0.0005572 | 0.0021744
2012-59-0083.mrg | 2012-59-0083.sam | 46 | 2001.27 | 2001.2 | 0.22505 | -0.067993 | 34.5986 | 34.5927 | 0.00020148 | -0.0059357 | 34.5986 | 34.5965 | 0.00035946 | -0.0021172
2012-59-0083.mrg | 2012-59-0083.sam | 43 | 1500.94 | 1501.01 | 0.24169 | 0.074097 | 34.5057 | 34.5027 | 0.00023623 | -0.0030136 | 34.5057 | 34.5057 | 0.0002724 | 1.1444e-05
2012-59-0083.mrg | 2012-59-0083.sam | 55 | 1253.28 | 1253.32 | 0.24321 | 0.039429 | 34.4554 | 34.4528 | 0.00026304 | -0.0025635 | 34.4554 | 34.4558 | 0.00033627 | 0.00044632
2012-59-0083.mrg | 2012-59-0083.sam | 53 | 1000.54 | 1000.51 | 0.23683 | -0.039307 | 34.3775 | 34.3728 | 0.00036008 | -0.0046654 | 34.3775 | 34.3755 | 0.00056957 | -0.0020142
2012-59-0083.mrg | 2012-59-0083.sam | 79 | 750.928 | 750.946 | 0.17198 | 0.017944 | 34.2727 | 34.2701 | 0.00024109 | -0.0025711 | 34.2727 | 34.2727 | 0.00027779 | -1.9073e-05
2012-59-0083.mrg | 2012-59-0083.sam | 50 | 501.443 | 501.478 | 0.24625 | 0.034821 | 34.0889 | 34.0856 | 0.00043872 | -0.0032692 | 34.0889 | 34.088 | 0.00041069 | -0.00085831
2012-59-0086.mrg | 2012-59-0086.sam | 79 | 1251.97 | 1252.01 | 0.23528 | 0.044434 | 34.4597 | 34.4514 | 0.00027136 | -0.0082664 | 34.4597 | 34.4546 | 0.00040448 | -0.0050888
2012-59-0086.mrg | 2012-59-0086.sam | 31 | 1002.32 | 1002.27 | 0.2345 | -0.052429 | 34.3754 | 34.3665 | 0.00017757 | -0.0088654 | 34.3754 | 34.3694 | 0.00032593 | -0.0059776
2012-59-0086.mrg | 2012-59-0086.sam | 76 | 751.524 | 751.596 | 0.24741 | 0.071899 | 34.2618 | 34.2595 | 0.000434 | -0.0023041 | 34.2618 | 34.2621 | 0.00052682 | 0.00027084
2012-59-0086.mrg | 2012-59-0086.sam | 118 | 501.387 | 501.376 | 0.24968 | -0.010773 | 34.0968 | 34.0935 | 0.00030313 | -0.0032654 | 34.0968 | 34.096 | 0.00042877 | -0.00078583
2012-59-0102.mrg | 2012-59-0102.sam | 109 | 2002.55 | 2002.51 | 0.28519 | -0.040161 | 34.5802 | 34.5736 | 0.00021434 | -0.0065537 | 34.5802 | 34.5777 | 0.00030487 | -0.0025444
2012-59-0102.mrg | 2012-59-0102.sam | 84 | 1502.15 | 1502 | 0.23134 | -0.15369 | 34.5095 | 34.5079 | 0.00022555 | -0.0015831 | 34.5095 | 34.5114 | 0.00030696 | 0.0018616
2012-59-0102.mrg | 2012-59-0102.sam | 57 | 1250.8 | 1250.7 | 0.24337 | -0.10156 | 34.4524 | 34.4478 | 0.00023666 | -0.0045662 | 34.4524 | 34.4509 | 0.00034966 | -0.0015068
2012-59-0102.mrg | 2012-59-0102.sam | 91 | 1001.07 | 1001.09 | 0.15144 | 0.020386 | 34.3681 | 34.3636 | 0.00021434 | -0.0045242 | 34.3681 | 34.3667 | 0.00031889 | -0.0014229
2012-59-0102.mrg | 2012-59-0102.sam | 43 | 752.456 | 752.41 | 0.25455 | -0.045654 | 34.2528 | 34.2535 | 0.00030025 | 0.00067902 | 34.2528 | 34.2561 | 0.00033789 | 0.003334
2012-59-0102.mrg | 2012-59-0102.sam | 39 | 500.475 | 500.443 | 0.22546 | -0.032257 | 34.074 | 34.0687 | 0.00018568 | -0.0052948 | 34.074 | 34.0716 | 0.00042574 | -0.0024223
2012-59-0102.mrg | 2012-59-0102.sam | 60 | 402.604 | 402.603 | 0.25084 | -0.00054932 | 33.9627 | 33.961 | 0.00031477 | -0.0017013 | 33.9627 | 33.9636 | 0.00034916 | 0.00085068
2012-59-0102.mrg | 2012-59-0102.sam | 111 | 301.603 | 301.557 | 0.16935 | -0.045715 | 33.8768 | 33.8733 | 0.00021117 | -0.0035438 | 33.8768 | 33.8758 | 0.00033755 | -0.0010109
2012-59-0102.mrg | 2012-59-0102.sam | 121 | 251.091 | 251.091 | 0.14058 | -0.00021362 | 33.8726 | 33.8716 | 0.00032512 | -0.00098419 | 33.8726 | 33.8742 | 0.00039369 | 0.0015869
2012-59-0102.mrg | 2012-59-0102.sam | 82 | 201.764 | 201.69 | 0.24282 | -0.074188 | 33.8744 | 33.8704 | 0.00021908 | -0.0040016 | 33.8744 | 33.873 | 0.00036811 | -0.0014267
2012-59-0102.mrg | 2012-59-0102.sam | 108 | 176.392 | 176.343 | 0.24814 | -0.04866 | 33.7758 | 33.7784 | 0.00039839 | 0.0026169 | 33.7758 | 33.7807 | 0.0007215 | 0.0048714
2012-59-0102.mrg | 2012-59-0102.sam | 43 | 151.374 | 151.291 | 0.23055 | -0.083221 | 33.5264 | 33.5269 | 0.0033372 | 0.0005188 | 33.5264 | 33.5278 | 0.0030761 | 0.0013618
2012-59-0102.mrg | 2012-59-0102.sam | 106 | 127.048 | 127.105 | 0.20789 | 0.057495 | 33.1642 | 33.1768 | 0.0033041 | 0.012611 | 33.1642 | 33.18 | 0.0040478 | 0.015778
2012-59-0102.mrg | 2012-59-0102.sam | 74 | 101.803 | 101.798 | 0.24217 | -0.0055008 | 32.7491 | 32.7449 | 0.0014918 | -0.0041771 | 32.7491 | 32.7496 | 0.00071428 | 0.00049591
2012-59-0102.mrg | 2012-59-0102.sam | 104 | 76.534 | 76.4527 | 0.20191 | -0.081322 | 32.6839 | 32.676 | 0.0010909 | -0.0078888 | 32.6839 | 32.6785 | 0.00057357 | -0.0054207
2012-59-0102.mrg | 2012-59-0102.sam | 121 | 51.142 | 51.1422 | 0.072927 | 0.00022507 | 32.6355 | 32.6318 | 0.0048955 | -0.0037117 | 32.6355 | 32.6355 | 0.0024449 | -4.9591e-05
2012-59-0102.mrg | 2012-59-0102.sam | 24 | 31.331 | 31.3669 | 0.25821 | 0.035917 | 32.3868 | 32.3784 | 0.047587 | -0.0084038 | 32.3868 | 32.4088 | 0.0037977 | 0.021996
2012-59-0102.mrg | 2012-59-0102.sam | 121 | 22.36 | 22.3599 | 0.19264 | -6.6757e-05 | 32.262 | 32.2865 | 0.029018 | 0.024521 | 32.262 | 32.2818 | 0.0082293 | 0.01984
2012-59-0102.mrg | 2012-59-0102.sam | 97 | 11.529 | 11.5943 | 0.24126 | 0.065319 | 31.952 | 31.9455 | 0.0011439 | -0.0064793 | 31.952 | 31.9471 | 0.00078157 | -0.0049191
2012-59-0102.mrg | 2012-59-0102.sam | 117 | 5.514 | 5.4989 | 0.22149 | -0.015128 | 31.9519 | 31.9441 | 0.00046435 | -0.0078411 | 31.9519 | 31.9463 | 0.00028327 | -0.0055523
2012-59-0104.mrg | 2012-59-0104.sam | 52 | 1503.13 | 1503.09 | 0.24827 | -0.039429 | 34.5128 | 34.5073 | 0.00022292 | -0.0054627 | 34.5128 | 34.5109 | 0.00033733 | -0.001873
2012-59-0104.mrg | 2012-59-0104.sam | 109 | 1251.81 | 1251.75 | 0.14028 | -0.063721 | 34.4556 | 34.4487 | 0.000226 | -0.0069389 | 34.4556 | 34.4517 | 0.000357 | -0.0038795
2012-59-0104.mrg | 2012-59-0104.sam | 47 | 1000.66 | 1000.78 | 0.21563 | 0.11694 | 34.3782 | 34.3733 | 0.0008876 | -0.0048637 | 34.3782 | 34.3763 | 0.00099448 | -0.0018959
2012-59-0104.mrg | 2012-59-0104.sam | 24 | 751.377 | 751.343 | 0.23498 | -0.034119 | 34.2737 | 34.2687 | 0.00027653 | -0.0049858 | 34.2737 | 34.2717 | 0.00039048 | -0.0020218
2012-59-0104.mrg | 2012-59-0104.sam | 121 | 501.632 | 501.632 | 0.083034 | 0 | 34.0496 | 34.0442 | 0.00027649 | -0.0054283 | 34.0496 | 34.0469 | 0.00041638 | -0.0027428
Create Fit Sheet