## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 March 2015 | Correction to header comment about salinity bottles in CHE files. G.G |

## PROCESSING NOTES

Cruise: 2012-71

Agency: OSD

Location: Strait of Georgia

Project: Strait of Georgia Moorings

Party Chief: Johannessen S.

Platform: Vector

Date: October 17, 2012 – October 21, 2012

Processed by: Germaine Gatien

Date of Processing: 15 January 2013 – 6 March 2013

Number of original HEX files: 12 (10 with surface data only for bottles)

Number of CTD files: 2	Number of bottle casts:	12

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1119), a SeaPoint Fluorometer (#2228) with a 3X cable, a Wet Labs Eco-AFL/FL Fluorometer (#2214), an SBE18 pH sensor (#0692) and an altimeter (no serial # available).

The deck unit was a Seabird model 11, serial #0425. The logging computer was #2.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, but no list of personnel. Useful comments were entered concerning problems in equipment and errors in file names.

The rosette logs were printed ahead with planned sampling included. Amendments were sometimes made to the rosette sheets but many salinity samples that were planned never reached the analyst and others that were not planned did. It is most likely that decisions were made on the fly, but not entered on the sheets, but this can only be guessed. This reduces one benefit of the rosette sheets as they are used to ensure that all samples are included in the bottle files. It is all too easy to misplace data in the analysis and processing steps. Having to check the analysts’ spreadsheets for all discrepancies is time-consuming.

CTD and CHE files were prepared for casts #1 and #3; all other casts were restricted to surface sampling so were used to create CHE (rosette) files only.

As has been seen in many cruises in the Strait of Georgia since July 2012, the near-surface salinity bottles have suspiciously high values compared to the CTD salinity. Even at depth the differences are a little high compared to recent offshore cruises (2012-13 and 2012-59) that used the same equipment. Salinity analysis was done within 4 weeks of collection, so evaporation will not explain this, and in any case that would not explain why the near-surface is worst. Two explanations are suggested for these results. Non-linearity in the Autosal could lead to errors that are higher for low salinity which would account for both the inshore cruises having poorer comparisons than offshore, and the results near the surface being worse than at depth. Alternately, incomplete flushing in the presence of high salinity gradients might explain the differences, especially near the surface Tests run by comparing results of this cruise with one from the same project in November 2011 were not able to pinpoint the reason for the large differences between bottles and CTD. This may be a complex issue with several contributing elements. Eliminating the possibility of non-linearity in the Autosal would be a good start towards understanding this phenomenon. There is evidence that the bottles are the problem, not the CTD. The best comparison available from August 2012 shows the CTD secondary salinity to be good to 0.001.

Two fluorometers were used for this cruise but there was no extracted chlorophyll sampling. The fluorescence was low, ranging from 0.15 to 0.65ug/L for the SeaPoint fluorometer. The WetLabs ECO fluorescence varied from about 0.9 to 1.5 times the SeaPoint fluorescence with the ratio increasing as fluorescence increased.

While there were not many dissolved oxygen samples to use for calibration of the SBE sensor, and there was a lot of scatter, a fit was found that was close to the results for the last 2 uses of this sensor.

The Oxygen:Dissolved:SBE data are considered to be roughly:

±0.2    mL/L	 above 50db

±0.05   mL/L	 from 50 to 100db

±0.02 	mL/L	 below 100db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen and salinity. Problems noted were misnaming of most files, the draw thermometers not working, difficulty with dispensers used for DO analysis and one sample not coming from the bottle intended.

Nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. Due to malfunctioning thermometers, the draw temperature was not recorded for DO sampling so concentration cannot be calculated in mass units as well as mL/L for the samples.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

No changes in configuration took place during the cruise.

The calibration constants were checked for all instruments and no errors were found. However, the pressure offset was updated in spring 2012 from the factory setting of -1.3 to -0.1. The new setting was found appropriate for subsequent 2012 cruises. The offset was updated in the configuration file which was then saved as 2012-71-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using configuration file 2012-71-ctd.xmlcon.

There is sometimes a problem converting rosette files with 2 bottles fired close together; a smaller window needs to be chosen to ensure that some CTD data is obtained for both bottles. For this cruise the usual 10s window did work, though the first bottle of cast #8 had fewer data records than ideal. But the setting worked well for most casts, so will be used.

After conversion the file names were corrected to match those in the log book.

The ROS files were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files and no significant outliers were found.

The BOT files were bin-averaged based on bottle numbers.

A preliminary header check turned up no off-scale fluorescence values.

Those files were used to create file ADDSAMP.csv. Sample numbers were added to that file based on information in the rosette log sheets. The addsamp.csv file was sorted on Event_Number and Sample_Number and then converted to CST files. Those will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2012-71-bot-hdr.txt; it may need further editing to reflect problems found during processing.

- DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2012-71oxy.xls which includes flags, comments and a precision study. Draw temperatures are not available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2012-71oxy.csv. That file was converted into individual *.OXY files.

- SALINITY

Salinity analysis was obtained in QF2012-71SAL.xls. The analysis was done 24 days after collection. The files were simplified and saved as 2012-71sal.csv. That file was then converted to individual SAL files. There were duplicates showing good correspondence; the only poor match was from a sample taken when the CTD was at 1.5m, so the Niskin would have been right at the surface and it is quite likely the Niskin contents were stratified.

- NUTRIENTS

The nutrient data were obtained in spreadsheet QF2012-71nuts.xls which included a report on precisions.  The file was simplified, reordered on sample numbers and the file was saved as 2012-71-nuts.csv. The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number since that is the usual method used. The output files were named MRGCLN1s.

Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

Salinity

Compare was run with pressure as reference channel. The fits against pressure show pressure dependence. When the bottom bottles and those above 100m are excluded, there is a reasonably flat fit, with the primary and secondary salinity both lower than the bottles by an average of 0.007 with a standard deviation for both data sets of 0.007. This is based on only 8 bottles. The bottles above 20db are particularly high, though at the surface some are low, likely explained by the local gradients being sufficiently high that the 1.5m distance from CTD to Niskin being significant.

This pattern is similar to those noted since July 2012 in several Strait of Georgia cruises. For some there was a delay in analysis of salinity samples which could lead to samples having higher salinity than expected due to evaporation, though that would not explain near-surface results being so strikingly bad. And for this cruise analysis was quick. Two recent cruises using the same equipment sampled offshore. During 2012-13 the CTD sensor compared very well with bottles below 500db, though there was a fair scatter; analysis was mostly within 7 weeks of collection. There is too much scatter to conclude that there was a trend to higher differences with decreasing salinity, but it looks possible. The results of 2012-59 are similar but the differences are greater and analysis was run about 2 months after collection. Again salinity-dependence is a possibility but that is not clear.

Salinity dependence could also be explained by incomplete flushing since local gradients can be fairly high in this region, particularly close to the surface. There is insufficient data to establish if there is a consistent offset between bottle depth and the depth at which a given salinity sample matches the CTD salinity. The effect could vary with location, season and from year to year as local gradients vary.

This problem has not been noted in previous years, but maybe that was because we didn’t focus on near-surface bottles. To check this, a few comparisons were examined from this same project, but in other years. For cruise 2010-16 in the same region, the comparison was very different with the CTD reading higher than bottles near the surface. The 1.5m difference in height was thought to explain that. For 2010-73 the CTD is lower than bottles above 25m but by no more than 0.02 and from 25 to 75m the CTD looks higher than bottles. For 2012-60 there is a lot of scatter above 50m with many cases of CTD lower than bottles, but those are right at the surface and there are also cases in the opposite direction. For cruise 2011-76 there are cases of the CTD being much lower than the bottles but these are only from bottles right at the surface, above 5m (aside from one clear outlier). There is nothing that looks like the recent results.

One further test was done by comparing the 2 deep casts of this cruise with the data from the same stations a year earlier. In 2011 the salinity analysis occurred about 6 weeks after collection and the comparison looked good with no large differences below 50m.

- For station SOGS the overall salinity gradient is higher in 2012 than 2011, but in the top 25m it is similar or slightly lower. The salinity in the top 30m is slightly lower in 2012. So we would not expect larger differences in the comparison of CTD and bottles in 2012 than 2011 if flushing is the issue. But evidence for non-linearity is also weak since the 2012 salinity values don’t seem sufficiently lower to cause such a large change in differences.

- For station SOGN the salinity gradient is significantly higher in 2012 than in 2011 above 75m and the salinity is higher above 100m in 2012. The differences in the comparison near the surface are much smaller for SOGN than SOGS as would be natural if the cause of outliers is either low salinity or poor flushing. So this doesn’t help distinguish between the possible causes. The comparison results for 2011 and 2012 are similar at SOGN.

- The descent rate of the CTD looks similar in the two years. This argues against another theory that quiet sea states in 2012 might have lead to poorer flushing of bottles given less shaking of Niskin bottles due to vertical motion during stops. SOGN looked noisier than SOGS for both years.

The cause may be complex with evaporation of samples, poor flushing and non-linearity of the Autosal all contributing and with the complication of the 1.5m vertical offset between bottles and CTD making interpretation especially tricky in very high gradients. To at least rule out one explanation, tests should be run to establish whether the Autosal has a linearity problem.

The flag (36) attached to sample #24 may be due to stratified contents in the bottle. The local salinity gradient was very high, so it is not surprising if the replicates differed significantly. No change is suggested to the flag.

The comparison depends too much on arbitrary identification of outliers to justify recalibration.

For full details for the COMPARE run see file 2012-71-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. Following recent recommendations from SeaBird recalibration will be based on a fit of differences between bottles and CTD DO with the offset forced to 0. The manufacturer warns that this will not fit well for values <2mL/L due to limitations in Winkler titrations at low values, but there are no values in that range from this cruise. Based on cruises 2012-13, 2012-59 and 2012-05 it was decided to allow an offset because the fits were poor otherwise. If the low DO bottles are excluded from the fits, the offset does not usually change significantly.

There are few points and a lot of scatter, so finding a fit is very dependent on precisely how outliers are chosen. If just one clear outlier is removed in the fit of differences against DO concentration the fit is

Bottle DO = 1.0552 * CTD DO

Or if an offset is allowed:

Bottle DO = 1.0279 * CTD DO + 0.1042

The scatter is very high – these are both poor fits. If 3 more outliers are excluded the fit is:

Bottle DO = 1.0457 * CTD DO + 0.0518

This result is very close to the results of the last 3 cruises which used the same equipment. The most recent result from 2012-06 (in the same region) was:

Bottle DO = 1.0464 * CTD DO + 0.0138

For 2012-59 off the west coast of Vancouver Island, the result was:

Bottle DO = 1.0465 * CTD DO + 0.0329

For full details of the comparison see 2012-71-dox-comp1.xls.

The only significant outlier was associated with a high standard deviation in the CTD DO and a very high DO gradient in both downcast and upcast. It came from a depth where the CTD DO sensor often seems to “overshoot” as it leaves the high gradient region. The sample is likely fine, so no flag should be added.

Fluorescence

There was no extracted CHL sampling, but there were 2 fluorometers, so a comparison was made of how those compared during bottle stops. The ratio of the WetLabs ECO fluorescence to the SeaPoint Fluorescence varied from about 0.9 to 1.5 with the values increasing as fluorescence increased. The range was small with SeaPoint fluorescence between 0.15 and 0.65ug/L.

For more detail see 2012-71-dual-fl-comp1.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no problems were detected.

At this point the data from the MRGCLN2 files were exported to a spreadsheet for comparison with the rosette files to ensure no data were misplaced or missing.

One problem was found. Niskin #11 did not fire for cast #3. Salinity sample #35 was taken from Niskin #12, so is a duplicate of Sample #36. The MRG file was fixed using the average of the duplicates. The DO sample was taken from Niskin #10 and sample #34 was used, so no change is needed for that.

##### Conversion of Full Files from Raw Data

All files were converted using 2012-71-ctd.con. After conversion the file names were changed so that the event # section matches the log book entries.

A few casts were examined and all expected channels are present. The descent rate was steady and high for the two casts that sampled below the surface.

The two temperature and conductivity channels are reasonably close during the downcasts but they differ much more during upcasts. The conductivity channels have some spikes in upcasts.

Altimetry looks fine. The SeaPoint fluorescence channel is lower than the WetLabs ECO fluorescence near the surface, but close below that. The lowest values are close on average, but the noise level is much higher in the ECO fluorescence. The pH, DO and transmissivity look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on a few casts to determine the offset between the DO voltage and the primary temperature. It is very hard to judge from offsets between upcasts and downcasts because the temperature is so noisy on the upcast. For other recent cruises using this equipment settings of 4.5 or 5s were found appropriate and for these data 4.5s looks like the better choice.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

The upcast data are extremely noisy so the usual tests for CELLTM settings do not give clear results. The same equipment was used for 2012-13, 2012-59 and 2012-06 and for those cruises results were best with a choice of (α = 0.02, β=7) for the primary and (α = 0.03, β=7) for the secondary.

CELLTM was run using (α = 0.02, β=7) and (α = 0.03, β=7) for the primary and secondary conductivity channels, respectively.

##### DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

- on the only 2 casts that sampled to 300db to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The only deep casts were plotted to check for agreement between the pairs of T and C sensors. These are shallow casts, but were examined to ensure there were no obvious problems. For comparison a selection of tests from other 2012 cruises with the same T/C sensors are shown.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2011-17-0121 | 800 | -0.0004 | -0.00008 | -0.0005 | High, V Noisy |
| 2012-05-0032 | 250 | 0.0001 | +0.00015 | +0.0018 | Steady, high |
| 2012-05-0039 | 350 | 0.0002 | +0.00014 | +0.0013 | Steady, high |
| 2012-16-0001 | 220 | -0.0001 | +0.00025 | +0.0026 | Steady, moderate |
| 2012-57-0013 | 270 | -0.0002 | +0.00025 | +0.0027 | Steady, high |
| 2012-58-0081 | 240 | +0.0001 | +0.00006 | +0.0006 | V.steady, high |
| 2012-13-0052 | 500 | +0.0002 | +0.0001 | +0.001 | Noisy, high |
| 2012-59-0102 | 500 | +0.0004 | +0.0003 | +0.0036 | XNoisy, high |
| 2012-71-0001 | 280 | -0.0002 N | -0.00002 | -0.0001 | Steady, high |
| 2012-71-0003 | 300 | -0.0002 | +0.00006 | +0.0006 | F.Steady, high |

The differences are small and do not suggest any systematic drift.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are a few slightly negative values in pressure in one cast, but they are at the end when the pumps were off, so the CTD might have just left the water.

Surface check was run and shows an average surface pressure for the cruise was 1.7db which looks about right for this cruise which was trying to sample very close to the surface.

The cross-reference check was compared with the log book. No problems were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and SAM files were exported to a spreadsheet. For the 2 deep CLN files, plots confirm the header entries. The CLN files for the surface-only casts will not be processed further so no changes were made to those.

For the bottle files, there are 4 cases where there are altimeter headers with values between 1 and 2db, but the CTD only sampled at the surface, more than 15m from the bottom. In those cases the header entry was removed from the SAM files and the final MERGE and CLEAN steps were rerun.

The Water Depth headers were also examined. No errors were found.

##### Shift

Note: From this point onwards only casts 1 and 3 will be processed for CTD files.

Fluorescence

Tests were run on the two deep casts to see what SHIFT values should be used to make the offset between the downcast and upcast fluorescence traces look like that of the temperature trace. There was so much noise that the signal is hard to find; even after filtering the fluorescence there is no clear pattern. Shifts of +24 records for the SeaPoint fluorometer and +6 records for the ECO fluorometer were used during all recent cruises using this equipment.

SHIFT was run on all casts to advance the ECO fluorescence channel by +6 records and to advance the SeaPoint fluorescence by +24 records.

Conductivity

For 2012-06 the conductivity shift results differed from previous cruises with the same equipment. Tests were run on cast #1 to see if the 2012-06 parameters suit this cruise as well and they did.

SHIFT was run twice on all casts using -0.8 records for the primary conductivity and -1.2 records for the secondary.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. The 2 casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary, though the upcast temperature is too noisy to have a lot of confidence in that conclusion.

pH

This pH channel does need alignment. The setting used for 2012-06 was +65 records and that appears to work quite well on these data as well, though again the tests are hard to interpret.

SHIFT was run on both casts using +65 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The conductivity sensors were both recalibrated in late March 2011 and were used for many cruises since then. There have not been many good comparisons with bottles; the best available were from 2012-05, 2012-13, 2012-59 and 2012-06. The first, analyzed after 1 week, indicated that primary salinity was low by about 0.001 and the secondary was high by about 0.0004.  The second, analyzed about 6 weeks after collection, showed the primary to be low by 0.002 and the secondary low by 0.0005. The difference between those results may be due to storage time or drift. 2012-59 had samples analyzed after 2 months storage; the primary salinity was found to be low by an average of 0.0047 and the secondary by 0.0015. That difference is about the size of error expected after 2 months storage. For 2012-06 with samples stored 7 months there was a lot of scatter and the salinity was found to be low by 0.0025.

2. Dissolved Oxygen

The DO sensor was recalibrated in April.2011 and has been used for 11 cruises since then. Only 2012-13, 2012-59 and 2012-06 used the bottle comparison method now recommended by SeaBird. Recalibration was done using a linear correction with slope/offset = 1.0435/0.0159 for 2012-13, 1.0465/0.0329 for 2012-59 and 1.0464/0.0138 for 2012-06.

3. Pressure

The pressure sensor has been used 9 times in 2012 and the offset was increased from -1.4 to -0.1 in April 2012. No further correction has been found necessary since then.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All data fell within the climatology.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration - There were no post-cruise calibrations available.

##### DETAILED EDITING

The secondary sensors have been selected for all recent uses of this equipment. COMPARE does not suggest which pair is better, so the secondary were selected.

CTDEDIT was used to remove or clean small spikes in salinity that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface and bottom records. Editing was light.

All edited files were copied to EDT.

##### Initial Recalibration

The pressure data appear to be ok with the updated configuration file, so further recalibration is unnecessary.

The results of the comparison of CTD salinity with bottle salinity suggest that it is low by 0.007, but the evidence is weak because of high scatter. The history of the secondary T/S sensor pair suggests that the salinity is within 0.001, so no recalibration will be applied. This can be revisited if post-cruise factory checks indicate there are larger errors.

The SBE Dissolved Oxygen channel does need recalibration but the question arises whether to use the results of 2012-06 or 2012-71.

A first attempt to recalibrate used the results of 2012-06. The CTD DO looked too low by about 0.04mL/L for most stops.

For a second attempt the results of this cruise were used:

Corrected DO = 1.0457 * CTD DO +0.0518

The MRGCLN2 and SAM files were recalibrated using file 2012-71-recal1.ccf.

COMPARE was rerun to ensure the recalibration was applied appropriately and it was.

There is a lot of scatter in the fit against DO but the average difference is very small. The fit against pressure looks odd but it is good in the low DO gradient zone between 125 and 300db.So the 2012-71 comparison will be used to recalibrate.

For details see 2012-71-dox-comp2.xls.

CALIBRATE was applied to the EDT files using file 2012-71-recal1.ccf to apply the above correction.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. ALIGNCTD corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.25m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles.

When the differences were plotted against DO concentration and a few outliers were excluded, the fits show do show some dependence on pressure and dissolved oxygen concentration, but there is too much scatter and the differences are too small to justify a second recalibration. The near-surface DO gradients are lower than during 2012-06 which probably accounts for the fact that the DO sensor performed better near the surface.

The Oxygen:Dissolved:SBE data are considered to be, very roughly:.

±0.2    mL/L	 above 50db

±0.05   mL/L	 from 50 to 100db

±0.02 mL/L 	below 100db

(See 2012-71-dox-comp3.xls.)

##### Special Fluorometer Processing

There was no CHL sampling so special files were not prepared for Dr. Peña.

There were no off-scale fluorescence data.

The two fluorometer channels were put through a filter, size 11 to reduce spikiness.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

From this point on 2 sets of CTD files were prepared; one is intended for the OSD_Archive and the other which will include pH data will be provided to the chief scientist. The same steps will be taken with the exception of what channels are removed. The special files will have extensions that end in pH.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, pH:SBE, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity and Fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

NOte #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered to be roughly:

±0.2    mL/L	 above 50db

±0.05   mL/L	 from 50 to 100db

±0.02 	mL/L	 below 100db

For details on the processing see processing report: 2012-71-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The surface values were 88% at SOGS and 76% at SOGN. The low values are likely due to mixing. The SBE DO values are close to bottles values at the surface, being slightly low at SOGN and slightly high at SOGS.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

From this point on 2 sets of files were produced, one for the OSD Archive without pH and one for the chief scientist with pH. The pH files have extensions that end in pH.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

Change Units was run to add a second version of Oxygen:Dissolved:SBE in mass units. Normally this would also be applied to the Oxygen:Dissolved channel, but the draw temperature is not available for this cruise due to a malfunctioning thermometer.

REORDER was run to get the 2 sets of DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and a few errors were found and fixed.

Header Check was run on the final files until all problems were resolved.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

General: No draw temperatures – thermometers wouldn’t work

DO Sampling – problems with pipette dispensers for chemicals

DIC alkalinity samples: Rep B collected before Rep A for casts 4-13.

Original file names off by one for casts 3-13. Ex. 2012-71-0003 was changed to 2012-71-0004.

3. Bottle 11 did not fire – extra O2 sampled was taken from bottle #10 (Daily Science log actually says taken from bottle #11, but rosette log and e-mail from Cindy Wright say it came from bottle #10).

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 1Apr2011 | Factory |  |  |
| Conductivity | 3396 | 29Mar2011 | Factory |  |  |
| Secondary Temp. | 2668 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 2754 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | Jan2012 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 29Mar2011 | Factory |  |  |
| pH | 692 | Dec2012 |  |  |  |
| SBE Fluorometer | 2228 |  |  |  |  |
| Eco-AFL Fluorometer | 2214 | Sept2012 |  |  |  |
| Pressure Sensor | 69698 | 15apr2011 | Factory |  |  |
| Altimeter | 43281 |  |  |  |  |