## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2012-77

Agency: Ocean Sciences Division, Sidney, B.C.

Project: Saanich Inlet Prawn Survey

Area: Saanich Inlet

Chief Scientist: Galbraith M.

Platform: Elvis

Date: 25 October 2012

Processed by: Germaine Gatien

Date of Processing: 26 October 2012 – 29 October 2012

Number of original CTD casts: 7

Number of casts processed: 7

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19+ SEACAT CTD (S/N#5299) was mounted with an SBE 43 DO sensor (# 1483).

SUMMARY OF QUALITY AND CONCERNS

CTD Daily Log sheets were available with details on surface soaking of the instrument.

The soak period for the CTD was at least 1 minute.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. However, for this cruise the descent rate was steady and the gradients were not high, so the errors are likely much smaller than that.

No salinity calibration sampling was available and there is no history for the temperature and conductivity sensors since the last factory calibration.

There was DO calibration information from a 2011 cruise and this was used to update parameters SOC and Voffset in the configuration file.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained. There were no notes indicating problems with the CTD.

The CTD was turned on in air, put in the water and soaked for 1 minute (or longer) before lowering.

The CTD sampled 4 times per second.

- Conversion of Raw Data

- The parameters in the configuration file used at sea were correct and were saved as 2012-77-ctd.xmlcon. There was no calibration sampling, so the results of 2011-08 were used to update the Soc/Voffset values for the dissolved oxygen sensor and the updated file was saved as 2011-77-ctd-new.xmlcon.

- The raw data were converted using conversion file 2012-77-ctd-new.con.

- After conversion the file names were changed to standard format.

- Examination of a few plots suggests that the surface was close to 0db both for the downcast and upcast judging by the conductivity, with the maximum error ~0.2db. There were soaks of at least 1 minute at about 1db, but most were a little longer than that. A 2-minute soak is recommended for this instrument because the pumps don’t come on for 30s and there is some equilibration after that; however, there is no evidence of significant problems with the downcast temperature and salinity for these data so the soak was probably adequate.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature. It looks smoother after this step.

- Align CTD

Tests were run to find appropriate advancements for T and DO. The results are not easy to interpret, but a value of +0.5s for temperature looks reasonable and is recommended by SeaBird.

The dissolved oxygen has been advanced by 4.4s in previous uses. This value looks best overall, but for some features, different values look better.

The temperature was advanced relative to pressure by 0.5s.

The dissolved oxygen was advanced relative to pressure by 4.4s.

- DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets and in the SeaBird headers. Spreadsheet 2012-77-hdr_merge.csv was prepared with that information.

Program MERGE CSV FILE TO HEADERS was run to add the headers to the CLN files.

8. Checking Headers

A cross-reference list was produced and checked against the log sheets. The only error found was that time was in PST, not UTC.

ADD TIME CHANNEL was used to add 8 hours to each header time.

Header Check was run and showed some negative DO values. This will be investigated later to see if recalibration is required.

The track plots look fine except for the first cast which appears to be on land. This is due to lack of fine detail in the coastline file; this cast was near shore. A better chart shows this position is in about 30m of water.

The surface check was run. The average surface pressure is +0.11db with very low surface conductivity values except for cast #3 for which the salinity is 15 at 0.19db

The pressure calibration looks appropriate and no correction is indicated.

9. SHIFT

Examination of downcast T-S plots suggests no further alignment of conductivity is required. Tests were run on one cast to be sure, and all adjustments made the few slightly unstable features worse rather than better.

SHIFT was not run.

10. Test Plots and Fix Time

Profiles were plotted for all casts and no serious problems were noted.

11. DELETE

CLIP was run to remove some initial records during the soak period.

Then DELETE was run on the CLIP files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings in the Delete log.

12. CTDEDIT

CTDEDIT was used to remove bad records near the top and bottom of casts and to smooth salinity lightly, as required. All casts required some editing except for #5.

Notes were entered in the headers about editing applied.

13. Study of negative DO values

Profiles of DO were examined in detail to determine how to remove negative values. All the negative values occur where DO suddenly shifts to anoxic conditions and temperature falls sharply. The oxygen voltage goes down gradually, but the DO concentration overshoots to values ~-0.07mL/L, then rises to about +0.05 then slowly settles to a stable value of ~+0.03mL/L. Adding an amount sufficient to remove negative values does not look appropriate. Instead CLEAN was run to replace negative DO values with pad values.

14. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and no further editing was needed.

15. Intercomparisons

Previous Use of CTD – This is the first known use of the CTD since it was last checked at the factory, but the DO sensor was used 3 times in 2011 with DO calibration sampling for one of those cruises.

Repeat Casts – There were no repeat casts.

Historic ranges – There is no local climatology for this area.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Oxygen:Voltage, Descent_Rate and Flag.

The HEADEDIT routine was used to fix channel names, units and formats and to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients.

For these data the descent rate was quite steady and temperature and salinity

gradients were not extreme, so the error is likely much smaller than that.

No calibration sampling was available for this cruise or any other cruise

since the last factory calibration.

The dissolved oxygen parameters in the configuration file were updated by using

the values for SOC and Voffset that were determined by calibration sampling

during cruise 2011-08 in April 2011..

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2012-77

Dates: 	Start:	25 October 2012	    End:	25 October 2012

Location: Saanich Inlet                 Vessel:    Elvis

Party Chief: Galbraith M.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SeacatPlus | 5299 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/SEACATPLUS v1.6b / 5299

Cruise ID#:	2012-77

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5299 | 12Jan11 | Factory |  |  |
| Conductivity | 5299 | 12Jan11 | Factory |  |  |
| SBE 43 DO sensor | 1483 | 24Dec10 | Factory |  |  |
| Pressure Sensor | 5299 | 29Dec10 | Factory |  |  |