## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 March 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2013-12

Agency: Ocean Sciences Division

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: April 10, 2013 –April 16, 2013

Processed by: Germaine Gatien

Date of Processing: 28 August 2013 -4 October  2013

Number of original HEX files: 76

Number of CTD files: 76	Number of bottle casts:	21

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1438) on the primary pump, a SeaPoint Fluorometer (#2228) with a 3X cable on the secondary pump, a Biospherical QSP-400 PAR sensor (#4601), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (no serial # available).

The deck unit was a Seabird model 11+, serial #0425. The logging computer was #2.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, Rosette log sheets and DO analysis logs were in good order.

There was good correspondence between the two CTD salinity channels and between those and the samples from bottles, which suggests that there has not been significant calibration drift since the last factory service. The salinity samples were analyzed within 8 days of collection, showing that good results can be obtained if analysis is prompt. There were outliers but they appear to be associated with shallow sampling and/or local variability and in two cases incomplete flushing of bottles is suspected in areas of calm conditions and high local gradients. There is no evidence of non-linearity in the salinometer.

The differences between the SBE dissolved oxygen sensor values and titrated samples had a relatively tight fit against the titrated values, and the slope of the fit for this particular instrument shows a slow increase with time since it was last calibrated at the factory.

The fluorometer values are low compared to extracted CHL. There were some concerns when this fluorometer was last used because values were low for CHL>0.8ug/L, but they were close to CHL at lower values. For this cruise the range of CHL is much larger and there are very few CHL values <0.8ug/L. As always, the fluorescence data must be considered nominal.

A pH sensor was attached to the CTD, but the data are not included in the files to be placed in the IOS Data Library because no field calibration information was available at the time of processing, and there are some concerns about how the sensors perform. Data including the pH channel were prepared for the use of the Chief Scientist.

The Oxygen:Dissolved:SBE data when the CTD is in motion are considered, very roughly, to be:

±0.6  mL/L from 0 to 50db

±0.2  mL/L from 50 to 100db

±0.1  mL/L from 100 to 200db

±0.05 mL/L below 200db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for extracted CHL, dissolved oxygen and salinity. There were problems with bottles failing to fire.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained

The calibration constants were checked for all instruments and no errors were found, though the date of one calibration was incomplete, so that was added. There were no changes through the cruise.

One file was saved as 2013-12-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All files were converted using 2013-12-ctd.xmlcon.

File #73 would not convert at first – there was an error in the headers (3 lines concatenated). When that was fixed, the file converted successfully.

A few casts were examined and all expected channels are present. As is usual for this program the descent rate varies from very steady in protected waters to extremely noisy at the mouth of Juan de Fuca Strait.

The two temperature and conductivity channels are fairly close during the downcasts and, as usual, much farther apart during the upcasts. There are many spikes in the conductivity, particularly in the primary but the secondary has many as well. These are generally not associated with significant spikes in temperature and are seen in both downcasts and upcasts. Spikes were not noted when this equipment was last used.

Altimetry is sometimes spiky, but there appears to be a useful signal at the bottom, and fluorescence, transmissivity, pH, PAR, SPAR and transmissivity look normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2013-12-ctd.xmlcon.

One ROS file will not be archived because bottles were fired as a test only and no samples were drawn and no sample numbers were assigned.

The files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files.

CTDEDIT was used to examine a few cases where the data seemed a little noisy, but the variations look real.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records. There were a few adjustments needed:

- Cast #29 - two bottles were fired at 20m because Niskin #13 had failed to close frequently, but on that occasion it did close and samples were taken from Niskin #13. Niskin #20 was not sampled and there is no sample number associated with it. That line was removed from the ADDSAMP file.

- Cast #33- the surface bottle (Niskin #13) did not fire, so a second bottle was fired (Niskin #14) and the sample # intended for Niskin #13 was reassigned to #14. The line for Niskin #13 was removed from the ADDSAMP file.

- Cast #41 - a second bottle was fired at 20m but was not needed since Niskin #13 did close, so the line referring to Niskin #20 was removed from the ADDSAMP file.

- Cast #43 - The log notes are not clear - 2 bottles were fired at 30m. The rosette log suggests both Niskin #11 and #18 were sampled with sample #134, but the Daily Log says that Niskin #11 fired properly, so the sample number was assigned to #11 and the line for #18 was dropped from the ADDSAMP file.

- Cast #46 -2 bottles were fired at the surface but the first closed and the second was not sampled. The line for Niskin 12 was removed from the ADDSAMP file.

- Cast #59 – 2 bottles were fired at 50m but the second (Niskin #20) was not needed so that line was removed from the ADDSAMP file.

- Cast #75 – sample #244 was not used.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-12-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2013-12chl*.xls. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2013-01chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-12oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2013-12oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2013-12SAL.xls. The analysis was done within 3 to 8 days of collection. The files were simplified and saved as 2013-12sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-12nuts.xls. This includes a precision study.

The file was simplified, reordered on sample numbers and the file was saved as 2013-12-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts to see what setting brought the vertical offset between the downcast and upcast DO traces into the best agreement with the temperature traces. When this sensor was used during 2013-01 tests found that an advance of +4.5s was best, on average, though results were better in places with either +4s or +5s. The tests for 2013-16 were even more confusing. There is a lot of variability in these data, so the results are hard to interpret, but a value of about +5s looks too high in some places, and 4.6s looks ok in some areas and too low elsewhere. A choice of 4.8s looks like a reasonable compromise.

ALIGNCTD was used to advance the DO Voltage by 4.8s relative to the pressure.

##### CELLTM

A variety of settings were used on 3 casts. The results showed little difference between them, presumably because the temperature gradients were low. Overall, the setting of (α = 0.02, β=7) looked best for both primary and secondary channels, but the differences were slight.

CELLTM was run using (α = 0.02, β=7) for the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a selection of deeper casts to examine differences between sensor pairs. None of the casts are very deep, but the differences were fairly steady for events 22 and 34.  And all were reasonably small at depth. The differences are similar to the results found during 2013-01, so calibration drift is likely small.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2013-12-0022 | 400 | -0.0002 | -0.00003 | -0.0002 | F.Steady; high |
| 2013-12-0034 | 330 | -0.0001 | -0.00004 | -0.0003 | Steady, high |
| 2013-12-0059 | 360 | +0.0001 | -0.00007 | -0.0008 | Noisy, medium |

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure shows a lot of scatter. 4 outliers were identified with large differences, of which 3 are associated with large standard deviations in CTD salinity during the 10s window. On average, when outliers are removed, the primary sensors are closer to the bottles, but the average difference for both pairs is <0.001 with standard deviations ~0.003. There is no significant pressure-dependence in either pair.

For the 4 outliers the CTD reads lower than the samples. For the two from 99m there is a lot of variability in the CTD data, so this looks like a mismatch with neither CTD nor Niskin sample being more or less reliable. For event #1 there is a difference of 0.06 and examination of the full data file suggests that the Niskin bottle contained water from shed wakes that passed through during the stop. Incomplete flushing is more likely in Saanich Inlet where conditions are usually quite calm. Once the shed wakes stopped there may have been little movement of the bottle to encourage flushing. The deepest outlier is from 175db and the CTD data indicates a lot of variability. Incomplete flushing might be an issue here as well.

There appears to be some time-dependence but since the slope of the fit is exactly the same for both sensor pairs, this is likely a geographic dependence with either flushing efficiency, varying sample depths or local variability accounting for the differences.

These results suggest that the salinometer is working well and if samples are analyzed quickly we can get useful results. They also show that incomplete flushing of bottles is likely significant in areas where salinity gradients are higher and conditions calm. This region is not one where we can expect good enough calibration sampling to justify recalibration of salinity, but with quick analysis of samples we should be able to detect significant drift in calibration.

A comparison of differences against primary salinity shows no suggestion of non-linearity in the salinometer. There is more noise at low salinity but this is expected because most low-salinity samples are near the surface. When the outliers are excluded the differences are close to 0 for salinity <32 and about +0.001 for the 4 samples with salinity >32. Given the scatter in the fits, these are close results.

All outliers are associated with high standard deviation in the CTD data and/or come from areas where local gradients are high. No flags were added.

Both salinity channels are close to bottles, on average, with the primary slightly closer. Given the noise level the difference between them is not significant.

For full details for the COMPARE run see file 2013-12-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There are two major outliers:

Sample #160 was already identified as a misfire and this confirms it. Flag 5 appropriate.

Sample #245 – no problem was found in analysis; CTD data bit noisy but not enough to explain difference. Flagged 4.

Minor outliers were excluded based on residuals. SeaBird recommend using a fit with offset=0, but for these data it looked better to allow an offset. Based on the fit, recalibration will be applied as follows:

CTD DO Corrected = CTD DO Raw* 1.0706 +0.0418

That correction was applied to the MRGCLN2 and SAM files. COMPARE was rerun to ensure the correction was appropriate and it was.

For full details of the comparisons see files 2013-12-dox-comp1.xls and 2013-12-dox-comp2.xlsx.

One other sample (cast #59, sample #176) was investigated because the analyst thought there might be a problem with the Niskin bottle.  The CTD data collected during the stop for bottle #11 is a little warmer and saltier and contains less oxygen than that from Niskin #20. A look at the downcast CTD profile reveals temperature and oxygen varying a lot from 45 to 65db, while salinity steadily increases with depth so the water column is stable. This looks like interleaving of two somewhat different water types. So it would be unwise to judge the performance of Niskin #11 on the basis of the variations in the 50db samples. The choice of which sample to use with the CTD data is not critical though, since what is in the bottle is not expected to correspond exactly to either of them.  The sample does not stand out as a significant outlier in COMPARE. So the comment was adjusted to:

OXY:  Value seems low.  Bot 20 was also closed at 20 m. Bot 20 [DO] = 4.452 mL/L. Temp and 	DO highly variable at this depth so difference likely not significant.

Plots of Titrated DO and CTD DO against CTD salinity were examined and significant outliers were in high gradient areas. No additional flags appear justified.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The CTD fluorescence values are an average of 34% of the extracted CHL samples, with a range of 15-50% with the exception of the 20db bottle from Saanich Inlet where the fluorescence was 70% of the CHL. That sample had the lowest CHL value of the comparison, and we often see a higher FL/CHL ratio for low CHL. There is no dependence on pressure or time.

During 2013-01 it was also found that the fluorescence seemed lower than expected for CHL>0.8ug/L. For that cruise the cable was said to be 30X and tests were done to see if 10X was better. It did look better for high CHL values but worse for lower values and dark values looked unrealistic.

For this cruise the cable is said to be 3X, so the only way to achieve values closer to CHL would be to assume a 1X cable. The comparison is excellent if that assumption is made, but the log and configuration files used at sea both indicate it was 3X and 1X is an unlikely choice.

For full details of the comparison see file 2013-12-fl-chl-comp1.xlsx.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  The minimum pressure was 0.47db. The speeds look fine and the fluorescence does not go off-scale.

Surface check was run and shows an average surface pressure for the cruise was 2.7db which is a little high for the Vector, but not unreasonable. The minimum pressure during the cruise was 0.47db and occurred at the end of event #22. The CTD was held at about that depth for 9s and then moved a little deeper until the cast ended. During that time at the minimum pressure the transmissivity spiked to zero and then back to high values. Since it didn’t stay low, this does not look as though the CTD was out of water; more likely there was some debris near the surface that interfered with light transmission. Both salinity values were low at the minimum pressure, but not unbelievably so. The primary salinity values reached ~20 when the CTD got back down to ~0.7db while the secondary reached a value ~27 at 0.7db. There were a few other times when salinity of one or the other sensors was quite low which could be due to local variability, or the low transmissivity might be associated with something that was interfering with flow to the pumped sensors. Fluorescence is zero for some of the time.

If the CTD was out of water when it read 0.5db, then the rosette would have been way out of the water and this seems unlikely with the pumps still turned on. It seems unlikely that the CTD pressure is significantly high, since most often drift in these sensors leads to values that are too low, but it is possible. However, there is insufficient evidence to warrant recalibration of pressure.

The cross-reference check was compared with the log book and the only problem noted is that the station name for event #1 was entered as Saanich Inlet, rather than the usual format and the one found in the log book, SI. The station name entries in the IOS and CLN files were changed to SI.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom and most altimetry headers look appropriate despite a very noisy signal near the bottom of some casts. All entries checked were appropriate. There were some cases where there was no header entry even though the CTD did get near the bottom; this was due to spikes. For bottle files for casts #41, 54 and 59, the CTD does appear to have been close to the bottom during the downcast, so header entries were added based on examination of the full profile. And for cast #34 a header was also added to the CTD file.

Depth entries are within 6m of those in the log book, mostly quite close. Since the ship may have drifted between preparing the header and writing in the log, it is not clear which value is more appropriate and this is a rough estimate anyway since the depth when the CTD reached bottom may well be different from either of these values. So no changes were made.

##### Shift

Fluorescence

SHIFT was run twice on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +55 and +75 records. The results vary from cast to cast and depth to depth. On a few 2012 cruises using this sensor, a setting of +65 records was found best, but for 2013-16 a value of +75 records was selected since it generally does a better job near the surface and most casts were very shallow. For this cruise +65 looks best. SHIFT was run on the pH:SBE channel using a setting of +65 records.

Conductivity

During 2013-01 the shift settings chosen for the conductivity channels were -0.5 records for the primary and +1.5 records for the secondary, but the CTD configuration was quite different for that cruise. During the last use, 2013-16, with similar configuration but very shallow sampling, the choice of -0.5s looked best for both channels.

Tests were run on 2 of the deeper casts using a variety of shifts. The choice of -0.5 records again proved to be best for the primary, but for the secondary -0.3 records was best.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary. At times the upcast vs downcast DO traces seem closer than those for temperature, but for some features the opposite is true. Overall the alignment looks good.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late March 2011 and were used for 1 cruise in 2011 with no salinity sampling, during 2013-01 with good sampling and 2013-16 with poor sampling. Both salinity channels were found to be within 0.001 of bottles, on average, during 2013-01.

Dissolved Oxygen

The sensor has been used only 3 times since its last recalibration in March 2011. For one of those a different method of checking calibration was used. For 2013-01 with a good range of DO values, the fit found had a slope of 1.066 and offset of +0.0189. For 2013-16 the offset was forced to =0 and the slope was 1.07.

Pressure

The sensor was recalibrated in April 2011 and was used for 1 cruise in 2011 as well as 2013-01 and 2013-12. The factory offset was used for all those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature and salinity profiles fell within the local climatology except for the 3 casts on the west coast of Texada Island which had low temperatures at depth, from 200m down, 250m down and around 300m at stations 7, 8 and 12, respectively. These look like real variations rather than indications of instrumental problems. Going a little outside the 3 standard deviation criterion is not unusual so close to shore.

Repeat Casts – There were no repeat casts and there is too much variability for comparison of nearby casts to have any value in assessing repeatability of the sensors.

Post-Cruise Calibration - There were no post-cruise calibrations available.

##### CHANNEL CHOICE STUDY

COMPARE does not help in choosing which sensor pair to select for archiving. The two salinity channels are very close in value during stops. T-S plots show only slight differences with sometimes one pair and sometimes the other looking smoother. Since the primary was used for 2013-01 and 2013-16 the alignment settings were the same for the primary as during those 2 cruises, the primary temperature and salinity channels were selected.

##### DETAILED EDITING

The primary temperature and salinity channels were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the cast.

Cast #45 required no editing. All other casts required some editing, mostly at the top and bottom of casts, with only casts #45 and #73-76 requiring heavier editing.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure does not appear to require recalibration.

No salinity recalibration will be applied. Based on the results of 2013-01 and the fact that the two sensor pairs are very close during stops, it is likely that there has not been significant calibration drift in any of the temperature or conductivity channels.

SBE Dissolved Oxygen was recalibrated by applying the following equation:

CTD DO Corrected = CTD DO Raw* 1.0706 +0.0418

CALIBRATE was run using file 2013-12-recal1.ccf to apply the correction to the EDT files.

COMPARE was rerun on the recalibrated data and the results were good.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure the CTD looked low by about 0.2mL/L for the highest DO values and high by about 0.05mL/L for DO=2mL/L. Fits of differences against CTD DO were studied; the fits were fairly insensitive to the methods of identifying outliers. For example, excluding all values of DO>6mL/L produced a similar result to that when they were included. And using only casts with no DO reversals in the profiles produced similar results. File 2013-12-recal2.ccf was prepared to apply the following correction based on the results of COMPARE:

DO corrected = DO*1.0394 – 0.1278

COMPARE was run again to ensure the correction was applied correctly and it was.

(See 2013-12-dox-comp3.xlsx and 2013-12-dox-comp4.xlsx for details. Note file 2013-12-dox-comp4a.xlsx was created to save a comparison versus pressure after final corrections to enable an estimate of SBE DO errors.)

##### Fluorescence Processing and special files for Angelica Peña

The COR2 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2012-04-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. There were a few small unstable features in the T-S plots but some are expected in this region. There is no evidence that these are due to instrumental problems. Profile plots showed that the transmissivity was close to 0 for cast #65 at station #70 on the north side of Juan de Fuca Strait; the altimeter indicates that the CTD was 7m off the bottom. The low values occur in a well-mixed 20m bottom layer. This was not a bottle cast.

##### Final CTD File Steps (REMOVE and HEADEDIT)

The steps in this section were run twice – first to produce CTD files, and then a second run to produce CTDpH files that include the pH:SBE channel.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6  mL/L from 0 to 50db

±0.2  mL/L from 50 to 100db

±0.1  mL/L from 100 to 200db

±0.05 mL/L below 200db

The two CTD salinity channels were generally within 0.001 of each other.

The comparison with bottles had a lot of scatter, but the average

difference showed the CTD primary salinity to be higher than bottles by

an average of 0.0002 when a few outliers were excluded. During a

cruise in February 2013 with a lot of deep sampling, the salinity was

found to be within 0.001 of bottle samples.

A pH sensor was attached to the CTD, but the data are not included in the

files to be placed in the IOS Data Library because no field calibration

information was available at the time of processing, and there are some

concerns about how the sensors perform.

For details on the processing see the report: 2013-12-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. There was a lot of variability with values between 70% and 90% in areas of strong vertical mixing such as Haro Strait and Discovery Passage. Values between 110% and 140% were found in the central and northern part of the Strait of Georgia. Values between 90% and 110% were found at a few stations in the southern part of the Strait of Georgia and near the mouth of Juan de Fuca Strait. A few cases of higher saturations were checked to see if the SBE DO near-surface waters were in reasonable agreement with bottles but unfortunately there was no DO sampling at those stations with values >130%. A few casts with saturation >120% were checked and they showed good correspondence between bottles and CTD DO near the surface. The high values are assumed to be due to blooms.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The rest of the steps in this section were run twice– first to produce CHE files, and then a second run to produce CHEpH files that include the pH channel.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and no errors were found.

A header check was run on the final files and it was discovered that the nutrient samples from cast #72 were missing. They had been identified as being from cast #77 which was not a bottle cast. The entry on the rosette sheet did look like #77.  The bottle merge steps had to be repeated for cast #72 after fixing nutrient files.

The CHE file for event #27 was deleted since there was no sampling.

The flag and values changes that were supposed to be made due to a misfire (samples #160 from cast #54) had not been made, so that was fixed.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

General – lot of problems with bottles

15. Bot #11 didn’t trip – trip mechanism replaced.

24. Bottle #13 did not fire

27. Test fire 16 bottles - #13 and #9 did not fire – no sampling

29. Out of order firing. Bottle 20 at same depth as 13 (20m). Bot 20 not sampled.

33. Bot #13 did not fire. Bot #14 fired at surface to substitute.

43. 2 bottles fired at 20m but 2nd not needed, not sampled.

46. 2 bottles fired at surface, but 2nd not needed, not sampled.

54. Blowing 20 knots, no chains bot #11 misfire.

59. 2 bottles fired at 50m, but only 1st sampled.

60. Hard landing for rosette

67. Winch stutter during downcast at 60m

72. Slow descent rate due to weather.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 5Apr2011 | Factory |  |  |
| Conductivity | 3321 | 30Mar2011 | Factory |  |  |
| Secondary Temp. | 4700 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 1766 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 1Apr2011 | Factory |  |  |
| PAR | 4601 | Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | Dec2010 |  |  |  |
| SBE Fluorometer | 2228 | n/a |  |  |  |
| Pressure Sensor | 77511 | 22Apr2011 | Factory |  |  |
| Altimeter | 16504 | 16Mar2011 |  |  |  |