## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 March 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |
| 4 Feb 2014 | Clarification: All data were converted using calibration control file 2013-13-ctd-test47.xmlcon after tests showed the dissolved oxygen sensor serial # entered in the configuration file used at sea was incorrect. GG |

## PROCESSING NOTES

Cruise: 2013-13

Agency: Ocean Sciences Division

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: June 27, 2013 –July 2, 2013

Processed by: Germaine Gatien

Date of Processing: 2 October 2013 -7 January 2014

Number of original HEX files: 85

Number of CTD files: 85	Number of bottle casts:	21 (plus a test cast with no sampling)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#0047), a SeaPoint Fluorometer (#2356) with a 10X cable, a Biospherical QSP-400 PAR sensor (#4601), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (no serial # available).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log was missing an equipment list. The Rosette log sheets and DO analysis logs were in good order.

The Dissolved Oxygen sensor listed in the configuration file (#1438) has the same serial number as the one listed on cruise 2013-43, a W.E. Ricker cruise that overlaps with this cruise. Since there are two sensors with very similar numbers (1438 and 1483), the first guess was that this was the source of confusion. Rosette data were converted first with 1438, then 1483, but neither led to a satisfactory comparison with bottles. Then the CTD technician found some rough notes suggesting a different sensor (#0047) might have been used. Tests run with that configuration looked much better, so the data were processed with that assumption. This incident reinforces the importance of doing a careful visual check of the set-up, entering information in the log book and taking dated photographs at the beginning of each cruise and every time sensors are changed.

Problems were noted in the CTD conductivity data, with sometimes the primary and sometimes the secondary having odd shifts or drifting values while the temperature traces were fine. Some interruption in the flow to the conductivity sensors might cause this and there was a report of many jelly fish in the general area where one such problem was found.  During downcasts, when these shifts were significant for one salinity channel they looked fine in the other. There was slightly more small-scale noise in the secondary channels, so the primary channels were selected for archiving in most cases. However, due to poor data in the primary salinity, the secondary channels were selected for events #3, 28, 56, 66 and 94.

The problems with CTD conductivity were also found during upcasts even while the CTD was stopped. These likely account for some severe outliers in the comparison of CTD salinity to calibration samples. Other minor outliers are possibly due to poor flushing of Niskin bottles. Analysis of the salinity samples was completed within 3 weeks of collection. On average the two salinity channels are well within 0.001 of bottles as has been found when these sensors were used for other cruises in 2013.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6  mL/L from 0 to 75db

±0.2  mL/L from 75 to 225db

±0.15 mL/L from 225 to 300db

±0.1 mL/L below 300db

The SBE fluorescence is higher than the extracted CHL for CHL<1.5ug/L but falls to about 60 % of the CHL for 2<CHL< 4.5ug/L. For CHL>5ug/L the fluorescence is about 75% of the extracted CHL. There is some evidence that inefficient flushing of Niskin bottles in the presence of strong chlorophyll gradients may account for some of this variability. There are significant sub-surface chlorophyll maxima at many sites.

The Surface PAR signal looks odd for many casts, particularly early in the cruise. Sometimes it looks as though there may have been a light shining on it so that there is a fairly steady, but unreasonably high signal even in the middle of the night. These data are always stated to be nominal.

A pH sensor was attached to the CTD, but the data are not included in the files to be placed in the OSD Data Library because no field calibration information was available at the time of processing, and there are some concerns about how the sensors perform. Data including the pH channel were prepared for the use of the Chief Scientist.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen and salinity. There was no equipment list in the Daily Science Log Book.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts. Where missing, the file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. There was confusion about what DO sensor was used, since the same instrument was listed in the configuration file of another cruise that was running at the same time. It is likely that the error was made for the other cruise

The calibration constants were checked for all instruments and no errors were found, though the date of one calibration was incomplete, so that was added. There were no changes through the cruise.

One file was saved as 2013-13-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All files were converted using 2013-13-ctd.xmlcon.

Cast #7 could not be converted at first because 3 lines in the header section were concatenated. When those were fixed in the HDR and HEX files, conversion worked.

A few casts were examined and all expected channels are present. The descent rate was mostly quiet with somewhat more variability at the mouth of Juan de Fuca Strait.

The two temperature and conductivity channels are fairly close during the downcasts and, as usual, much farther apart during the upcasts.

Altimetry looks useful and fluorescence, transmissivity, pH, PAR, SPAR and transmissivity look normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2013-13-ctd.xmlcon.

File #41 will not be processed further because a bottle was fired as a test only and no samples were drawn and no sample number was assigned.

The files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records. There were a few adjustments needed:

- Cast #1 – some bottles fired just as tests had no sampled numbers attached, so those bottles were removed from the list. 29

- Cast #44 – a second bottle was fired in case Niskin 10 misfired, but it did not. Niskin #20 was removed from the list.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-13-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2013-13chl*.xls. The file included comments and flags and an event-number column, but there were no entries in the latter, so those were added. One error was found in sample number and the analyst was alerted to that. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2013-13chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-13oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2013-13oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2013-13SAL.xls. The analysis was done within 3 to 8 days of collection. The files were simplified and saved as 2013-13sal.csv. That file was then converted to individual SAL files. (A typo in 1 sample # was corrected.)

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-13nuts.xls. This includes a precision study.

Event numbers and station names were added. Then the file was simplified, reordered on sample numbers and saved as 2013-13-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts to see what setting brought the vertical offset between the downcast and upcast DO traces into the best agreement with the temperature traces. This sensor has been recalibrated and serviced since it was last used. Tests were run on 3 casts using a variety of settings and the best setting was between 2 and 3s.

ALIGNCTD was used to advance the DO Voltage by 2.5s relative to the pressure.

At this stage some corrections were made to the SeaBird headers. These are described in section 11. The errors were found in the first attempt at processing these data before it was discovered that there was a different DO sensor than indicated in the configuration file.

##### CELLTM

A variety of settings were tested on 4 casts. The results varied somewhat with all choices investigated bringing downcast and upcast traces closer, but given the noisy upcast data it is difficult to choose which is best. A setting of (α = 0.0245, β=9.5) looked reasonable overall for both sensors. This is different from the last time these sensors were used, but at that time the temperature gradients were low overall.

CELLTM was run using (α = 0.0245, β=9.5) for the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 deeper casts to examine differences between sensor pairs. The data from one cast of 2013-12 which used the same equipment 2 months earlier is included for comparison. None of the casts are very deep, but the differences are consistent and small. The differences are similar to the results found during 2013-01, so calibration drift is likely small.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2013-12-0059 | 360 | +0.0001 | -0.00007 | -0.0008 | Noisy, medium |
| 2013-13-0051 | 350 | -0.0002 | -0.00006 | -0.0006 | F.Steady, high |
| 2013-13-0058 | 350 | -0.0001 | -0.00006 | -0.0006 | Steady, high |
| 2013-13-0077 | 340 | -0.0001 | -0.00006 | -0.0006 | Steady, high |

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure shows 4 severe outliers in the primary salinity and 3 in the secondary. The source of trouble appears to be noise in the CTD data in some cases and poor flushing of the bottles in others.

Cast #1 at 75db, both CTD salinity channels differ from bottle by ~0.16. There are values matching the bottle salinity that appear in a shed wake at the beginning of the stop.

Cast #3 at 201db has a high standard deviation in both CTD salinity channels. So this may be natural variability.

Cast # 44 is in central SoG where a good comparison is expected. The upcast CTD data look strange. There is a sudden shift in both salinity traces at the bottom with the secondary furthest from the downcast.

Cast #56 at 300db looks fine in the secondary comparison but the primary salinity is lower than the bottle by 0.04. There is an offset in the primary trace.

A check was made to see if the 4 bottles above look reasonably close to the downcast salinity.

| Cast | Bottle Pressure | Bottle Sal | Upcast Sal0 | Upcast Sal1 | Downcast Sal0 | Downcast Sal1 | Downcast level where Sal0 = Bottle Salinity |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 75 | 30.5875 | 30.4276 | 30.4267 | 34.4404 | 34.4396 | 77.3 |
| 3 | 201 | 31.7568 | 31.7271 | 31.7307 | 34.6734 | 34.6718 | 211 |
| 44 | 301 | 30.9473 | 30.8646 | 30.6392 | 30.9473 | 30.9472 | 301 |
| 56 | 300 | 30.9500 | 30.9103 | 30.9495 | 30.9517 | 31.9508 | 298 |

The cast #1 salinity gradient near 75db is high enough that the difference between the bottle and CTD is not as significant as it might seem. The Saanich Inlet cast is often out of line with others due to a sensor that can’t respond quickly enough to the very high dissolved oxygen gradient.

For cast #3 the bottle is closer to the upcast than downcast as one would expect. The bottle value is the same as that seen 10m deeper in the downcast. This could be due to poor flushing but this is in Haro Strait so local variability is a possibility.

Cast #44 shows good correspondence to the downcast values. This is an area where we expect good comparisons and the upcast data look very odd, so the CTD is most likely the cause of the large differences.

Cast #56 is much like cast #44 with the bottle looking close to downcast salinity and the upcast secondary salinity. There were some obvious problems seen in the upcast primary CTD salinity data.

These observations indicate that we cannot blame the bottle values for the noise in this comparison. The analysis was done quickly (2 to 3 weeks). There are likely two things at work – poor CTD performance on upcasts and inadequate flushing of bottles. The bottle salinity is likely reliable in saying what is in the bottles. All 4 casts had steady descent rates where the mixing of the bottle contents may not do as good a job of reaching ambient conditions.

Plots of the full casts show that the problems with casts #44 and 56 are due to bad conductivity, not temperature. The pumps were on throughout those casts. A few other casts with no bottle firing show similar problems, so it is not firing that causes the trouble. Fluorescence data look odd during casts with offsets in both primary and secondary conductivity. There was a report of jellyfish problems during a net cast close to one of the affected casts, so perhaps there were biological materials disturbing the flow to the sensors. It is harder to see effects on the dissolved oxygen because of hysteresis effects, but for cast #50 and 56 there do appear to be small offsets when the secondary conductivity was affected. Thus far no downcast data has been observed to have a problem.

When a few other outliers are excluded based on high standard deviations in the CTD data, the primary CTD salinity is found to be well within 0.001, but the standard deviation is higher in the primary data. Both fits are reasonably flat with pressure.

Both channels also look quite flat against time with the secondary slightly better than the primary.

For full details for the COMPARE run see file 2013-13-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The first attempt produced unsatisfactory results. There were no extreme outliers but there is a lot of scatter. The range of DO values is small and there are no values < 1mL/L and only one sample with DO < 1.8mL/L, so the slope of the fit of differences against DO value is very sensitive to setting the offset to 0 or leaving it free. If the offset was not set to zero it was very large, ~0.1mL/L. Normally with no low DO sampling it is wise to set the offset to 0, but it makes for a poor fit this time.

Several attempts were made at identifying outliers and some points looked out of line in both plots against pressure and against DO value. A fit of

CTD DO Corrected = CTD DO * 1.0739   (R2 value 0.4112)

When this sensor was used during 2013-12 the fit found was:

CTD DO Corrected = CTD DO * 1.0706 + 0.0418

which is roughly equivalent to 1.079 *CTD DO. When that offset was forced on the 2013-13 data, the slope found was 1.0637 and the fit looks even worse than the one through the origin.

Given the problems found in the salinity comparison due to problems with the upcast CTD data, a study was made to see if the dissolved oxygen data were affected in similar ways. The DO samples from the same casts/levels as the 4 salinity outliers were checked and only 1 is an outlier in the DO comparison. That was from cast #1 in Saanich Inlet, a site where outliers are very common.

There was a concern that the dissolved oxygen sensor mounted might not be the one entered in the configuration file, #1438. Cruise 2013-43 which overlaps with this one is said to have had #1438, so one of these is wrong.  SeaBird number these sensors in pairs with just two digits reversed. For example, the Water Properties group own sensors with IDs #1438 and #1483. There was no entry for the sensor ID in the log for this cruise, though it follows an April cruise for which #1438 was entered in both log and configuration file.

To try to clarify what sensor was used, a separate set of bottle files were prepared using the configuration for DO sensor #1483 which had been recalibrated in March 2013. The fit looks similar to that found above (assuming DO sensor #1438) but the slope is significantly higher. This is unlikely in a newly calibrated sensor. #1438 was calibrated in early 2011 and used many times previous to this cruise. So it seems unlikely that sensor #1483 was in use.

Finally, based on a note about the configuration found by the CTD technician, the data were reconverted using the assumption that the DO sensor was #0047. This produced much better results with an overall fit after exclusion of outliers based on residuals of

CTD DO Corrected = CTD DO * 1.0752  (R2 value 0.7895)

Or if a non-zero offset is allowed the fit was

CTD DO Corrected = CTD DO * 1.0796  - 0.0166

A few casts stood slightly out of line, and when all casts except 60, 69 and 77 were used the fit (excluding some outliers) was

CTD DO Corrected = CTD DO * 1.0688

The fit for casts 60, 69 and 77 (excluding some outliers) was

CTD DO Corrected = CTD DO * 1.0902

When non-zero offsets were allowed there was little difference between the latter two fits. The odd casts were to the west of Texada Island and have a reversal in DO at mid-depths, which may account for the different fit. But with only 3 casts that stand out, there is insufficient data to justify applying different fits for different regions, especially when one considers that the fit may not suit the whole profile nor other nearby casts with no DO sampling. There was a lot of variability in the profiles in that area. With all data included in the fit, there was a higher R2 value at 0.7895 than the 0.6321 and 0.6566 R2 values for the other two fits.

File 2013-13-recal1.ccf to apply the corrections:

CTD DO Corrected = CTD DO * 1.0752

This was applied to the SAM and MRGCLN2 files and COMPARE was then rerun to ensure this correction was appropriate and it was.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no significant outliers were found. No additional flags appear justified.

Note that the DO sensor is likely recorded incorrectly for other summer Vector cruises. However, the bottle comparison for 2013-12 does support the records for that cruise that indicate sensor #1438 was on CTD #0585 in April 2013.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

CTD fluorescence reads higher than extracted CHL for CHL<1.5ug/L, then gradually reduces to about  0.6 times the CHL when CHL is between 2 and 4.5ug/L, but there are a few cases when it is close to 1. Then when CHL is CHL>4.5ug/L, the ratio is about 0.75.

There is no obvious geographic pattern to explain this distribution, but it is possible that variations are affected by how well bottles flush and the magnitude and sign of the local CHL gradient. To see if this might explain some of the lower values, a few profiles were examined in detail.

- Cast #24 – The downcast SBE fluorescence has only a slight decrease with depth between the surface and 20db, (~0.7ug/L) but there is a steep gradient starting at 20db and the fluorescence at 25db is ~0.5ug/L. If there were poor flushing it would have little effect at 2 and 10db, but we might expect a poorer match at 20db. In fact, the ratio of FL/CHL is ~1.1 at all 3 depths.  The effect of the 1.5m offset between bottle and CTD would lead to the fluorescence being slightly lower than that of the bottle sample if flushing were perfect, though with such a low gradient it would not be noticeable.

- Cast #33 – The downcast SBE fluorescence is fairly well-mixed in the top 8db at ~0.5ug/L, then increases fairly sharply to 12db reaching 0.8ug/L. It then decreases slowly reaching 0.68 at 20db. If there was poor flushing we would expect CHL to be too high at 2db and it is almost 2X the fluorescence. At 10db and 20db poor flushing would make it slightly low, but in fact it is slightly higher than the fluorescence. If the flushing were perfect the offset between bottle and CTD would lead to CHL being low at the surface which it is not. It would be slightly higher lower down which it is.

- Cast #65 – The downcast SBE fluorescence starts at about 2ug/L, has a very sharp gradient to about 6.5 at 8db, then falls to 2.2 at 10db and 1.5 at 20db. Poor flushing would have a big effect, though in what direction would depend on how bad the flushing is. But the CHL might be too high at 2db, and it is about twice the fluorescence from either upcast or downcast. At 10m the CHL is slightly low. This would support an explanation of poor flushing. If there were good flushing, then the offset between bottle and CTD would lead to the fluorescence being much higher than the CHL, and that is not the case.

So poor flushing may account for some of the noise in these comparisons. Flushing would vary from bottle to bottle. The offset between bottle and CTD could also account for some differences if flushing were good, and might reduce the effect of poor flushing to some extent.

The majority of cases where fluorescence is lower than extracted CHL are from near-surface samples. So a study was made of the 2m comparisons. For this cruise there are subsurface maxima in fluorescence around 8 to 12m, except in areas of strong mixing where the top 25m are either well-mixed or show a gradual decline with depth. So for the surface samples, small inefficiencies in flushing would likely lead to a FL/CHL ratio >1 in strong mixing areas and <1 where the subsurface maximum exists. If flushing were very poor so the bottle contains water from below the maximum, this would be much more complicated. Looking at the bottle comparison it is clear that the surface samples from events #3, 20, 24 and 27 from Haro Strait and nearby areas have a ratio close to 1, as do a few events from farther north in the strait (#40, 51, 72, 77, 79) where there is either no sub-surface maximum or only a slight one or a very complex profile. Where a significant sub-surface maximum is present the ratio is much smaller. Thus it appears that inefficient flushing is a significant issue in comparing SBE fluorescence and bottles.

A few downcast profiles were examined to see if a vertical flushing offset could be established that would explain the FL/CHL ratio found in the bottle comparison, but the complexity precludes simple answers. Further complications come from temporal changes between the downcast SBE fluorescence which is the “cleanest” data available and chlorophyll samples collected later. A statistical study might produce useful conclusions.

For full details of the comparison see file 2013-13-fl-chl-comp1.xlsx.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  The minimum pressure was 0.66db at the end of a cast. The speeds look fine. The fluorescence appears to be off-scale at the end of cast #55, but the pumps were turned off, this was not a bottle cast and the downcast is unaffected.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.8db which is a little high for the Vector, but not unreasonable

NOTE: The following problems with station names and water depths (those written in italics) were identified during a first attempt at processing these data, before it was discovered that the DO sensor was wrongly entered in the configuration file. So on the second time through, the changes were applied to the SeaBird headers of files after the ALIGN DO step.

The cross-reference check was compared with the log book and two errors were found and corrected in the headers:

- Cast #8 had station number as 12; this should be 102

- Cast #41 was said to be at station 105 and this should be 106.

Depth entries are mostly within 5m of those in the log book, mostly quite close. Since the ship may have drifted between preparing the header and writing in the log, it is not clear which value is more appropriate and this is a rough estimate anyway since the depth when the CTD reached bottom may well be different from either of these values. So changes were not made where the changes were small.

For event #2 the depth was wrong – the value from the previous cast had been entered – so that was changed to match the log book.

For events #5, 65, 78 and 89 the log book entry looks more appropriate so those values were changed in the files.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for about half the casts and the headers look appropriate despite a very noisy signal near the bottom of some casts.

After the second time through Conversion to IOS Headers and CLEAN, the Header check and track plots were rerun and the results were the same.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +55 and +75 records. The results vary from cast to cast and depth to depth. On a few 2012 cruises using this sensor, a setting of +65 records was found best, for 2013-16 a value of +75 records was selected since it generally does a better job near the surface and most casts were very shallow. For 2013-12 +65 looked best.

For this cruise a setting of +70 looks like a reasonable choice.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Conductivity

During 2013-01 the shift settings chosen for the conductivity channels were -0.5 records for the primary and +1.5 records for the secondary, but the CTD configuration was quite different for that cruise. During 2013-16, with similar configuration but very shallow sampling, the choice of -0.5s looked best for both channels, while during 2013-12 the choice made was -0.5 for the primary and -0.3s for the secondary.

Tests were run on 2 of the deeper casts using a variety of shifts and the same results were found as for 2013-12 though the traces were not particularly noisy so the corrections were small.

SHIFT was run using -0.5 records for the primary conductivity and -0.3 records for the secondary.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. At times the upcast vs downcast DO traces seem closer than those for temperature, but often the opposite is true. The upcasts were so different from downcasts that it is very hard to judge the alignment.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late April 2011 and were used for 1 cruise in 2011 with no salinity sampling, during 2013-01 and 2013-12 with good sampling and 2013-16 with poor sampling. Both salinity channels were found to be within 0.001 of bottles, on average, during 2013-01 and 2013-12. The 2013-01 cruise included many deep samples where low salinity gradients ensure that incomplete flushing does not have a large effect on comparisons.

Dissolved Oxygen

The sensor has not been for any other cruises since it was repaired and recalibrated in April 2012.

Pressure

The sensor was recalibrated in April 2011 and was used for 1 cruise in 2011 as well as 2013-01, 2013-16 and 2013-12. The factory offset was used for all those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. In a few cases there was no local climatology available. All temperature and salinity profiles fell within the local climatology in the Strait of Georgia except for event #74 where the temperature was slightly high and salinity slightly low around 25db. There were many excursions from the climatology in Juan de Fuca Strait. Near the mouth the casts to the south and centre of the strait had high temperatures and low salinities in the top 50m (stations 101, 102, 74, 75) and further to the east (stations 68-70) had some slightly high temperatures and low salinities around 100db. These look like real variations rather than indications of instrumental problems. Going a little outside the 3 standard deviation criterion is not unusual so close to shore and in this region.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration - There were no post-cruise calibrations available at the time of processing.

##### CHANNEL CHOICE STUDY

COMPARE does not help in choosing which sensor pair to select for archiving. The two salinity channels are very close to bottle values during stops. T-S plots show only slight differences with sometimes one pair and sometimes the other looking smoother. Examination of the data in CTDEDIT suggests the primary may have less small-scale instability. Since the primary was used for other recent cruises using this equipment, the primary temperature and salinity channels were selected for most casts, but in a few cases the secondary were selected because the primary had sections of poor data.

##### DETAILED EDITING

The primary temperature and salinity channels were edited except for casts #3, 28, 56, 66 & 94 for which the secondary channels were selected.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the cast. In a few cases some bad salinity points were removed and temperature left in place.

All files required some editing.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure does not appear to require recalibration.

From COMPARE we found that both salinity channels were within 0.001 of bottles, on average, but there was a lot of scatter. The T and C sensors had been deployed on 4 previous cruises but only 2 had calibration sampling and in both cases both primary and secondary salinity were within 0.001 of bottles. No salinity recalibration will be applied.

SBE Dissolved Oxygen was recalibrated by applying the following equation:

CTD DO Corrected = CTD DO * 1.0752

CALIBRATE was run using file 2013-13-recal1.ccf to apply the correction to the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles – bottles excluded in the previous comparison were excluded and there was no CTD data available sufficiently close to some bottles. The differences show that the CTD DO is higher than the bottles by an average of 0.05mL/L. The deepest bottles are very close to the CTD while near the surface the CTD reads higher. This is likely explained mostly by poor flushing so that the bottles contain water from lower in the water column than the pressure indicated in the CHE files. This effect is most notable in higher gradients and may be worse when the descent rate is steady, as was the case for most of this cruise. So these differences cannot be ascribed to calibration errors. During 2013-12 the CTD read a little lower and the data were recalibrated.

No recalibration will be applied to these data.

(See 2013-13-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2013-13-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. There were some unstable features in the T-S plots for some casts, but they all occur in areas known for active mixing. The profiles were examined again in CTDEDIT and no further editing was appropriate. There is no evidence that these are due to instrumental problems.

Profile plots were examined on screen and most variables looked normal. However, early in the cruise the Surface PAR values were usually high when the PAR values showed no signal and the time of day shows no signal is expected. After cast #26 the SPAR looks normal at night, so perhaps the odd nighttime values were caused by ship lights. When there was a PAR signal the SPAR usually had much higher values, but this is may be because the CTD data starts below the surface. Where there is PAR data above 2db the values are closer and PAR is sometimes higher than SPAR. As usual, these data are stated to be nominal, and it will be up to users to decide what is reasonable.

##### Final CTD File Steps (REMOVE and HEADEDIT)

The steps in this section were run twice – first to produce CTD files, and then a second run to produce CTDpH files that include the pH:SBE channel.

REMOVE was run on all casts except #3, 28, 56, 66 and 94 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

REMOVE was run on cast #3, 28, 56, 66 and 94 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6  mL/L from 0 to 75db

±0.2  mL/L from 75 to 225db

±0.15  mL/L from 225 to 300db

±0.1 mL/L below 300db

The two CTD salinity channels were generally within 0.001 of each other.

The comparison with bottles had a lot of scatter, but the average

differences showed that both CTD salinity channels were well within

0.001 of bottle values when a few outliers were excluded. During a

cruise in February 2013 with a lot of deep sampling, both salinity

channels were found to be within 0.001 of bottle samples.

A pH sensor was attached to the CTD, but the data are not included in the

files to be placed in the OSD Data Library because no field calibration

information was available at the time of processing, and there are some

concerns about how the sensors perform.

For details on the processing see the report: 2013-13-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. There was a lot of variability with low values in areas of strong vertical mixing such as Haro Strait (50%-70%) and eastern Juan de Fuca Strait and Discovery Passage (70%-90%). Values between 90% and 130% were found at the mouth of Juan de Fuca Strait and in most of the Strait of Georgia, with some northern sites having higher saturations. A few cases of high saturations were checked to see if the SBE DO near-surface waters were in reasonable agreement with bottles but unfortunately there was only one cast with DO sampling at those stations with values >130%. That casts had near-surface SBE DO values a little higher than bottle DO. A few casts with saturation between 110% and 130% were checked and for those the bottles and SBE DO are close overall. The high values are assumed to be due to blooms and fluorescence does tend to be higher in the areas of high DO saturation.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The rest of the steps in this section were run twice– first to produce CHE files, and then a second run to produce CHEpH files that include the pH channel.

REMOVE was run on all casts except #3 and 56 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

REMOVE was run on casts #3 and 56 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with different units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and no errors were found.

A header check was run on the final files. It was discovered that some nutrient samples were missing, so the bottle file preparation was repeated once those were found.

No further errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

1. Test cast all bottles fired but only 1-9 sampled.

4. Cast aborted for ship repositioning

7. pressure on deck 1

40. Bottle 10 misfire – possible interference with sea cable – added extra zap strap

41. Bottle 10 tested – good – no sampling

44. Bottle 20 also fired at 50m in case 10 misfired

55. Jelly fish plugged the net during NET cast just before this CTD cast

64. Winch operator change at170db

72. No O2 sampling

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 5Apr2011 | Factory |  |  |
| Conductivity | 3321 | 30Mar2011 | Factory |  |  |
| Secondary Temp. | 4700 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 1766 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 47 | 1Apr2011 | Factory |  |  |
| PAR | 4601 | Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | Dec2010 |  |  |  |
| SBE Fluorometer | 2356 | n/a |  |  |  |
| Pressure Sensor | 77511 | 22Apr2011 | Factory |  |  |
| Altimeter | 16504 | 16Mar2011 |  |  |  |