## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31 March 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2013-14

Agency: Ocean Sciences Division

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: September 25, 2013 – September 30, 2013

Processed by: Germaine Gatien

Date of Processing: 30 January 2014 - 18 February 2014

Number of original HEX files: 70

Number of CTD files: 70	Number of bottle casts:	21

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#0047), a SeaPoint Fluorometer (#2356) with a 10X cable, a Biospherical QSP-400 PAR sensor (#4601), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (no serial # available).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order except that the wrong dissolved oxygen sensor was entered in the equipment list.

CTD #0585 was used for 3 successive Vector cruises in mid-2013, and it is believed that it was equipped with the same sensors. In processing data from the first of those cruises, 2013-13, it was found that the dissolved oxygen sensor mounted on the CTD was not the one listed in the configuration file. This was discovered because the same sensor was said to be in use on another ship during a cruise that overlapped with 2013-13. The comparison of the CTD Dissolved Oxygen with calibration samples also looked odd. The CTD technician found some rough notes suggesting that sensor #0047 was in use. Tests were run converting data with parameters for several different DO sensors and then comparing the results with bottle samples and this confirmed that #0047 was the sensor in use. Comparisons between bottles and CTD dissolved oxygen for 2013-55 and this cruise also look reasonable when the parameters for sensor #0047 are used. There was no equipment list in the log for 2013-13 and the lists for 2013-55 and 2013-14 had the wrong serial number for the DO sensor.

The PAR sensor listed in the log book does not match that listed in the configuration file. Based on a study of what sensors were used on a series of cruises, it is believed that the log is correct. So the configuration file was changed, but some doubt remains.

These discrepancies reinforce the importance of doing a careful visual check of the set-up and taking dated photographs at the beginning of each cruise and every time sensors are changed.

The salinity analysis was delayed enough that some evaporation of samples is expected. The comparison of CTD with bottles suggests that the primary sensor is low by ~0.0100 and the secondary low by ~0.0036. However, a post-cruise factory calibration shows that the primary is low by ~0.0036 and the secondary is high by ~0.0032. This implies that the bottles are reading too high by an average of 0.006. It is possible that some of that difference is due to poor flushing of Niskin bottles.

The salinity was recalibrated based on a post-cruise calibration. The drift in CTD salinity calibration appears to have been fairly sudden, with the differences between primary and secondary increasing significantly between July and late September.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.25 mL/L from 0 to 50db

±0.15 mL/L from 50 to 125db

±0.1 mL/L from 175 to 200db

±0.02 mL/L below 200db

When some outliers are removed the ratio of SeaBird fluorescence to extracted chlorophyll is about 1.5 for CHL<0.5 and then decreases to about 0.8 at 1ug/L and 0.5 at 2ug/L. The outliers all come from 2m and 10m in the Strait of Georgia and Saanich Inlet. For those the ratio is about 1.5 even when CHL is ~2ug/L.

For event #13 the pumps were not turned on when bottles #1 to #12 were fired, so the pumped CTD channels have been removed. Only the surface bottle has the full array of CTD data. A second drop was made so that a complete CTD file with pumps on is available.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen, extracted chlorophyll and salinity.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. There was confusion about what DO sensor was used, since the same instrument was listed in the configuration file of another cruise that was running at the same time. Tests run for 2013-13 and notes from the CTD technician confirm that the sensor was #0047 and it was assumed to be the sensor used for this cruise as well. This will be checked when the dissolved oxygen bottle data and CTD data are compared.

Another discrepancy was found between the log and configuration files. The PAR sensor is said to be S/N 4601 in the log but S/N 4615 in the configuration file. The sensor used in the configuration file for 2013-13 was 4601 but there was no list in the log. For 2013-55 PAR S/N 4601 is entered in the log and the configuration file. While it is possible that it was changed before this cruise, the technician could not remember doing that.

Tests were run on a few daylight casts to see if the PAR data makes more sense with 4601 or 4615 parameters. The 4615 parameters lead to higher PAR values. For 3 casts with fairly high SPAR values and data available at about 2db, the PAR value is higher than SPAR for 2 of them and lower for 1. The 4601 values are closest to SPAR for 2 and the 4615 closest for one. This is too random to be useful. When light levels are very low, once again the 4601 readings are slightly lower than 4615 with both lower than SPAR. So these tests might suggest #4601 is the better choice, but the results are far from convincing. The PAR may be shaded by the ship near the surface and none of these sensors have been calibrated since 2011, so a good match is not expected.

Next, checks were made of other cruises to see if PAR #4601 or #4615 could be found in use elsewhere in September. The history shows sensor #4601 was mounted on CTD #0585 between February and early May 2013 and moved between the Tully and Vector, as needed. Most sensors were then moved to CTD #0443 which was used for May/June cruises on the Tully. The sensors were moved back to CTD#0585 and were used for Vector cruises in late June, July and for this cruise in late September. In late June a Tully cruise, 2013-50 went out at the same time as a Vector cruise, so CTD #0443 was used for the Tully with sensors that had not been used on any of the earlier cruises, including PAR #4615. Cruise 2013-13 went out at the same time with CTD #0585 and PAR #4601 and the same equipment was used for 2013-55 in July. CTD #0443 went out on the Tully in August and September with PAR #4615. The Tully was back before this cruise left in late September, so either CTD was available, but CTD #0585 was used. So it is most likely PAR #4601 was mounted. There is no obvious reason why it would have been switched as there were no other cruises in September that required a PAR sensor and no reports of problems in sensor #4601.

So it will be assumed that the log is correct and the configuration file is wrong.

The calibration constants were checked for all instruments and no errors were found, though the date of one calibration was incomplete, so that was added. There were no changes through the cruise.

One file was saved as 2013-14-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All files were converted using 2013-14-ctd.xmlcon.

A few casts were examined and all expected channels are present. The descent rate was sometimes very quiet but it was very noisy with some complete reversals at the mouth of Juan de Fuca Strait and during some casts late in the cruise. The temperature channels were close during downcasts while the conductivity appears to be a little further apart than usual. The upcast differences were larger and more variable.

The altimetry looks useful and fluorescence, transmissivity, pH, PAR and SPAR look normal. The fluorescence dark value is <0.1ug/L.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2013-14-ctd.xmlcon. One file (#61) contains almost no data and is not listed as a bottle file in the log and there is no rosette sheet. This file will not be processed.

The files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. There was some noise in the primary salinity in cast #26 which was cleaned using CTDEDIT. The output was copied to BOT.

Note: Cast #13 is extremely noisy because the bottles were fired when the pumps were off during the first drop (except for the surface bottle). A CTD file can be prepared based on the 2nd drop of the CTD, but the bottles were fired on the first drop. So the pumped channel values will need to be replaced in the CHE file for bottles #1 to #12. Comparisons were examined and the differences between downcast and upcast are much too large to justify using the downcast data to make estimates of upcast pumped CTD channels.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2013-14-bot-hdr.txt; it may need further editing to reflect problems found during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2013-14chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2013-14chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-14oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2013-14oxy.csv.

An error was found in event #s – bottles were from event #7 not #6.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2013-14SAL.xls. The analysis was done within 45-51 days of collection. The files were simplified and saved as 2013-14sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-14nuts.xls. This includes a precision study.

Event numbers and station names were added. Then the file was simplified, reordered on sample numbers and saved as 2013-14-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts to see what setting brought the vertical offset between the downcast and upcast DO traces into the best agreement with the temperature traces. Tests were run on 3 casts using a variety of settings and the best setting was between 3.5s and 4s. The upcast temperature is noisy making a judgment difficult, but +3.7s looks like the best choice overall. This is higher than found for 2013-13 in June/July when the advance used was +2.5s.

ALIGNCTD was used to advance the DO Voltage by 3.7s relative to the pressure.

##### CELLTM

A variety of settings were tested on 4 casts. The results varied somewhat with all choices investigated bringing downcast and upcast traces closer in T-S space, but given the noisy upcast data it is difficult to choose which is best. Overall a setting of (α = 0.03, β=7) looked best for the primary conductivity while (α = 0.02, β=9) was better for the secondary conductivity.

CELLTM was run using (α = 0.03, β=7) and (α = 0.02, β=9) for the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. The data from one cast each from 2013-12 (April) and 2013-13 (June/July) which used the same equipment 5 months and 2 months earlier are included for comparison. There was a notable difference between the differences for 2013-13 and 2013-14 (September), so a preliminary processing of the deepest cast from 2013-55 (July) was done for further comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2013-12-0059 | 360 | +0.0001 | -0.00007 | -0.0008 | Medium, noisy |
| 2013-13-0051 | 350 | -0.0002 | -0.00006 | -0.0006 | F.Steady, high |
| 2013-55-0139 | 260 | +0.0004 | -0.00015 | -0.0018 | Steady, low |
| 2013-14-0044 | 345 | +0.00005 | +0.0006 | +0.0060 | F.Steady, high |
| 2013-14-0052 | 345 | +0.0001 | +0.0006 | +0.0062 | Steady, high |
| 2013-14-0074 | 310 | -0.0001 | +0.0006 | +0.0063 | F.Steady, high |

Temperature differences are small for all 4 cruises. Conductivity and salinity differences were small for the first 2 cruises and then became slightly larger for 2013-55 and then much larger for 2013-14. But the changes are not monotonic with the change from 2013-13 to 2013-55 being of opposite sign to the change from either of those to 2013-14. This may be reading too much into the shallow observations which are all that are available from 2013-55. However, it does suggest that there was a fairly dramatic change between late in 2013-55 and the middle of 2013-14. For 2013-14 the conductivity seems noisier than usual. Plots of conductivity itself don’t make it clear that one sensor is worse than the other. Conductivity calibrations were checked again and no error was found. Salinity comparisons are the only way to resolve which sensors are more reliable, but the analysis was delayed and there are no really deep casts or areas of very low salinity gradients, so this may not be as helpful as we would like. A few checks show the bottle values are higher than both salinity channels, and closer to the secondary.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure shows 1 severe outlier in the primary salinity and 2 in the secondary. One other bottle is a clear outlier in both. All outliers appear to be due to noise in the CTD data. When those outliers are excluded the fits are noisy but fairly flat. The primary salinity is low by an average of 0.0100 and the secondary is low by 0.0036. So the secondary is higher than the primary by 0.0064 which is close to the difference found in the previous section. There was a sufficiently long wait before analysis that we might expect some evaporation of sample, which would make the CTD salinity look lower than it really is. That effect is expected to be random. There is also a possibility that Niskin bottles did not flush fully which would also lead to CTD salinity looking relative low, but many of these casts had a noisy descent rate which is likely to encourage flushing. There is also no suggestion of depth dependence in the differences except that there is more scatter in shallower waters. The noise in the comparison is greatest early in the cruise, but that corresponds to some of the shallowest casts and noisiest descent rates so the CTD data is noisier.  A plot of differences against salinity shows no obvious trend, so non-linearity in the salinometer is not a factor.

Some preliminary results of a post-cruise calibration became available later, and will be discussed in section 14.

For full details for the COMPARE run see file 2013-14-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

SeaBird recommend doing a fit with the offset set to 0. This does not always seem to produce the best fits, but for these data there are no DO values <1.5mL/L, so it is wise to force the offset to 0.

Excluding differences that were clearly out of line and then others based on residuals the following fit was found:

CTD DO Corrected = CTD DO * 1.084   (R2 value 0.96)

The outliers come from casts #1 in Saanich Inlet which is always challenging for the DO sensor due to steep gradients, cast #13 for which the pumps were not turned on and 1 near-surface bottle from cast #44 for which the CTD data were very noisy.

There is some confusion about which DO sensor was used for a series of cruises on the Vector on CTD #0585. Tests run for 2013-13 led to the conclusion that sensor #0047 was used. When other sensor calibrations were used the bottle comparisons looked unusual. Bottle comparisons for this cruise and for 2013-55 look reasonable when the sensor is assumed to be #0047.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no significant outliers were found. No additional flags appear justified.

Fluorescence

A first run of COMPARE turned up some errors:

For casts #51 and 54 the rosette log indicates that, as usual, there were samples from 20m, 10m and 2m.

For cast #51 the samples should be 192, 193, 195.

For cast #54 the samples should be 204, 205, 207 .

But the spreadsheet has sample numbers 194, 195, 197, 204, 205, 207.

This leads to a CHL sample at 150m of cast #54 which is obviously wrong.

Cast #51 has values entered for 5m and 2m which are the depths that correspond to the sample #s but not what we expect.

It looks like samples 192, 193 and 195 for some reason got recorded as 194, 195 and 197 (if someone thought the 20m sample was 194, the other numbers would follow logically). Judging by the fluorescence, the values make sense as 192, 193 and 195.  Those sample numbers were changed in the CHL file and the merge process repeated.

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. In Juan de Fuca Strait the Fluorometer read a little lower than the extracted CHL and in Saanich Inlet and the Strait of Georgia it read higher. For most of the samples the ratio of fluorescence / Extracted CHL is about 1.5 for low CHL and gradually reduces to about 0.5 when CHL is ~2.5ug/L. However, there are outliers for which the fluorescence is higher than the extracted CHL for CHL>1.5ug/L. All the outliers come from 2m and 10m between the southern Strait of Georgia and the northern tip of Texada Island. These outliers are likely due to CDOM from the Fraser River.

For full details of the comparison see file 2013-14-fl-chl-comp1.xlsx.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

A text editor was used to remove the first CTD drop of cast #13 because the pumps were not on. The CTD was returned to the surface and the cast rerun with the pumps on. Note that the bottles were fired during the first drop, so the CTD channels will not be useful in comparisons, with the exception of the surface bottle for which the pumps had been turned on.

##### Checking Headers

The header check was run.  No obvious errors were found. The minimum pressure was 0.4db at the end of a cast with pumps on and salinity values were fairly low but the CTD was clearly in water. The speeds look fine. The fluorescence maximum was high during cast #13, but that was a spike and the pumps were not on at the time.

A cross-reference list was checked against the log book and two station names (casts #9 and 10) were found to be wrong. They were corrected in the IOS and CLN files and in the bottle files for cast #10.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.6db which is reasonable for the Vector.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for about half the casts and the headers look appropriate despite a very noisy signal near the bottom of some casts. However, in some cases there was no entry due to such noise and an estimate can be made by examining the plots of altimetry at the bottom. Those estimates were entered into the headers of the CLN files and SAMAVG files.

There were also some water depths that did not match the entries in the log; this is most likely due to ship drift between the times when the header and log entries were made; in most cases the header looks best based on maximum depth sampled and corresponding altimetry with the following exceptions:

- For cast #4 the water depth from the previous station was repeated in the header, so the entry was changed to match the log.

- For cast #56 the log entry looks better, so a change was made for that cast.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Conductivity

During 2013-13 the shift settings chosen for the conductivity channels were -0.5 records for the primary and -0.3 records for the secondary. Tests were run on 4 casts using a variety of shifts and the best results were found with -0.5 records for both channels.

SHIFT was run using -0.5 records for the primary and secondary conductivity.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +60 and +80 since the best results during 2013-13 was +70 records. That setting looks right for these data as well.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The alignment looks good.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late April 2011 and were used for 8 other cruises since then. Many of those cruises had either no calibration sampling or very limited sampling. Both salinity channels were found to be within 0.001 of bottles, on average, during 2013-01, 2013-12 and 2013-13. The 2013-01 cruise included many deep samples where low salinity gradients ensure that incomplete flushing does not have a large effect on comparisons. For 2013-12 and 2013-13 analysis was prompt thus avoiding any problems with evaporation of samples. For 2013-55 salinity analysis was done quickly but sampling was shallow and likely affecting by poor flushing of Niskin bottles. The deepest bottles suggested that both channels were within 0.001 of bottles but the evidence was weak.

Dissolved Oxygen

The sensor was calibrated at the factory in April 2012 and this is the 3rd use since then. The linear recalibration for 2013-13 had a slope of 1.075; the offset was forced to 0. For 2013-55 a linear fit thorough the origin had a slope of 1.067, but there were some doubts about the fit as poor flushing of Niskin bottles was likely a problem.

Pressure

The sensor was recalibrated in April 2011 and has been used for 8 other cruises since then. The factory offset was used for all those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. In a few cases there was no local climatology available. All temperature and salinity profiles fell within the local climatology except:

- Temperatures were a little high and salinity a little low around 20 to 30m at the southern station at the mouth of Juan de Fuca Strait.

- Temperatures were a little high and salinity a little low around 30 to 40m just to the east of Texada and Cortes Islands. Further to the east of those islands temperatures were within the 3-standard deviation limits.

- Salinity was well above the maximum from 50m down at cast #31. This position is close to where the climatology used changes from that for the South Strait of Georgia to the Gulf Islands file. If this cast were plotted with the Gulf Islands data it looks fine.

None of these outliers look significant. The 3-standard deviation climatology is too severe for this region especially close to shore.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – Preliminary data from a post-cruise calibration became available late in processing. A few casts were converted using the new conductivity parameters to try to determine which sensors are the more appropriate choice. The results are summarized in the following table:

| CTD #585 |  | primary |  | secondary |  | primary | Secondary |
| --- | --- | --- | --- | --- | --- | --- | --- |
| cast | pressure | old cal | new cal | old cal | new cal | old-new | old-new |
| 7 | 215 | 33.8190 | 33.8230 | 33.8257 | 33.8222 | -0.0040 | 0.0035 |
| 7 | 2.6 | 31.9304 | 31.9340 | 31.9347 | 31.9313 | -0.0036 | 0.0034 |
| 10 | 254 | 33.8255 | 33.8296 | 33.8324 | 33.8289 | -0.0041 | 0.0035 |
| 10 | 2.7 | 31.7091 | 31.7127 | 31.7129 | 31.7096 | -0.0036 | 0.0033 |
| 48 | 340 | 31.2484 | 31.2521 | 31.2540 | 31.2508 | -0.0037 | 0.0032 |
| 48 | 3 | 27.0625 | 27.0657 | 27.0641 | 27.0613 | -0.0032 | 0.0028 |
| 72 | 273 | 31.0577 | 31.0614 | 31.0632 | 31.0600 | -0.0037 | 0.0032 |
| 72 | 2.9 | 28.5000 | 28.5035 | 28.5049 | 28.5019 | -0.0035 | 0.0030 |
| 75 | 319 | 31.1444 | 31.1481 | 31.1510 | 31.1477 | -0.0037 | 0.0033 |
| 75 | 2 | 28.4993 | 28.5027 | 28.5005 | 28.4975 | -0.0034 | 0.0030 |
|  |  |  |  | Average |  | -0.0036 | 0.0032 |
|  |  |  |  | Average | Deep | -0.0039 | 0.0034 |
|  |  |  |  | Average | Shallow | -0.0035 | 0.0032 |

So the primary salinity was too low by about 0.0036 when it arrived back at the factory and the secondary was too high by 0.0032. While the differences are slightly higher at depth, this does not look like a strong conclusion given that the deepest cast was close to the average. A constant correction looks appropriate and there is no evidence to favour either primary or secondary sensors. While there has obviously been significant drift since 2013-13 there is no evidence of significant drift through this cruise.

##### CHANNEL CHOICE STUDY

COMPARE would suggest that the primary sensors are closer to the bottles, but since both need correction this is not important. The bottles are not fully trusted due to delayed analysis and poor Niskin flushing may be a factor. It is clear from section 8 that the bottles differ by ~0.006. The post-cruise factory calibration shows that the two channels differed by about 0.0068. The agreement between the channel differences in September 2013 and January 2014 means that we should be able to rely on the post-cruise calibrations as a guide to recalibration. Either sensor pair could be recalibrated. Both salinity channels show slight pressure dependence in calibration drift, though it is slightly lower in the secondary.

A few casts were opened in CTD edit to see whether there were significant differences that would affect data quality and editing effort required. The primary T-S plots have fewer unstable features so that less editing would be required.

##### DETAILED EDITING

The primary temperature and salinity channels were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. In a few cases some bad salinity points were removed and temperature left in place.

All files required some editing.

All ED1 files were copied to EDT.

NOTE: The DELETE program was run on the wrong files so that the alignment of the pH sensor was not included. So the files with pH shifted were put through DELETE. MERGE was then run choosing all but the pH channel from the EDT files and only pH from the delpH files. This ensures pH is removed for records that were removed in editing. The output files were named EDTMRG.

##### Initial Recalibration

The pressure does not appear to require recalibration.

From COMPARE we found that both salinity channels were lower than bottles, but the results do not agree with the post-cruise calibration. This is not unexpected because there was some delay in salinity analysis which has been found to lead to bottle salinity values that are too high, probably due to evaporation of samples. There may also be errors due to poor flushing of Niskin bottles which will also lead to salinity sample values that are too high. The post-cruise calibration indicates that the primary salinity is low by about 0.0036. This is the best information available. No other cruise has been found that used these sensors between 2013-14 and the factory calibration.

CALIBRATE was run using file 2013-14-recal1.ccf to apply these two corrections to the SAM and MRGCLN2 files:

Salinity:T0:C0 Corrected =   Salinity:T0:C0 + 0.0036.

CTD DO Corrected = CTD DO * 1.084

COMPARE was rerun for salinity and dissolved oxygen using the recalibrated values. The results confirm that the recalibrations were done properly.

CALIBRATE was then run on the EDTMRG files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles – there were no CTD data available sufficiently close to some bottles, mostly at the surface. The differences show that the CTD DO is higher than the bottles by an average of about 0.02mL/L, but the results vary greatly depending on how outliers are identified. As usual, the cast in Saanich Inlet is out of line with other casts. This is presumed to be due to a very steady descent rate in the presence of a high gradient which likely leads to poor flushing of Niskin bottles and there may be problems with the response of the DO sensor.

There is a little pressure dependence in the differences which may be due to forcing the trend line through the origin to determine the correction applied. Looking at the plot of differences against pressure also shows that the CTD DO values tend to be relatively high in the regions where the DO gradient was highest, typically around 10-40m and around 100-125m. The fact that the CTD DO reads a little high could be due to incomplete Niskin flushing. However, the differences are small which suggests that flushing of bottles is not a major problem except in the high DO gradient areas. The fits against pressure are not very convincing because the variable gradient means a linear fit is not appropriate. Nonetheless, it is useful for making a rough estimate of the likely errors in the CTD DO data in different pressure ranges.

No further recalibration will be applied to these data.

See 2013-14-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2013-14-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots look fine. Profiles mostly look fine though there are a few deep spikes in transmissivity and fluorescence. No editing was applied as these features are at sites with fairly strong mixing.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH, PAR and Reference PAR data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.25 mL/L from 0 to 50db

±0.15 mL/L from 50 to 125db

±0.1 mL/L from 175 to 200db

±0.02 mL/L below 200db

Salinity data were recalibrated based on a post-cruise factory calibration.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2013-14-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values <70% were all found in areas known for strong vertical mixing such as Haro Strait, the eastern part of Juan de Fuca Strait and 1 cast near Discovery Passage. In Juan de Fuca Strait the saturation rises from 50-70% in the east to 90% to 110% at the mouth. In most of the Strait of Georgia saturation is 90-110% but is between 70 and 90% for some of the casts in the central part. Weather conditions varied a lot through the cruise, so mixing probably explains the low values. These values do not suggest any problem with calibration.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A text editor was used to enter pad values in the following channels except for bottle #13 at the surface: Temperature:Primary, Salinity:T0:C0, Fluorescence:URU:Seapoint and Oxygen:Dissolved:SBE.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, fix the chief scientists name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check was run on all files and no errors were found.

A header check was run on the final files.

No further errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

11. Chains on.

13. CTD lowered with pumps off, returned to surface and repeated with pumps on. Bottles fired on first cycle, so pumps off except for surface bottle.

32/33. Suspension of science program due to engineering problem – 7 hours between casts.

51. Chains on.

61. Rosette file produced but no bottles from that cast and no data in the file.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 5Apr2011 | Factory |  |  |
| Conductivity | 3321 | 30Mar2011 | Factory |  |  |
| Secondary Temp. | 4700 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 1766 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 47 | 1Apr2011 | Factory |  |  |
| PAR | 4601 | Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | Dec2010 |  |  |  |
| SBE Fluorometer | 2356 | n/a |  |  |  |
| Pressure Sensor | 77511 | 22Apr2011 | Factory |  |  |
| Altimeter | 16504 | 16Mar2011 |  |  |  |