## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2013-16

Agency: MEAD

Location: Johnstone Strait

Project: Aquaculture Shellfish Carrying Capacity

Party Chief: Sutherland T.

Platform: Vector

Date: April 2, 2013 –April 10, 2013

Processed by: Germaine Gatien

Date of Processing: 14 August 2013 - 27 August 2013

Number of original HEX files: 89

Number of CTD files: 89	Number of bottle casts:	40

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1438), a SeaPoint Fluorometer (#2228) with a 3X cable,), a Biospherical QSP-400 PAR sensor (#4601), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (no serial # available).

The deck unit was a Seabird model 11+, serial #0425. The logging computer was #2.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list, but the altimeter serial number was missing from the log and from the configuration file. There was no personnel list.

There were no rosette log sheets; none were available on board.

There were Dissolved Oxygen analysis log sheets.

There were no notes in the log detailing what was an unusual nutrient sampling protocol. There were samples with IDs like 19 and 19a. The values made it clear that these were not duplicates from a single Niskin. Fortunately, personnel were available to explain that the samples with an “a” after the number came from a bucket sample a few metres above the surface Niskin. The bucket samples have no CTD data to go with them since no bottle was fired there; the CTD was usually turned off when it was within 1.5 to 2m of the surface and the near-surface temperature and salinity gradients are likely high enough that the 2m CTD data are not representative of the waters from which the bucket samples came. The bucket samples are not included in the CHE files, but may be found in the nutrient spreadsheet prepared by the analyst.

There are other mysteries in the nutrient data. Two samples labelled as 128-2 and 168-2 do not have values close to samples 128, 128a, 168 or 168a. They do not look like surface samples. The source is unknown. They were not included in bottle files, but may be found in the nutrient spreadsheet.

The event numbers were wrong in the salinity analysis spreadsheet.

In all but one case the salinity samples came from the bottom of casts. Even that one exception was only 4m off the bottom. It is understood that this cruise was mostly in shallow water and focused on near-bottom and surface samples, so there were limited choices. But it would be better to sample about 10m off the bottom as the samples may be contaminated by sediments, there may be bottom currents that lead to significant differences over the 1.5m offset between the bottle and CTD and the flushing may not be as efficient close to the bottom.

The CTD salinity was higher than the bottles in every case, which is likely explained by poor flushing of bottles. Normally we expect poor flushing to lead to CTD salinity looking low, but with bottom bottles the effect is the opposite. The comparison of bottles with downcast CTD profiles did not turn up any simple patterns such as a typical vertical offset.

There was reasonably good correspondence between the two CTD salinity channels which suggests that there has not been significant calibration drift since the last factory service.

A pH sensor was attached to the CTD, but the data are not included in the files to be placed in the IOS Data Library because no field calibration information was available at the time of processing, and there are some concerns about how the sensors perform. Data including the pH channel were prepared for the use of the Chief Scientist.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8 mL/L from 0 to 35db

±0.3 mL/L from 35 to 150db

±0.08 mL/L below 150db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained as well as analysis sheets for dissolved oxygen and salinity. No CTD equipment problems were noted other than some difficulty with pressure at the dock. That was resolved before any CTD casts. There was no personnel list, but there was a cruise report with that information.

There were no rosette log sheets available at sea; some notes were made on the dissolved oxygen analysis log sheets about salinity sampling. There are no notes about nutrient sampling.

The nutrient analyst provided a spreadsheet to assist with figuring out the sampling. This was especially helpful since the assignment of sample numbers was non-standard.

The files from cast #162 were misnamed in acquisition; the names were fixed in the raw files.

Nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained

The calibration constants were checked for all instruments and no errors were found. There were no changes through the cruise.

One file was saved as 2013-16-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All files were converted using 2013-16-ctd.xmlcon.

After conversion 3 station names were corrected based on log notes.

A few casts were examined and all expected channels are present. The descent rate is often low, but is generally >0.3m/s between 10m and bottom -10m, so the usual DELETE cut-off should be ok.

The two temperature and conductivity channels are fairly close during the downcasts and, as usual, much farther apart during the upcasts. Even during stops there is a lot of noise. However, the casts are quite shallow and in an area where mixing may be significant, so these variations may be real. The dissolved oxygen and pH have similar patterns due to slower response times. .

Altimetry is sometimes spiky, but there appears to be a useful signal at the bottom, and fluorescence, PAR, SPAR and transmissivity look normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2013-16-ctd.xmlcon.

There were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. A few outliers were found in the salinity channels of events #133 and #157. CTDEDIT was used to remove the outliers.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. The dissolved oxygen analysis sheets were used to determine the correspondence between sample numbers and Niskin numbers for each cast. Sample #181 was used twice, so 8181 was used for the case during event #146 (there was no DO or SAL sampling for that one, just nutrients).

At this stage it was realized that there is a problem with the nutrient sampling. There are many samples that duplicate other sample numbers, but with an “a” attached. The values show that the 2 samples do not come from the same bottle. In the absence of rosette sheets and no note in the Daily Science Log book, this caused a halt in work until cruise participants could be found to explain what happened. It seemed likely that the extra bottle was from some separate system to get near-surface samples. It was confirmed by the CTD technician that the “a” samples came from a surface bucket sample.

The “a” samples will not be included in the CHE files since there is no CTD data to go with them. The CTD generally stopped at least 1.5m below the surface, and the gradients are likely high enough that the data would not be representative of the waters from which the bucket samples came.

There is one other case where a non-standard sample number had to be assigned. Sample #181 was accidentally used on 2 different casts, so for the sample originally named #181 during cast #146 the number was changed to 8181. This may cause a few problems with the merging process, but only affects nutrients since there was no DO or salinity sampling.

Sample #s were added to the ADDSAMP file based on the dissolved oxygen analysis sheets.

The addsamp.csv file was then converted to CST files. The CST files will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Because file #146 has sample numbers that are not increasing consecutively, the file was edited to put the last bottle at the top. This will enable merging based on sample numbers.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-16-bot-hdr.txt; it may need further editing to reflect problems found during processing.

Dates of creation were added to the names of spreadsheets from analysts.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-16oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary rows and columns and the file was then saved as 2013-16oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2013-16SAL.xls. The analysis was done within 8 to 16 days of collection. The files were simplified and saved as 2013-16sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-16nuts.xls. There was no report on precisions because there were no duplicates.  The file was simplified, reordered on sample numbers, the bucket samples were removed and the file was saved as 2013-16-nuts.csv. Two other records were removed because the source could not be found. They were labelled 128-2 and 168-2 and the analyst had changed those names to 99128 and 99168. They look like deep samples, whereas 128, 128a, 168 and 168a are near-surface samples. The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files. Some manipulation was needed for event #146 to make this work.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests are difficult with these data since casts are shallow and variability is high, so the dissolved oxygen sensor can’t keep up well. The last time this sensor was used during 2013-01 tests found that an advance of +4.5s was best, on average, though results were better in places with either +4s or +5s. Testing on one of the deeper casts from this cruise was inconclusive, but 4.5s was reasonable.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

As for the DO alignment step, tests were inconclusive for these data. So again the results of 2013-01 were used.

CELLTM was run using (α = 0.02, β=9) for the primary and (α = 0.03, β=9) for the secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

(The usual run to study differences between channels was skipped since the casts are too shallow for meaningful results.)

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure indicates that the CTD is higher than the bottles in every case. The differences are, very roughly, inversely proportional to pressure. There are a number of factors other than calibration drift that may affect the comparison:

- The data are mostly from above 100m in areas of high salinity gradients and temporal variations, so the vertical offset between the Niskin bottle and the CTD is significant. The CTD is liable to read higher than the sample especially in the shallower waters.

- The descent rate of the CTD indicates that the CTD moved smoothly through the waters which likely leads to poorer flushing than in open waters or in windy conditions.

- The samples came from the bottom of each cast with one exception. If flushing is incomplete the bottle will contain water from higher in the water column. This is different from most cruises when the samples are taken at least 10m above the bottom to that poor flushing leads to CTD salinity being lower than bottles.

- Samples taken from the bottom may be contaminated by sediment.

- CTD values near the bottom may be affected by bottom currents while the sample may come from above the current.

- The only cast that was not sampled at the bottom of the cast was from only 2m above the bottom, so there was likely little flushing.

The difference between the two CTD salinity channels during bottle stops was <0.001. There were only 3 cases with differences were >0.002. The sensors have been used on only 2 cruises since they were last calibrated and during their use in 2013-01 (with lots of deep open-water sampling) both produced average salinity that was within 0.001 of bottles. So there has likely been no significant drift in calibration.

Plots of differences against event number show no obvious patterns.

File 2013-16-bottle-study.xlsx contains data that were examined to see if there were patterns in the differences related to local gradients, areas or pressure. No obvious patterns were found.

This comparison is not suitable for judging the salinity accuracy and no attempt has been made to identify outliers in the salinity analysis.

For full details for the COMPARE run see file 2013-16-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. There are many cases of the CTD DO being higher than the titrated samples; drift in DO calibration usually leads to low values, not high. Almost all of these cases are from 2m or 10m, levels of very high DO gradients where small mismatches due to vertical offset between bottle and CTD or incomplete flushing of the bottles may lead to large errors. Also the DO sensor often has a peak just below the surface. As well as the high DO values, there are also some relatively low DO values at the surface.

Various attempts were made to remove outliers based on different criteria. In all cases it looks best to set the origin to 0, which is recommended by SeaBird but does not always look best. But for these data there are no DO values <2 and some constraint needs to be placed on the offset. The trendline with offset = 0 has a slope of 1.07.  When this sensor was last used, during 2013-01, there were many data with DO<2mL/L, so an offset was allowed; the fit was slope=1.0664 and offset=0.0189. Had the offset been set to 0 the slope would have been 1.0694 which is very close to the result from this cruise.

So recalibration will be applied as follows:

CTD DO Corrected = CTD DO Raw* 1.07

Flags –There is too much scatter in the comparison to justify either adding or adjusting quality flags. Some of the outliers could be due to analysis problems but there are just as likely to be due to high vertical and temporal DO gradients or incomplete flushing.

For full details of the comparisons see files 2013-16-dox-comp1.xls.

Fluorescence

There are no extracted chlorophyll data

Plots of Titrated DO and CTD DO against CTD salinity were examined and significant outliers were in high gradient areas. No additional flags appear justified.

##### Test Plots and Channel Check

There are no casts deep enough to use for study of differences between sensors.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  The minimum pressure was 1.08db. The speeds look fine and the fluorescence does not go off-scale.

Surface check was run and shows an average surface pressure for the cruise was 2.6db which is a little high for the Vector, but not unreasonable. Drift in pressure is generally towards low values, not high. Plotting a few casts shows salinity dropping off sharply at about 2db, but it is still high enough that the CTD was definitely in the water. There were no pressure values <1db so we do not have sufficient evidence to consider recalibration. There was a check of pressure on the deck, but we have no experience with how to use that check. This is the first in a series of cruises using this CTD in 2013; if processing of the other cruises indicate that the pressure needs recalibrating, that can be done later.

The cross-reference check was compared with the log book and a few errors were identified:

- The station name was wrong for cast 14; that was fixed in the IOS and CLN files. It was not a rosette cast.

- The cast at MAL 40 had been saved as 2013-16-0074 but should have been 2013-16-0083. The file name was changed in the IOS and CLN files and the event number was fixed in the CLN file. This was a rosette cast, so the merges had to be redone since the samples had been saved with the correct event number. The relevant files were corrected and the merges rerun. The cast lists were adjusted and the COMPARE routine was rerun for dissolved oxygen. Salinity was not sampled at that station so the salinity comparison is ok.

- Event 202 had a NMEA latitude of 49 34.49 N but the log had 49 33.51 N. Both positions are reasonable, but the NMEA entry is less likely to have an error. The log entry is assumed to include a typo with 33’ instead of 34’. The file was not changed.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom and the altimetry headers look appropriate despite some very noisy altimetry near the bottom of many casts. In cases where the CTD did not get within 15m of the bottom, no header was created. A few bottom depth entries were checked and were close to the log entries.

##### Shift

Fluorescence

Plots were examined to estimate what SHIFT value should be used to make the offset between the downcast and upcast fluorescence traces look like that of the temperature trace. The results are difficult to judge because the upcast temperature is noisy but an advance of +24 records looks reasonable for the SeaPoint fluorometer and that setting is the one generally found appropriate. The upcast data are very noisy for both sensors, so these tests are not easy to interpret.

SHIFT was run twice on all casts to advance the SeaPoint fluorescence by +24. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +55 and +75 records. The results vary from cast to cast and depth to depth. On a few 2012 cruises using this sensor, a setting of +65 records was found best, but since many casts are shallow a value of +75 records was selected since it generally does a better job near the surface.

Conductivity

These casts are rather shallow to apply the usual test for alignment of the two conductivity channels. The values used when the sensors were last used were -0.5 records for the primary and +1.5 records for the secondary, but the CTD configuration was quite different for that cruise. Tests were run on 2 of the deeper casts using those shifts. The primary choice of -0.5 records proved to be a good choice, but using +1.5 records for the secondary produced bad results. Tests indicate that -0.5 records is the best choice for the secondary.

SHIFT was run twice on all casts using -0.5s for both conductivity channels.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but this does not appear necessary. At times the DO seems closer but since there are real variations in temperature between downcast and upcast, this test is not very reliable.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late March 2011 and were used for 1 cruise in 2011 with no salinity sampling and during 2013-01 with good sampling. Both salinity channels were found to be within 0.001 or bottles, on average.

Dissolved Oxygen

The sensor has been used only twice since its last recalibration in March 2011, and for one cruise a different method of checking calibration was used. For 2013-01 with a good range of DO values, the fit found had a slope of 1.066 and offset of +0.0189.

Pressure

The sensor was recalibrated in April 2011 and was used for 1 cruise in 2011 as well as 2013-01. The factory offset was used for both those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. There was no local climatology available for the Malaspina Inlet casts. For all others, the temperature and salinity profiles fell within the local climatology.

Repeat Casts – There were no repeat casts and there is too much variability for comparison of nearby casts to have any value in assessing repeatability of the sensors.

Post-Cruise Calibration - There were no post-cruise calibrations available.

##### CHANNEL CHOICE STUDY

COMPARE does not help in choosing which sensor pair to select for archiving. The two salinity channels are very close in values during stops. T-S plots show only slight differences with sometimes one pair and sometimes the other looking smoother. Since the primary was used for 2013-01 and the alignment settings were the same for the primary as during 2013-01, the primary temperature and salinity channels were selected.

##### DETAILED EDITING

The primary temperature and salinity channels were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the cast.

All casts required light editing, mostly at the top and bottom of casts. Editing was a little heavier for cast #45 because the data came from the upcast.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure does not appear to require recalibration, but pressure should be monitored.

No salinity recalibration will be applied. Based on the results of 2013-01 and the fact that the two sensor pairs are very close during stops, it is likely that there has not been significant calibration drift in any of the temperature or conductivity channels.

SBE Dissolved Oxygen was recalibrated by applying the following equation:

Corrected DO = 1.0701 * SBE DO

CALIBRATE was run using file 2013-16-recal1.ccf to apply the corrections to the SAM files.

COMPARE was rerun to ensure the correction worked well and it did. The results show values slightly low at the lower end of the range and slightly high at the higher end. (See 2013-16-DO-comp2.xlsx.)

CALIBRATE was then applied to the EDT and MRGCLN2 files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure the CTD looked slightly high near the surface, but was generally close to the bottles, especially for DO < 4mL/L.

No further DO correction is justified.

The results looked good in both plots of differences against pressure and DO concentration.

(See 2013-16-dox-comp3.xls for details.)

##### Special Fluorometer Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. There were a few small unstable features in the T-S plots but there is no evidence that these are due to instrumental problems; they may well be real. Profiles turned up no problems.

##### Final CTD File Steps (REMOVE and HEADEDIT)

The steps in this section were run twice – first to produce CTD files, and then a second run to produce CTDpH files that include the pH channel.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE,Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1.0  mL/L from 0 to 35db

±0.3  mL/L from 35 to 150db

±0.08 mL/L below 150db

The comparison of CTD salinity with bottle samples did not produce useful results.

The two CTD salinity channels are in reasonable agreement with each other, and

the temperature and conductivity sensors have only been used for 2 other

cruises since their last factory calibration.

A pH sensor was attached to the CTD, but the data are not included in the

files to be placed in the IOS Data Library because no field calibration

information was available at the time of processing, and there are some

concerns about how the sensors perform.

For details on the processing see the report: 2013-16-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. There was a lot of variability with many very high values (up to 175%) in the southern part of Malaspina Inlet. The casts in more open waters had values between 95% and 130%. A few cases of higher saturations were checked to see if the SBE DO near-surface waters were in reasonable agreement with bottles and they were overall, but there were significant differences with bottles sometimes higher than CTD and sometimes lower. This is likely due to high temporal and spatial variability.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The rest of the steps in this section were run twice– first to produce CHE files, and then a second run to produce CHEpH files that include the pH channel.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE,Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and no errors were found.

A header check was run on the final files and no problems were found.

Track plots look ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

# Pressure sensor not working at first, cleared the line and then it was fine.

CTD reading on deck = 0.861db

14. Station name was MAL7 in the log but MAL8 in the file – MAL7 looks right; file changed to MAL7. 45. Archived on upcast only. There were 8 bottles fired.

74/83. Log gives CTD cast at MAL40 as event #83, but file was saved as event #74. File name and event number were changed to 83.

140. Station name wrong –should be MAL45 – noted in log – changed in file.

146. No O2

184. Error in Station name – should be GH1- error noted in log - changed in file.

191. Error in station name – should be BS5 – error noted in log - changed in file.

202. Latitude in log different from NMEA position by 1’. Presume typo in log. No change made to file.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information CTD #506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 5Apr2011 | Factory |  |  |
| Conductivity | 3321 | 30Mar2011 | Factory |  |  |
| Secondary Temp. | 4700 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 1766 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 1Apr2011 | Factory |  |  |
| PAR | 4601 | Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | Dec2010 |  |  |  |
| SBE Fluorometer | 2228 | n/a |  |  |  |
| Pressure Sensor | 77511 | 22Apr2011 | Factory |  |  |
| Altimeter | 16504 | 16Mar2011 |  |  |  |