## Notes
Analysis: | Hugh MacLean, onboard CCGS JP Tully using IOS Kit #1. Data processing by Kenny Scozzafava
Chief Scientist: | Marie Robert
Program: | Line P
Cruise Dates: | 20-August-13 to 06-September-13
2013-10-25 |  | KS
*Important* The water traps connected to the titrant and standard bottles on the Dosimats were both left empty for the duration of the cruise. The traps
are meant to maintain a sealed and moist environment inside the bottle in order to help control evaporation and the potential change in the concentration of
the solution that will occur if the bottle is left open to air. This may not have been so critical on the first standardization (Aug 25) but subsequent
standardizations (Aug 29 and Sept 3) may be suspect due to the unknown change this may have caused in both the titrant and standard concentrations.
The standard volume was also very low after the second standardization. The fact that thio titers did not change significantly is a bit of a mystery but
it may be due to the lab being close to 100% humidity thereby slowing concentration changes inside the bottles. This data is likely still ok but should still be viewed
with a cautious eye.
Precision is also relatively low (0.010 mL/L), even after the removal of known bad values and outliers selected by the Chauvenet criteria.
It is recommended that this data be trusted to the second decimal place and not the third and is reported as such in the QF spreadsheet.
Draw temperature thermometer was malfunctioning. It was swapped at sample #424 with a spare that seemed to work properly from that point forward.
Event 9, Station P4 - Samplers had many problems with bubbles being injected into samples via the Alkaline Iodide bottle top dispenser during this cast.
Because the concentration of dissolved oxygen assumes that accurate of additions of 1 mL of each reagent, there is a good likelihood that these samples may
have reduced accuracy due to the problems with the bottle top dispenser.
Event 31, Station P12- Flasks were not shaken a second time and were missing a water seal.
2013-11-04 |  | KS
Event 31, Sample 205 (OXY-2) - There was a bubble on the colourimeter probe lens that led to an incomplete titration and a value much lower than its replicate and
duplicate Niskin. Value was discarded and the replicates were left out of the precision calculation.
Event 89, Sample 591 - Both reps suffered from poor titrations that had no endpoints. The values are flagged as a 5 and the values padded.
They were also left out of the precision calculation.
Event 89, Sample 607 (OXY-1) - There was a bubble on the colourimeter probe lens during the tiration. There was no endpoint. Just a short straight line. Value was not reported.
Duplicates were also left out of the precision calculation.
2013-11-05 |  | KS
Event 31, Sample 210 - Replicate outliers according to the Chauvenet criteria. There were no notes to point to a reason why values may be so far apart.
2014-03-05 |  | GG
Samples 25-32 are from cast #5, not #3
Sample #79 | Flagged 5 with pad value |  | Analyst flagged 4 and it is a severe outlier in comparison between bottle and CTD DO
Sample #206 | Flagged 3 |  | Analyst noted noisy titration; outlier in comparison of bottle and CTD DO
2014-04-08 |  | KS
Modified all appropriate worksheets to reflect corrections from March 05, 2014
2014-04-09 |  | KS
Event 89, sample 591 - Modified comment for conciseness. The portion of the comment that stated "top valve completely open" was transposed incorrectly from the oxy analysis log sheet and should have applied to sample 592.
Event 89, Sample 592 - Added to comment that the Niskin top valve was open. The titration of this sample was also problematic. Value flagged with a 4.

## QF 2013-18 OXY FINAL
HEADER: | Oxygen samples were analyzed at sea using an automated Winkler titration system (Brinkmann Dosimat model 665) for titrations,
Dosimat model 865 for dispensing standards and a PC 950 Colorimeter controlled by AutoOxy software V3.4) with modifications based on
Carpenter (1965) and adhering to WOCE protocols (Culberson 1991). For details including a duplicate analysis, see document QF2013-18OXY.xls.
Cruise_Number | Event_Number | Station | Sample_Number | Oxygen:Dissolved | Flag:Oxygen:Dissolved | Temperature:Draw | Comment
2013-18 | 2 | P1 | 8 | 1.0949 | 6 | 7.1
2013-18 | 2 | P1 | 10 | 6.7658 | 6 | 12.3
2013-18 | 5 | P2 | 25 | 1.3254 | 6 | 7.3 | OXY: Cast originally labeled as 3
2013-18 | 5 | P2 | 26 | 1.2997 | 2 | 7.3 | OXY: Cast originally labeled as 3
2013-18 | 5 | P2 | 27 | 1.827 | 2 | 7.5 | OXY: Cast originally labeled as 3
2013-18 | 5 | P2 | 28 | 1.7956 | 2 | 7.6 | OXY: Cast originally labeled as 3
2013-18 | 5 | P2 | 29 | 2.2444 | 2 | 7.9 | OXY: Cast originally labeled as 3
2013-18 | 5 | P2 | 30 | 3.4715 | 2 | 9.1 | OXY: Cast originally labeled as 3
2013-18 | 5 | P2 | 31 | 5.7917 | 2 | 11.2 | OXY: Cast originally labeled as 3
2013-18 | 5 | P2 | 32 | 7.3568 | 2 | 12.9 | OXY: Cast originally labeled as 3
2013-18 | 9 | P4 | 79 | -99 | 5 | 4.4 | OXY: Niskin had a bit of a leak. Titration failure. Out of line in the depth profile. Bubbles may have been introduced via the Alkaline Iodide bottle top dispenser. Severe outlier in comparison between bottle and CTD DO.
2013-18 | 9 | P4 | 80 | 0.40443 | 3 | 4.6 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 81 | 0.37704 | 36 | 5 | OXY-1: Bubbles may have been introduced via the Alkaline Iodide dispenser. OXY-2: Titration ended abruptly after the endpoint but otherwise ok. No problems were observed by the analyst.
2013-18 | 9 | P4 | 82 | 0.22988 | 3 | 5.5 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 83 | 0.20934 | 3 | 6 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 84 | 0.39328 | 4 | 6.4 | OXY: Poor titration curve led to an errant endpoint. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 85 | 1.0254 | 4 | 8.1 | OXY: Sample had problems with bubbles but there is no further explanation by the analyst. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 86 | 1.5451 | 3 | 7.4 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 87 | 1.8892 | 3 | 8.1 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 88 | 2.4754 | 36 | 8.95 | OXY-1: Bubbles may have been introduced via the Alkaline Iodide dispenser. OXY-2: Bubbles in sample.
2013-18 | 9 | P4 | 89 | 2.9601 | 3 | 8.9 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 90 | 3.2193 | 3 | 8.8 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 91 | 3.6393 | 3 | 9 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 92 | 4.9123 | 3 | 8.8 | OXY: Small bubble in sample. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 93 | 6.2673 | 36 | 9.1 | OXY-1: Noisy titration but endpoint ok. Bubbles may have been introduced via the Alkaline Iodide dispenser. OXY-2: Odd titration curve led to an errant endpoint. Adjusted to more appropriate position.
2013-18 | 9 | P4 | 94 | 6.5182 | 3 | 9.7 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 95 | 6.274 | 3 | 15 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 96 | 5.8311 | 3 | 17.7 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 9 | P4 | 97 | 5.7535 | 3 | 18.1 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-18 | 23 | P8 | 144 | 1.3901 | 36 | 3.85 | OXY-1: Titration ended abruptly after the endpoint but otherwise ok. No problems were observed by the analyst. OXY-2: Very messy titration curve led to an ambiguous endpoint.
2013-18 | 23 | P8 | 145 | 0.6467 |  | 4.1
2013-18 | 23 | P8 | 146 | 0.38427 | 3 | 4.5 | OXY: Poor titration curve led to errant endpoint.
2013-18 | 23 | P8 | 147 | 0.26099 | 3 | 5.3 | OXY: Titration changed directions very slowly rendering the endpoint ambiguous. It did seem late so was adjusted accordingly.
2013-18 | 23 | P8 | 148 | 0.21069 |  | 5.7
2013-18 | 23 | P8 | 149 | 0.3901 |  | 6.1
2013-18 | 23 | P8 | 150 | 1.1281 |  | 6.7
2013-18 | 23 | P8 | 151 | 2.0957 |  | 7.1
2013-18 | 23 | P8 | 152 | 2.6963 | 6 | 7.6
2013-18 | 23 | P8 | 153 | 2.9569 |  | 8
2013-18 | 23 | P8 | 154 | 3.2039 |  | 8.3
2013-18 | 23 | P8 | 155 | 3.4789 |  | 8.4
2013-18 | 23 | P8 | 156 | 4.1555 |  | 8.4
2013-18 | 23 | P8 | 157 | 4.8036 |  | 8.5
2013-18 | 23 | P8 | 158 | 6.5049 |  | 9.1
2013-18 | 23 | P8 | 159 | 6.1695 |  | 11.6
2013-18 | 23 | P8 | 160 | 6.2921 | 6 | 15.45
2013-18 | 23 | P8 | 161 | 5.7858 |  | 17.7
2013-18 | 23 | P8 | 162 | 5.7757 |  | 17.8
2013-18 | 31 | P12 | 203 | 2.3243 | 3 | 3.1 | OXY: Titration was quite noisy but endpoint seems to be in best position.
2013-18 | 31 | P12 | 204 | 2.2842 |  | 3.2
2013-18 | 31 | P12 | 205 | 2.2837 | 2 | 3.4 | OXY: Replicate value not reported. OXY-2 sample had an incomplete titration and a value much lower than its rep and duplicate Niskin. Sample was discarded.
2013-18 | 31 | P12 | 206 | 1.8474 | 3 | 3.5 | OXY: Titration noisy but endpoint ok. Corrected depth value that was inputted incorrectly before analysis. Outlier in comparison between CTD and bottle DO.
2013-18 | 31 | P12 | 207 | 1.3023 |  | 3.5
2013-18 | 31 | P12 | 208 | 0.60638 |  | 4.2
2013-18 | 31 | P12 | 209 | 0.33157 |  | 5.1
2013-18 | 31 | P12 | 210 | 0.3091 | 46 | 5.55 | OXY: Replicate outliers
2013-18 | 31 | P12 | 211 | 0.28636 |  | 6
2013-18 | 31 | P12 | 212 | 0.603 |  | 6.5
2013-18 | 31 | P12 | 213 | 1.3933 |  | 6.6
2013-18 | 31 | P12 | 214 | -99 | 1 | 7 | OXY: Titration ended abruptly before reaching the endpoint. Sample lost.
2013-18 | 31 | P12 | 215 | 2.8542 |  | 7.5
2013-18 | 31 | P12 | 216 | 3.7595 |  | 7.9
2013-18 | 31 | P12 | 217 | 3.7708 |  | 8.3
2013-18 | 31 | P12 | 218 | 4.745 | 6 | 8.3
2013-18 | 31 | P12 | 219 | 5.6371 |  | 8.3
2013-18 | 31 | P12 | 220 | 6.2609 |  | 8.5
2013-18 | 31 | P12 | 221 | 6.7474 |  | 8.7
2013-18 | 31 | P12 | 222 | 6.7914 | 6 | 10.1
2013-18 | 31 | P12 | 223 | 6.5758 |  | 15.5
2013-18 | 31 | P12 | 224 | 5.9212 |  | 17.1
2013-18 | 31 | P12 | 225 | 5.902 |  | 17.3
2013-18 | 42 | P16 | 293 | 2.7486 | 3 | 4.1 | OXY: Niskin bottle spigot was leaking prior to sampling
2013-18 | 42 | P16 | 294 | 2.6589 | 3 | 3.9 | OXY: Poor titration curve led to an errant endpoint.
2013-18 | 42 | P16 | 295 | 2.6686 |  | 4.4
2013-18 | 42 | P16 | 296 | 2.4148 | 3 | 4 | OXY: Poor titration curve led to an ambiguous endpoint.
2013-18 | 42 | P16 | 297 | 1.9153 |  | 4.3
2013-18 | 42 | P16 | 298 | 1.2352 | 26 | 4.75 | OXY-1: Messy titration. Endpoint seems ok.
2013-18 | 42 | P16 | 299 | 0.53929 |  | 5.4
2013-18 | 42 | P16 | 300 | 0.31936 |  | 5.4
2013-18 | 42 | P16 | 301 | 0.27284 |  | 5.7
2013-18 | 42 | P16 | 302 | 0.33513 |  | 6.1
2013-18 | 42 | P16 | 303 | 0.46176 |  | 6.4
2013-18 | 42 | P16 | 304 | 1.2634 |  | 6.9
2013-18 | 42 | P16 | 305 | 2.0911 | 6 | 7.35
2013-18 | 42 | P16 | 306 | 2.6421 |  | 7.5
2013-18 | 42 | P16 | 307 | 3.0542 |  | 8.3
2013-18 | 42 | P16 | 308 | 3.7858 | 6 | 8.4
2013-18 | 42 | P16 | 309 | 4.9236 |  | 8.3
2013-18 | 42 | P16 | 310 | 5.7195 |  | 8.4
2013-18 | 42 | P16 | 311 | 6.7736 |  | 8.3
2013-18 | 42 | P16 | 312 | 6.7934 | 6 | 8.8
2013-18 | 42 | P16 | 313 | 6.5315 |  | 9.9
2013-18 | 42 | P16 | 314 | 6.1984 |  | 15.7
2013-18 | 42 | P16 | 315 | 5.9082 |  | 16.7
2013-18 | 42 | P16 | 316 | 5.8958 |  | 16.6
2013-18 | 59 | P20 | 409 | 2.9991 |  | 3.2
2013-18 | 59 | P20 | 410 | 2.8226 |  | 3.2
2013-18 | 59 | P20 | 411 | 2.8253 | 6 | 3.85
2013-18 | 59 | P20 | 412 | 2.4572 |  | 3.8
2013-18 | 59 | P20 | 413 | 1.906 | 3 | 4 | OXY: An odd step in the flat portion of the curve led to a very late endpoint. Adjusted back.
2013-18 | 59 | P20 | 414 | 1.2097 |  | 4
2013-18 | 59 | P20 | 415 | 0.51464 |  | 4.7
2013-18 | 59 | P20 | 416 | 0.315 |  | 4.8
2013-18 | 59 | P20 | 417 | 0.33924 |  | 5.3
2013-18 | 59 | P20 | 418 | 0.44191 | 6 | 5.65
2013-18 | 59 | P20 | 419 | 0.67906 |  | 6
2013-18 | 59 | P20 | 420 | 1.4587 |  | 6.2
2013-18 | 59 | P20 | 421 | 2.2455 |  | 6.3
2013-18 | 59 | P20 | 422 | 2.9144 |  | 6.7
2013-18 | 59 | P20 | 423 | 3.6699 | 6 | 7.35
2013-18 | 59 | P20 | 424 | 4.2444 |  | 7.7
2013-18 | 59 | P20 | 425 | 5.4353 | 4 | 7.2 | OXY: A bubble on the probe lens led to a big vertical jump in the curve near the endpoint. Endpoint is unclear.
2013-18 | 59 | P20 | 426 | 6.7039 |  | 7.4
2013-18 | 59 | P20 | 427 | 6.8367 |  | 7.7
2013-18 | 59 | P20 | 428 | 7.0097 |  | 7.8
2013-18 | 59 | P20 | 429 | 7.1789 | 3 | 8.9 | OXY: Niskin bottle spigot was leaking prior to sampling
2013-18 | 59 | P20 | 430 | 6.9653 |  | 12.4
2013-18 | 59 | P20 | 431 | 5.8318 | 6 | 16
2013-18 | 59 | P20 | 432 | 5.8369 |  | 16
2013-18 | 89 | P26 | 590 | 3.1986 | 4 | 3.2 | OXY: Poor titration curve led to an errant endpoint.
2013-18 | 89 | P26 | 591 | -99 | 5 | 2.7 | OXY: Both reps suffered from poor titrations that had no endpoints. The values are flagged as a 5 and the values padded
2013-18 | 89 | P26 | 592 | 2.9608 | 4 | 2.6 | OXY: Niskin top valve was completely open. Very messy titration curve led to a very ambiguous endpoint.
2013-18 | 89 | P26 | 593 | 2.5637 |  | 2.7
2013-18 | 89 | P26 | 594 | 1.9442 |  | 2.9
2013-18 | 89 | P26 | 595 | 1.2634 | 3 | 3.1 | OXY: Messy curve near endpoint. Endpoint ambiguous.
2013-18 | 89 | P26 | 596 | 0.59621 |  | 3.5
2013-18 | 89 | P26 | 597 | 0.52261 | 36 | 3.85 | OXY-1: Niskin spigot was open prior to sampling.
2013-18 | 89 | P26 | 598 | 0.50318 |  | 4.1
2013-18 | 89 | P26 | 599 | 0.5255 |  | 4.5
2013-18 | 89 | P26 | 600 | 0.66103 |  | 4.9
2013-18 | 89 | P26 | 601 | 0.79258 |  | 5.1
2013-18 | 89 | P26 | 602 | 1.4041 |  | 5.1
2013-18 | 89 | P26 | 603 | 1.8185 | 6 | 5.2
2013-18 | 89 | P26 | 604 | 2.6658 |  | 5.2
2013-18 | 89 | P26 | 605 | 3.3929 |  | 5.2
2013-18 | 89 | P26 | 606 | 4.6222 |  | 5.6
2013-18 | 89 | P26 | 607 | 6.6148 | 2 | 5.9 | OXY: Replicate value not reported. OXY-1: Bubble on lens. There was no endpoint. Just a short straight line.
2013-18 | 89 | P26 | 608 | 6.9848 |  | 6.2
2013-18 | 89 | P26 | 609 | 6.9771 | 3 | 6.5 | OXY: Endpoint was ambiguous due to a messy titration
2013-18 | 89 | P26 | 610 | 7.0566 | 3 | 8.6 | OXY: Niskin bottle spigot was leaking prior to sampling
2013-18 | 89 | P26 | 611 | 5.9969 |  | 14.1
2013-18 | 89 | P26 | 612 | 5.9239 | 3 | 14.9 | OXY: Niskin bottle spigot was leaking prior to sampling
2013-18 | 89 | P26 | 613 | 5.9284 |  | 14.9
2013-18 | 105 | P4 | 647 | 0.4546 | 3 | 4.8 | OXY: Niskin bottle top valve was leaking. New O-ring was installed after this cast.
2013-18 | 105 | P4 | 648 | 0.40014 |  | 4.5
2013-18 | 105 | P4 | 649 | 0.20968 | 6 | 5.15
2013-18 | 105 | P4 | 650 | 0.22531 |  | 5.5
2013-18 | 105 | P4 | 651 | 0.49135 | 3 | 5.9 | OXY: Niskin bottle was leaking in an unspecified location. A new O-ring was installed prior to next cast.
2013-18 | 105 | P4 | 652 | 1.0295 |  | 6.5
2013-18 | 105 | P4 | 653 | 2.1586 |  | 6.7
2013-18 | 105 | P4 | 654 | 2.6619 |  | 7 | OXY: Niskin bottle spigot was leaking prior to sampling. A new O-ring was installed prior to next cast.
2013-18 | 105 | P4 | 655 | 2.6243 |  | 7.1
2013-18 | 105 | P4 | 656 | 3.0273 | 6 | 7.55
2013-18 | 105 | P4 | 657 | 5.5629 |  | 7.8
2013-18 | 105 | P4 | 658 | 6.5817 |  | 9.2
2013-18 | 105 | P4 | 659 | 6.2525 |  | 14.8
2013-18 | 105 | P4 | 660 | 6.2107 |  | 12.6
2013-18 | 105 | P4 | 661 | 6.2157 |  | 12.6
2013-18 | 105 | P4 | 662 | 5.6616 |  | 17.5

## Precision
Precision Statement for Replicate Samples Drawn From a Single Niskin Bottle:
The mean difference for Dissolved Oxygen for the range 0.209 to 6.797 mL/L was 0.010, k = 48 with 2 outlier(s) removed
Precision Statement for Samples Drawn From Duplicate Niskin Bottles:
The mean difference for Dissolved Oxygen for the range 0.377 to 2.825 mL/L was 0.014, k = 10 with 0 outlier(s) removed
Note: Although the precision statement for samples drawn from duplicate Niskin bottles is calculated using the same formula as the precision statement for duplicates
samples drawn from one single Niskin, this process is mainly used to identify problem samples and is not being used as a measure of analysis precision.
Precision Analysis and the Determination of Outliers
Precision was determined by analyzing replicate samples drawn from one Niskin
Outliers are discarded on the basis of Chauvenet's criteria. The statistic is calculated using the following formula:
The maximum deviation, dmax, is compared with the individual residuals from the original concentrations (see Table 1, columns N, O and P).
If a replicate's residual is greater than dmax, this reading can be rejected based upon Chauvenet's criterion
Table 1. Replicate Analysis for Samples Drawn From a Single Niskin Bottle | [20 empty] | Table 2. Critical values for Chauvenet's Criteria
Event | Station | Sample_ Number | Depth | OXY-1 | OXY-2 | OXYAvg | Flask1 | Flask2 | Draw T1 | Draw T2 | Draw TAvg | Diff | Residual1 (Xi-Xm) | Residual2 (Xi-Xm) | Residual3 (Xi-Xm) | Σ(Xi-Xm)2 | n | Reject? | OXY-1 Comment | OXY-2 Comment | n | Chauvenet
2 | P1 | 8 | 111 | 1.1236 | 1.0662 | 1.0949 | 3070 | 3076 | 7.1 | 7.1 | 7.1 | 0.057336 | 0.028668 | -0.028668 |  | 0.0016437 | 2 |  |  |  | 3 | 1.37
2 | P1 | 10 | 5 | 6.7689 | 6.7627 | 6.7658 | 1019 | 620 | 12.3 | 12.3 | 12.3 | 0.0061987 | 0.0030994 | -0.0030994 |  | 1.9212e-05 | 2 |  |  |  | 4 | 1.53
5 | P2 | 25 | 105 | 1.3256 | 1.3253 | 1.3254 | 3029 | 639 | 7.3 | 7.3 | 7.3 | 0.00027905 | 0.00013952 | -0.00013952 |  | 3.8933e-08 | 2 |  |  |  | 5 | 1.64
9 | P4 | 81 | 1250 | 0.37304 | 0.38104 | 0.37704 | 2033 | 712 | 4.5 | 5.5 | 5 | 0.0079927 | -0.0039963 | 0.0039963 |  | 3.1941e-05 | 2 |  | OXY-1: Bubbles may have been introduced via the Alkaline Iodide dispenser. | OXY-2: Titration ended abruptly after the endpoint but otherwise ok. No problems were observed by the analyst. Bubbles may have been introduced via the Alkaline Iodide dispenser. | 6 | 1.73
9 | P4 | 88 | 200 | 2.474 | 2.4768 | 2.4754 | 3209 | 715 | 8.8 | 9.1 | 8.95 | 0.0028589 | -0.0014295 | 0.0014295 |  | 4.0867e-06 | 2 |  | OXY-1: Bubbles may have been introduced via the Alkaline Iodide dispenser. | OXY-2: Bubbles in sample. Bubbles may have been introduced via the Alkaline Iodide dispenser. | 7 | 1.8
9 | P4 | 93 | 75 | 6.2628 | 6.2717 | 6.2673 | 2086 | 631 | 9.1 | 9.1 | 9.1 | 0.0089092 | -0.0044546 | 0.0044546 |  | 3.9687e-05 | 2 |  | OXY-1: Noisy titration but endpoint ok. Bubbles may have been introduced via the Alkaline Iodide dispenser. | OXY-2: Odd titration curve led to an errant endpoint. Adjusted to more apprpriate poaition. Bubbles may have been introduced via the Alkaline Iodide dispenser. | 8 | 1.87
23 | P8 | 144 | 2000 | 1.398 | 1.3822 | 1.3901 | 3151 | 3207 | 3.8 | 3.9 | 3.85 | 0.015845 | 0.0079227 | -0.0079227 |  | 0.00012554 | 2 |  | OXY-1: Titration ended abruptly after the endpoint but otherwise ok. No problems were observed by the analyst. | OXY-2: Very messy titration curve led to an ambiguous endpoint. | 9 | 1.91
23 | P8 | 152 | 250 | 2.7109 | 2.6816 | 2.6963 | 3198 | 3182 | 7.5 | 7.7 | 7.6 | 0.029363 | 0.014681 | -0.014681 |  | 0.00043108 | 2 |  |  |  | 10 | 1.97
23 | P8 | 160 | 25 | 6.3078 | 6.2764 | 6.2921 | 3201 | 3206 | 15.4 | 15.5 | 15.45 | 0.031388 | 0.015694 | -0.015694 |  | 0.00049261 | 2 |  |  |  | 20 | 2.24
31 | P12 | 210 | 1000 | 0.36356 | 0.25464 | 0.3091 | 1174 | 1178 | 5.4 | 5.7 | 5.55 | 0.10892 | 0.054459 | -0.054459 |  | 0.0059316 | 2 | Y | OXY-1: Replicate outliers |  | 30 | 2.39
31 | P12 | 218 | 150 | 4.7549 | 4.7351 | 4.745 | 1119 | 1114 | 8.2 | 8.4 | 8.3 | 0.019897 | 0.0099483 | -0.0099483 |  | 0.00019794 | 2 |  |  |  | 50 | 2.57
31 | P12 | 222 | 50 | 6.7913 | 6.7915 | 6.7914 | 1144 | 1136 | 10 | 10.2 | 10.1 | 0.0002032 | -0.0001016 | 0.0001016 |  | 2.0645e-08 | 2 |  |  |  | 80 | 2.74
42 | P16 | 298 | 2000 | 1.2332 | 1.2373 | 1.2352 | 1413 | 1407 | 4.4 | 5.1 | 4.75 | 0.004089 | -0.0020445 | 0.0020445 |  | 8.3601e-06 | 2 |  | OXY-1: Messy titration. Endpoint seems ok. |  | 100 | 2.81
42 | P16 | 305 | 300 | 2.098 | 2.0842 | 2.0911 | 1411 | 1402 | 7.2 | 7.5 | 7.35 | 0.013807 | 0.0069033 | -0.0069033 |  | 9.5311e-05 | 2 |  |  |  | 150 | 2.93
42 | P16 | 308 | 175 | 3.788 | 3.7835 | 3.7858 | 1430 | 1409 | 8.3 | 8.5 | 8.4 | 0.0045254 | 0.0022627 | -0.0022627 |  | 1.024e-05 | 2 |  |  |  | 200 | 3.02
42 | P16 | 312 | 75 | 6.7901 | 6.7966 | 6.7934 | 1416 | 1421 | 8.7 | 8.9 | 8.8 | 0.0065581 | -0.0032791 | 0.0032791 |  | 2.1505e-05 | 2
59 | P20 | 411 | 3500 | 2.8265 | 2.824 | 2.8253 | 1192 | 1188 | 3.8 | 3.9 | 3.85 | 0.0025271 | 0.0012635 | -0.0012635 |  | 3.1931e-06 | 2
59 | P20 | 418 | 800 | 0.44389 | 0.43993 | 0.44191 | 1134 | 1146 | 5.6 | 5.7 | 5.65 | 0.0039579 | 0.001979 | -0.001979 |  | 7.8326e-06 | 2
59 | P20 | 423 | 200 | 3.6698 | 3.67 | 3.6699 | 1126 | 1119 | 7.4 | 7.3 | 7.35 | 0.00028356 | -0.00014178 | 0.00014178 |  | 4.0203e-08 | 2 |  | Sample was redrawn
59 | P20 | 431 | 10 | 5.831 | 5.8327 | 5.8318 | 1123 | 1110 | 15.9 | 16.1 | 16 | 0.0017396 | -0.00086981 | 0.00086981 |  | 1.5131e-06 | 2
89 | P26 | 597 | 1250 | 0.52025 | 0.52496 | 0.52261 | 1174 | 1178 | 3.7 | 4 | 3.85 | 0.0047072 | -0.0023536 | 0.0023536 |  | 1.1079e-05 | 2 |  | OXY-1: Niskin spigot was open prior to sampling. | Incorrect flask ID. Should be 1178
89 | P26 | 603 | 250 | 1.8174 | 1.8196 | 1.8185 | 1130 | 1126 | 5.1 | 5.3 | 5.2 | 0.0022375 | -0.0011187 | 0.0011187 |  | 2.5031e-06 | 2
105 | P4 | 649 | 1000 | 0.21041 | 0.20895 | 0.20968 | 2033 | 712 | 5.1 | 5.2 | 5.15 | 0.0014572 | 0.00072861 | -0.00072861 |  | 1.0617e-06 | 2
105 | P4 | 656 | 200 | 3.0276 | 3.027 | 3.0273 | 3209 | 3038 | 7.5 | 7.6 | 7.55 | 0.00060768 | 0.00030384 | -0.00030384 |  | 1.8463e-07 | 2
[15 empty] | Sum | 0.0090802
[15 empty] | Count |  | 48
[15 empty] | σ | 0.0139
[15 empty] | dmax/σ for n (Table 2) | 2.57
[15 empty] | Thus dmax = 2.57 x σ | 0.035722
[15 empty] | Mean Diff | 0.013987
[15 empty] | MIN | 0.20895
[15 empty] | MAX | 6.7966
2 Outlier(s) Removed
[15 empty] | Mean Diff = | 0.0098595
Table 1. Replicate Analysis for Samples Drawn From a Single Niskin Bottle | [20 empty] | Table 2. Critical values for Chauvenet's Criteria
Event | Station | Sample_ Number | Depth | OXY-1 | OXY-2 | OXYAvg | Flask1 | Flask2 | Draw T1 | Draw T2 | Draw TAvg | Diff | Residual1 (Xi-Xm) | Residual2 (Xi-Xm) | Residual3 (Xi-Xm) | Σ(Xi-Xm)2 | n | Reject? | OXY-1 Comment | OXY-2 Comment | n | Chauvenet
3 | P2 | 27/28 | 75 | 1.827 | 1.7956 | 1.8113 | 196 | 2040 | 7.5 | 7.6 | 7.55 | 0.031437 | 0.015718 | -0.015718 |  | 0.00049414 | 2 |  |  |  | 3 | 1.37
9 | P4 | 80/81 | 1250 | 0.40443 | 0.37704 | 0.39074 | 468 | 2033/712 | 4.6 | 5 | 4.8 | 0.027394 | 0.013697 | -0.013697 |  | 0.00037523 | 2 |  | OXY-1: Bubbles may have been introduced via the Alkaline Iodide dispenser. | OXY-2: Bubbles may have been introduced via the Alkaline Iodide dispenser. | 4 | 1.53
31 | P12 | 204/205 | 3000 | 2.2842 | 2.2837 | 2.2839 | 1127 | 1192/1188 | 3.2 | 3.3 | 3.25 | 0.00047752 | 0.00023876 | -0.00023876 |  | 1.1401e-07 | 2 |  |  | OXY-2: Bubble on lens led to an incomplete titration and a value much lower than its rep and duplicate Niskin. Sample discarded. | 5 | 1.64
42 | P16 | 294/295 | 3500 | 2.6589 | 2.6686 | 2.6637 | 1434 | 1442 | 3.9 | 4.4 | 4.15 | 0.0097045 | -0.0048522 | 0.0048522 |  | 4.7088e-05 | 2 |  | OXY-1: Poor titration curve led to an errant endpoint. |  | 6 | 1.73
59 | P20 | 410/411 | 3500 | 2.8226 | 2.8253 | 2.824 | 1127 | 1192/1188 | 3.2 | 3.85 | 3.525 | 0.0026411 | -0.0013205 | 0.0013205 |  | 3.4876e-06 | 2 |  |  |  | 7 | 1.8
[21 empty] | 8 | 1.87
[21 empty] | 9 | 1.91
[21 empty] | 10 | 1.97
[21 empty] | 20 | 2.24
[21 empty] | 30 | 2.39
[21 empty] | 50 | 2.57
[21 empty] | 80 | 2.74
[21 empty] | 100 | 2.81
[21 empty] | 150 | 2.93
[21 empty] | 200 | 3.02
[15 empty] | Sum | 0.00092005
[15 empty] | Count |  | 10
[15 empty] | σ | 0.010111
[15 empty] | dmax/σ for n (Table 2) | 1.97
[15 empty] | Thus dmax = 1.97 x σ | 0.019918
[15 empty] | Mean Diff | 0.014331
[15 empty] | MIN | 0.37704
[15 empty] | MAX | 2.8253
0 Outlier(s) Removed
[15 empty] | Mean Diff = | 0.014331

## Flags
Institute of Ocean Sciences - Quality Flags for Water Sampling |  | Last Updated: February 28, 2013
Double click document to use.
Current version of entire quality flag document can be found on OSDShare Network Drive: | O:\Documentation\Quality Flags

## Oxy Data Processing
File Name | Event | Station | Sample_ Number | Depth | Raw Oxygen: Dissolved | AutoOxy Thiosulfate_ titer | Regression Thiosulfate_ titer | Flask_ Number | Flask_ Volume | Blank | Standard | KIO3 Std volume ml | KIO3 Normality | THIO N | AutoOxy Oxygen: Dissolved | AutoOxy Oxygen: Dissolved | diff | Flag | Temp: Draw | Comment
2013-180002.oxy | 2 | P1 | 8 | 111 | 1.124 | 0.209 | 0.209 | 3070 | 145.33 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.1236 | 1.1236 | 0 |  | 7.1
2013-180002.oxy | 2 | P1 | 8 | 111 | 1.066 | 0.191 | 0.191 | 3076 | 139.72 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.0662 | 1.0662 | 0 |  | 7.1
2013-180002.oxy | 2 | P1 | 10 | 5 | 6.769 | 1.21 | 1.21 | 1019 | 142.45 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7689 | 6.7689 | 0 |  | 12.3
2013-180002.oxy | 2 | P1 | 10 | 5 | 6.761 | 1.139 | 1.1392 | 620 | 134.32 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7615 | 6.7627 | 0.0011924 |  | 12.3
2013-180003.oxy | 5 | P2 | 25 | 105 | 1.324 | 0.221 | 0.2213 | 3029 | 130.81 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.3237 | 1.3256 | 0.0018373 |  | 7.3
2013-180003.oxy | 5 | P2 | 25 | 105 | 1.325 | 0.227 | 0.227 | 639 | 134.23 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.3253 | 1.3253 | 0 | 2 | 7.3 | OXY: Cast originally labeled as 3
2013-180003.oxy | 5 | P2 | 26 | 100 | 1.3 | 0.227 | 0.227 | 632 | 136.83 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.2997 | 1.2997 | 0 |  | 7.3 | OXY: Cast originally labeled as 3
2013-180003.oxy | 5 | P2 | 27 | 75 | 1.83 | 0.318 | 0.3174 | 196 | 136.95 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.8305 | 1.827 | 0.0035074 |  | 7.5 | OXY: Cast originally labeled as 3
2013-180003.oxy | 5 | P2 | 28 | 75 | 1.793 | 0.327 | 0.3275 | 2040 | 143.75 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.7928 | 1.7956 | 0.0027826 |  | 7.6 | OXY: Cast originally labeled as 3
2013-180003.oxy | 5 | P2 | 29 | 50 | 2.244 | 0.41 | 0.41 | 3121 | 144.4 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.2444 | 2.2444 | 0 |  | 7.9 | OXY: Cast originally labeled as 3
2013-180003.oxy | 5 | P2 | 30 | 25 | 3.47 | 0.599 | 0.5992 | 569 | 137.06 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.4703 | 3.4715 | 0.0011682 |  | 9.1 | OXY: Cast originally labeled as 3
2013-180003.oxy | 5 | P2 | 31 | 10 | 5.796 | 1.041 | 1.0402 | 1050 | 143.02 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.7962 | 5.7917 | 0.0044752 |  | 11.2 | OXY: Cast originally labeled as 3
2013-180003.oxy | 5 | P2 | 32 | 5 | 7.36 | 1.341 | 1.3404 | 1054 | 145.21 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 7.3601 | 7.3568 | 0.0033051 |  | 12.9 | OXY: Cast originally labeled as 3
2013-180009.oxy | 9 | P4 | 79 | 1318 | 0.114 | 0.026 | 0.026 | 1094 | 148.09 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | -99 | -99 | 0 | 5 | 4.4 | OXY: Niskin had a bit of a leak. Titration failure. Out of line in the depth profile. Bubbles may have been introduced via the Alkaline Iodide bottle top dispenser. Severe outlier in comparison between bottle and CTD DO.
2013-180009.oxy | 9 | P4 | 80 | 1250 | 0.408 | 0.074 | 0.0734 | 468 | 135.7 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.40797 | 0.40443 | 0.0035402 | 3 | 4.6 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 81 | 1250 | 0.373 | 0.072 | 0.072 | 2033 | 143.99 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.37304 | 0.37304 | 0 | 3 | 4.5 | OXY-1: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 81 | 1250 | 0.399 | 0.07 | 0.067 | 712 | 130.66 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.39943 | 0.38104 | 0.018394 | 3 | 5.5 | OXY-2: Titration ended abruptly after the endpoint but otherwise ok. No problems were observed by the analyst. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 82 | 1000 | 0.232 | 0.044 | 0.0437 | 2090 | 135.3 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.23166 | 0.22988 | 0.0017754 | 3 | 5.5 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 83 | 800 | 0.209 | 0.04 | 0.04 | 510 | 134.44 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.20934 | 0.20934 | 0 | 3 | 6 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 84 | 600 | 0.4 | 0.076 | 0.0748 | 3038 | 142.3 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.40003 | 0.39328 | 0.0067472 | 4 | 6.4 | OXY: Poor titration curve led to an errant endpoint. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 85 | 400 | 1.027 | 0.19 | 0.1897 | 2049 | 144.21 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.027 | 1.0254 | 0.0016642 | 4 | 8.1 | OXY: Sample had problems with bubbles but there is no further explanation by the analyst. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 86 | 300 | 1.547 | 0.26 | 0.2597 | 611 | 132.11 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.547 | 1.5451 | 0.0018189 | 3 | 7.4 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 87 | 250 | 1.889 | 0.346 | 0.346 | 1053 | 144.45 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.8892 | 1.8892 | 0 | 3 | 8.1 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 88 | 200 | 2.474 | 0.426 | 0.426 | 3209 | 136.29 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.474 | 2.474 | 0 | 3 | 8.8 | OXY-1: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 88 | 200 | 2.477 | 0.436 | 0.436 | 715 | 139.32 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.4768 | 2.4768 | 0 | 4 | 9.1 | OXY-2: Bubbles in sample. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 89 | 175 | 2.959 | 0.508 | 0.5082 | 3111 | 136.14 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.959 | 2.9601 | 0.0011762 | 3 | 8.9 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 90 | 150 | 3.214 | 0.586 | 0.587 | 2014 | 144.65 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.2138 | 3.2193 | 0.0055301 | 3 | 8.8 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 91 | 125 | 3.639 | 0.599 | 0.599 | 621 | 130.79 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.6393 | 3.6393 | 0 | 3 | 9 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 92 | 100 | 4.916 | 0.894 | 0.8934 | 3142 | 144.69 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 4.9157 | 4.9123 | 0.0033171 | 3 | 8.8 | OXY: Small bubble in sample. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 93 | 75 | 6.263 | 1.126 | 1.126 | 2086 | 143.22 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.2628 | 6.2628 | 0 | 3 | 9.1 | OXY-1: Noisy titration but endpoint ok. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 93 | 75 | 6.298 | 1.08 | 1.0756 | 631 | 136.68 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.2975 | 6.2717 | 0.025772 | 3 | 9.1 | OXY-2: Odd titration curve led to an errant endpoint. Adjusted to more appropriate position. Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 94 | 50 | 6.518 | 1.197 | 1.197 | 2065 | 146.28 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.5182 | 6.5182 | 0 | 3 | 9.7 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 95 | 25 | 6.273 | 1.108 | 1.1082 | 3009 | 140.73 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.2729 | 6.274 | 0.0011373 | 3 | 15 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 96 | 10 | 5.83 | 0.97 | 0.9703 | 2055 | 132.61 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.8293 | 5.8311 | 0.001812 | 3 | 17.7 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180009.oxy | 9 | P4 | 97 | 5 | 5.76 | 0.982 | 0.981 | 3047 | 135.84 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.7594 | 5.7535 | 0.0058941 | 3 | 18.1 | OXY: Bubbles may have been introduced via the Alkaline Iodide dispenser.
2013-180023.oxy | 23 | P8 | 144 | 2000 | 1.398 | 0.255 | 0.255 | 3151 | 143.15 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.398 | 1.398 | 0 | 3 | 3.8 | OXY-1: Titration ended abruptly after the endpoint but otherwise ok. No problems were observed by the analyst.
2013-180023.oxy | 23 | P8 | 144 | 2000 | 1.382 | 0.244 | 0.244 | 3207 | 138.49 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.3822 | 1.3822 | 0 | 4 | 3.9 | OXY-2: Very messy titration curve led to an ambiguous endpoint.
2013-180023.oxy | 23 | P8 | 145 | 1500 | 0.653 | 0.112 | 0.111 | 3187 | 131.48 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.65279 | 0.6467 | 0.0060926 |  | 4.1
2013-180023.oxy | 23 | P8 | 146 | 1250 | 0.384 | 0.068 | 0.068 | 3193 | 131.63 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.38427 | 0.38427 | 0 | 3 | 4.5 | OXY: Poor titration curve led to errant endpoint.
2013-180023.oxy | 23 | P8 | 147 | 1000 | 0.279 | 0.051 | 0.048 | 3196 | 132.41 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.27914 | 0.26099 | 0.018147 | 3 | 5.3 | OXY: Titration changed directions very slowly rendering the endpoint ambiguous. It did seem late so was adjusted accordingly.
2013-180023.oxy | 23 | P8 | 148 | 800 | 0.211 | 0.04 | 0.04 | 3189 | 133.59 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.21069 | 0.21069 | 0 |  | 5.7
2013-180023.oxy | 23 | P8 | 149 | 600 | 0.387 | 0.07 | 0.0706 | 3191 | 134.95 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.38654 | 0.3901 | 0.0035601 |  | 6.1
2013-180023.oxy | 23 | P8 | 150 | 400 | 1.127 | 0.19 | 0.1902 | 3188 | 131.61 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.1269 | 1.1281 | 0.0012173 |  | 6.7
2013-180023.oxy | 23 | P8 | 151 | 300 | 2.093 | 0.36 | 0.3605 | 3184 | 135.87 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.0928 | 2.0957 | 0.0029464 |  | 7.1
2013-180023.oxy | 23 | P8 | 152 | 250 | 2.71 | 0.467 | 0.4671 | 3198 | 136.51 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.7104 | 2.7109 | 0.00058647 |  | 7.5
2013-180023.oxy | 23 | P8 | 152 | 250 | 2.682 | 0.459 | 0.459 | 3182 | 135.6 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.6816 | 2.6816 | 0 |  | 7.7
2013-180023.oxy | 23 | P8 | 153 | 200 | 2.956 | 0.541 | 0.5412 | 3177 | 145.09 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.9558 | 2.9569 | 0.0011026 |  | 8
2013-180023.oxy | 23 | P8 | 154 | 175 | 3.204 | 0.578 | 0.578 | 3178 | 143.12 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.2039 | 3.2039 | 0 |  | 8.3
2013-180023.oxy | 23 | P8 | 155 | 150 | 3.48 | 0.639 | 0.6388 | 3166 | 145.75 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.48 | 3.4789 | 0.0010976 |  | 8.4
2013-180023.oxy | 23 | P8 | 156 | 125 | 4.158 | 0.758 | 0.7576 | 3160 | 144.9 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 4.1577 | 4.1555 | 0.0022082 |  | 8.4
2013-180023.oxy | 23 | P8 | 157 | 100 | 4.805 | 0.864 | 0.8638 | 3171 | 143.06 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 4.8047 | 4.8036 | 0.0011185 |  | 8.5
2013-180023.oxy | 23 | P8 | 158 | 75 | 6.511 | 1.085 | 1.084 | 3157 | 132.87 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.511 | 6.5049 | 0.0060279 |  | 9.1
2013-180023.oxy | 23 | P8 | 159 | 50 | 6.175 | 1.042 | 1.0411 | 3200 | 134.5 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.1749 | 6.1695 | 0.0053583 |  | 11.6
2013-180023.oxy | 23 | P8 | 160 | 25 | 6.308 | 1.086 | 1.086 | 3201 | 137.21 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.3078 | 6.3078 | 0 |  | 15.4
2013-180023.oxy | 23 | P8 | 160 | 25 | 6.275 | 1.135 | 1.1352 | 3206 | 144.07 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.2753 | 6.2764 | 0.0011105 |  | 15.5
2013-180023.oxy | 23 | P8 | 161 | 10 | 5.786 | 1.058 | 1.058 | 3204 | 145.59 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.7858 | 5.7858 | 0 |  | 17.7
2013-180023.oxy | 23 | P8 | 162 | 5 | 5.775 | 1.052 | 1.0522 | 3181 | 145.05 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.7746 | 5.7757 | 0.0011029 |  | 17.8
2013-180031.oxy | 31 | P12 | 203 | 3276 | 2.324 | 0.401 | 0.401 | 1177 | 136.45 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.3243 | 2.3243 | 0 | 3 | 3.1 | OXY: Titration is quite noisy but endpoint seems to be in best position.
2013-180031.oxy | 31 | P12 | 204 | 3000 | 2.284 | 0.393 | 0.393 | 1127 | 136.05 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.2842 | 2.2842 | 0 |  | 3.2
2013-180031.oxy | 31 | P12 | 205 | 3000 | 2.284 | 0.386 | 0.386 | 1192 | 133.66 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.2837 | 2.2837 | 0 |  | 3.3
2013-180031.oxy | 31 | P12 | 205 | 3000 | 1.957 | 0.342 | 0.342 | 1188 | 137.94 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | -99 | -99 | 0 | 5 | 3.4 | OXY-2: Bubble on lens led to an incomplete titration and a value much lower than its rep and duplicate Niskin. Sample discarded.
2013-180031.oxy | 31 | P12 | 206 | 2500 | 1.846 | 0.323 | 0.3232 | 1188 | 137.94 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.8462 | 1.8474 | 0.0011606 | 3 | 3.5 | OXY: Titration noisy but endpoint ok. Corrected depth value that was inputted incorrectly before analysis. Outlier in comparison between CTD and bottle DO.
2013-180031.oxy | 31 | P12 | 207 | 2000 | 1.299 | 0.223 | 0.2236 | 1150 | 134.5 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.2988 | 1.3023 | 0.0035722 |  | 3.5
2013-180031.oxy | 31 | P12 | 208 | 1500 | 0.605 | 0.113 | 0.1132 | 1175 | 142.95 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.60526 | 0.60638 | 0.0011194 |  | 4.2
2013-180031.oxy | 31 | P12 | 209 | 1250 | 0.332 | 0.061 | 0.061 | 1189 | 135.58 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.33157 | 0.33157 | 0 |  | 5.1
2013-180031.oxy | 31 | P12 | 210 | 1000 | 0.355 | 0.064 | 0.0655 | 1174 | 133.59 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.35457 | 0.36356 | 0.0089923 | 4 | 5.4 | OXY-1: Replicate outliers
2013-180031.oxy | 31 | P12 | 210 | 1000 | 0.258 | 0.049 | 0.0485 | 1178 | 137.21 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.25756 | 0.25464 | 0.0029172 | 4 | 5.7 | OXY-2: Replicate outliers
2013-180031.oxy | 31 | P12 | 211 | 800 | 0.285 | 0.054 | 0.0543 | 1134 | 138.21 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.28463 | 0.28636 | 0.0017375 |  | 6
2013-180031.oxy | 31 | P12 | 212 | 600 | 0.603 | 0.101 | 0.101 | 1146 | 127.78 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.603 | 0.603 | 0 |  | 6.5
2013-180031.oxy | 31 | P12 | 213 | 400 | 1.393 | 0.237 | 0.237 | 1141 | 133.44 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.3933 | 1.3933 | 0 |  | 6.6
2013-180031.oxy | 31 | P12 | 214 | 300 | 2.437 | 0.435 | 0.435 | 1182 | 141.23 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | -99 | -99 | 0 | 1 | 7 | OXY: Titration ended abruptly before reaching the endpoint. Sample lost..
2013-180031.oxy | 31 | P12 | 215 | 250 | 2.854 | 0.491 | 0.4911 | 1124 | 136.39 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.8537 | 2.8542 | 0.000587 |  | 7.5
2013-180031.oxy | 31 | P12 | 216 | 200 | 3.764 | 0.615 | 0.6143 | 1130 | 129.88 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.7639 | 3.7595 | 0.0043182 |  | 7.9
2013-180031.oxy | 31 | P12 | 217 | 175 | 3.769 | 0.613 | 0.6133 | 1126 | 129.29 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.7689 | 3.7708 | 0.0018592 |  | 8.3
2013-180031.oxy | 31 | P12 | 218 | 150 | 4.758 | 0.789 | 0.7885 | 1119 | 132.01 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 4.758 | 4.7549 | 0.0030339 |  | 8.2
2013-180031.oxy | 31 | P12 | 218 | 150 | 4.736 | 0.834 | 0.8339 | 1114 | 140.12 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 4.7356 | 4.7351 | 0.00057115 |  | 8.4
2013-180031.oxy | 31 | P12 | 219 | 125 | 5.64 | 0.944 | 0.9436 | 1106 | 133.37 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.6395 | 5.6371 | 0.002402 |  | 8.3
2013-180031.oxy | 31 | P12 | 220 | 100 | 6.261 | 1.007 | 1.007 | 1105 | 128.27 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.2609 | 6.2609 | 0 |  | 8.5
2013-180031.oxy | 31 | P12 | 221 | 75 | 6.746 | 1.107 | 1.1072 | 1158 | 130.88 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7462 | 6.7474 | 0.0012242 |  | 8.7
2013-180031.oxy | 31 | P12 | 222 | 50 | 6.794 | 1.171 | 1.1705 | 1144 | 137.4 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7942 | 6.7913 | 0.0029131 |  | 10
2013-180031.oxy | 31 | P12 | 222 | 50 | 6.786 | 1.103 | 1.1039 | 1136 | 129.66 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7859 | 6.7915 | 0.0055615 |  | 10.2
2013-180031.oxy | 31 | P12 | 223 | 25 | 6.576 | 1.097 | 1.097 | 1128 | 133.02 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.5758 | 6.5758 | 0 |  | 15.5
2013-180031.oxy | 31 | P12 | 224 | 10 | 5.921 | 1.009 | 1.009 | 1123 | 135.78 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.9212 | 5.9212 | 0 |  | 17.1
2013-180031.oxy | 31 | P12 | 225 | 5 | 5.902 | 0.942 | 0.942 | 1110 | 127.26 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.902 | 5.902 | 0 |  | 17.3
2013-180042.oxy | 42 | P16 | 293 | 3682 | 2.747 | 0.469 | 0.4693 | 1446 | 135.3 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.7468 | 2.7486 | 0.0017754 | 3 | 4.1 | OXY: Niskin bottle spigot was leaking prior to sampling
2013-180042.oxy | 42 | P16 | 294 | 3500 | 2.665 | 0.442 | 0.441 | 1434 | 131.4 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.665 | 2.6589 | 0.0060963 | 3 | 3.9 | OXY: Poor titration curve led to an errant endpoint.
2013-180042.oxy | 42 | P16 | 295 | 3500 | 2.669 | 0.491 | 0.491 | 1442 | 145.71 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.6686 | 2.6686 | 0 |  | 4.4
2013-180042.oxy | 42 | P16 | 296 | 3000 | 2.421 | 0.408 | 0.407 | 1445 | 133.37 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.4209 | 2.4148 | 0.0060049 | 3 | 4 | OXY: Poor titration curve led to an ambiguous endpoint.
2013-180042.oxy | 42 | P16 | 297 | 2500 | 1.915 | 0.318 | 0.318 | 1440 | 130.98 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.9153 | 1.9153 | 0 |  | 4.3
2013-180042.oxy | 42 | P16 | 298 | 2000 | 1.233 | 0.21 | 0.21 | 1413 | 133.23 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.2332 | 1.2332 | 0 | 2 | 4.4 | OXY-1: Messy titration. Endpoint seems ok.
2013-180042.oxy | 42 | P16 | 298 | 2000 | 1.237 | 0.225 | 0.225 | 1407 | 142.36 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.2373 | 1.2373 | 0 |  | 5.1
2013-180042.oxy | 42 | P16 | 299 | 1500 | 0.539 | 0.094 | 0.094 | 1433 | 132.4 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.53929 | 0.53929 | 0 |  | 5.4
2013-180042.oxy | 42 | P16 | 300 | 1250 | 0.328 | 0.061 | 0.0595 | 1427 | 136.98 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.32813 | 0.31936 | 0.0087665 |  | 5.4
2013-180042.oxy | 42 | P16 | 301 | 1000 | 0.273 | 0.05 | 0.05 | 1432 | 132.53 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.27284 | 0.27284 | 0 |  | 5.7
2013-180042.oxy | 42 | P16 | 302 | 800 | 0.335 | 0.061 | 0.061 | 1422 | 134.16 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.33513 | 0.33513 | 0 |  | 6.1
2013-180042.oxy | 42 | P16 | 303 | 600 | 0.458 | 0.087 | 0.0876 | 1417 | 143.36 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.45841 | 0.46176 | 0.0033483 |  | 6.4
2013-180042.oxy | 42 | P16 | 304 | 400 | 1.265 | 0.215 | 0.2147 | 1443 | 133.03 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.2652 | 1.2634 | 0.0018061 |  | 6.9
2013-180042.oxy | 42 | P16 | 305 | 300 | 2.098 | 0.372 | 0.372 | 1411 | 140.05 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.098 | 2.098 | 0 |  | 7.2
2013-180042.oxy | 42 | P16 | 305 | 300 | 2.084 | 0.355 | 0.355 | 1402 | 134.53 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.0842 | 2.0842 | 0 |  | 7.5
2013-180042.oxy | 42 | P16 | 306 | 250 | 2.642 | 0.464 | 0.464 | 1428 | 139.09 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.6421 | 2.6421 | 0 |  | 7.5
2013-180042.oxy | 42 | P16 | 307 | 200 | 3.055 | 0.51 | 0.5098 | 1429 | 132.42 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.0554 | 3.0542 | 0.0012097 |  | 8.3
2013-180042.oxy | 42 | P16 | 308 | 175 | 3.782 | 0.628 | 0.629 | 1430 | 131.98 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.7819 | 3.788 | 0.0060691 |  | 8.3
2013-180042.oxy | 42 | P16 | 308 | 175 | 3.785 | 0.671 | 0.6707 | 1409 | 140.83 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.7852 | 3.7835 | 0.0017047 |  | 8.5
2013-180042.oxy | 42 | P16 | 309 | 150 | 4.925 | 0.855 | 0.8548 | 1406 | 138.18 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 4.9247 | 4.9236 | 0.0011586 |  | 8.3
2013-180042.oxy | 42 | P16 | 310 | 125 | 5.719 | 0.95 | 0.95 | 1426 | 132.36 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.7195 | 5.7195 | 0 |  | 8.4
2013-180042.oxy | 42 | P16 | 311 | 100 | 6.771 | 1.225 | 1.2256 | 1424 | 144.17 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7703 | 6.7736 | 0.0033292 |  | 8.3
2013-180042.oxy | 42 | P16 | 312 | 75 | 6.79 | 1.14 | 1.14 | 1416 | 133.88 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7901 | 6.7901 | 0 |  | 8.7
2013-180042.oxy | 42 | P16 | 312 | 75 | 6.793 | 1.194 | 1.1946 | 1421 | 140.09 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7932 | 6.7966 | 0.0034276 |  | 8.9
2013-180042.oxy | 42 | P16 | 313 | 50 | 6.531 | 1.099 | 1.099 | 1414 | 134.15 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.5315 | 6.5315 | 0 |  | 9.9
2013-180042.oxy | 42 | P16 | 314 | 25 | 6.198 | 1.019 | 1.019 | 1408 | 131.07 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.1984 | 6.1984 | 0 |  | 15.7
2013-180042.oxy | 42 | P16 | 315 | 10 | 5.909 | 1.089 | 1.0888 | 1404 | 146.73 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.9092 | 5.9082 | 0.0010901 |  | 16.7
2013-180042.oxy | 42 | P16 | 316 | 5 | 5.898 | 0.968 | 0.9677 | 1401 | 130.83 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.8976 | 5.8958 | 0.001837 |  | 16.6
2013-180059.oxy | 59 | P20 | 409 | 4024 | 2.999 | 0.516 | 0.516 | 1177 | 136.45 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.9991 | 2.9991 | 0 |  | 3.2
2013-180059.oxy | 59 | P20 | 410 | 3500 | 2.831 | 0.486 | 0.4845 | 1127 | 136.05 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.8315 | 2.8226 | 0.0088273 |  | 3.2
2013-180059.oxy | 59 | P20 | 411 | 3500 | 2.823 | 0.476 | 0.4766 | 1192 | 133.66 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.823 | 2.8265 | 0.003595 |  | 3.8
2013-180059.oxy | 59 | P20 | 411 | 3500 | 2.821 | 0.491 | 0.4915 | 1188 | 137.94 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.8211 | 2.824 | 0.0029015 |  | 3.9
2013-180059.oxy | 59 | P20 | 412 | 3000 | 2.457 | 0.409 | 0.409 | 1180 | 131.75 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.4572 | 2.4572 | 0 |  | 3.8
2013-180059.oxy | 59 | P20 | 413 | 2500 | 1.942 | 0.331 | 0.325 | 1150 | 134.5 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.9418 | 1.906 | 0.035722 | 3 | 4 | OXY: An odd step in the flat portion of the curve led to a very late endpoint. Adjusted back.
2013-180059.oxy | 59 | P20 | 414 | 2000 | 1.21 | 0.221 | 0.221 | 1175 | 142.95 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.2097 | 1.2097 | 0 |  | 4
2013-180059.oxy | 59 | P20 | 415 | 1500 | 0.515 | 0.092 | 0.092 | 1189 | 135.58 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.51464 | 0.51464 | 0 |  | 4.7
2013-180059.oxy | 59 | P20 | 416 | 1250 | 0.313 | 0.057 | 0.0574 | 1174 | 133.59 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.3126 | 0.315 | 0.002398 |  | 4.8
2013-180059.oxy | 59 | P20 | 417 | 1000 | 0.345 | 0.064 | 0.063 | 1178 | 137.21 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.34507 | 0.33924 | 0.0058344 |  | 5.3
2013-180059.oxy | 59 | P20 | 418 | 800 | 0.441 | 0.081 | 0.0815 | 1134 | 138.21 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.441 | 0.44389 | 0.0028958 |  | 5.6
2013-180059.oxy | 59 | P20 | 418 | 800 | 0.44 | 0.075 | 0.075 | 1146 | 127.78 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.43993 | 0.43993 | 0 |  | 5.7
2013-180059.oxy | 59 | P20 | 419 | 600 | 0.679 | 0.118 | 0.118 | 1141 | 133.44 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 0.67906 | 0.67906 | 0 |  | 6
2013-180059.oxy | 59 | P20 | 420 | 400 | 1.457 | 0.262 | 0.2623 | 1182 | 141.23 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 1.457 | 1.4587 | 0.0016998 |  | 6.2
2013-180059.oxy | 59 | P20 | 421 | 300 | 2.249 | 0.388 | 0.3874 | 1124 | 136.39 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.2491 | 2.2455 | 0.003522 |  | 6.3
2013-180059.oxy | 59 | P20 | 422 | 250 | 2.919 | 0.478 | 0.4773 | 1130 | 129.88 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 2.9187 | 2.9144 | 0.0043182 |  | 6.7
2013-180059.oxy | 59 | P20 | 423 | 200 | 3.67 | 0.597 | 0.597 | 1126 | 129.29 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.6698 | 3.6698 | 0 |  | 7.4 | Sample was redrawn
2013-180059.oxy | 59 | P20 | 423 | 200 | 3.672 | 0.61 | 0.6097 | 1119 | 132.01 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 3.6719 | 3.67 | 0.0018203 |  | 7.3
2013-180059.oxy | 59 | P20 | 424 | 175 | 4.244 | 0.748 | 0.748 | 1114 | 140.12 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 4.2444 | 4.2444 | 0 |  | 7.7 | Draw temperature thermometer was changed from this point forward.
2013-180059.oxy | 59 | P20 | 425 | 150 | 5.435 | 0.91 | 0.91 | 1106 | 133.37 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.4353 | 5.4353 | 0 | 4 | 7.2 | OXY: A bubble on the probe lens led to a big vertical jump in the curve near the endpoint. Endpoint is unclear.
2013-180059.oxy | 59 | P20 | 426 | 125 | 6.705 | 1.119 | 1.1188 | 1138 | 133.08 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.7051 | 6.7039 | 0.0012036 |  | 7.4
2013-180059.oxy | 59 | P20 | 427 | 100 | 6.838 | 1.122 | 1.1218 | 1158 | 130.88 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.838 | 6.8367 | 0.0012242 |  | 7.7
2013-180059.oxy | 59 | P20 | 428 | 75 | 7.016 | 1.209 | 1.208 | 1144 | 137.4 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 7.0156 | 7.0097 | 0.0058262 |  | 7.8
2013-180059.oxy | 59 | P20 | 429 | 50 | 7.182 | 1.167 | 1.1666 | 1136 | 129.66 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 7.1814 | 7.1789 | 0.0024718 | 3 | 8.9 | OXY: Niskin bottle spigot was leaking prior to sampling
2013-180059.oxy | 59 | P20 | 430 | 25 | 6.967 | 1.162 | 1.1617 | 1128 | 133.02 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 6.9671 | 6.9653 | 0.0018063 |  | 12.4
2013-180059.oxy | 59 | P20 | 431 | 10 | 5.833 | 0.994 | 0.9937 | 1123 | 135.78 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.8327 | 5.831 | 0.001769 |  | 15.9
2013-180059.oxy | 59 | P20 | 431 | 10 | 5.839 | 0.932 | 0.931 | 1110 | 127.26 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.839 | 5.8327 | 0.0062978 |  | 16.1
2013-180059.oxy | 59 | P20 | 432 | 5 | 5.84 | 1.043 | 1.0425 | 1142 | 142.24 | 0.0027 | 0.716 | 10.0271 | 0.010025 | 0.14092 | 5.8397 | 5.8369 | 0.0028126 |  | 16
2013-180089.oxy | 89 | P26 | 590 | 4318 | 3.204 | 0.551 | 0.551 | 1177 | 136.45 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 3.1986 | 3.1986 | 0 | 4 | 3.2 | OXY: Poor titration curve led to an errant endpoint.
2013-180089.oxy | 89 | P26 | 591 | 4000 | 3.043 | 0.522 | 0.522 | 1127 | 136.05 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | -99 | -99 | 0 | 5 | 2.7 | OXY-1: Bubble on lens. Titration a straight line. No endpoint.
2013-180089.oxy | 89 | P26 | 591 | 4000 | 3.051 | 0.514 | 0.514 | 1192 | 133.66 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | -99 | -99 | 0 | 5 | 2.8 | OXY-2: Titration a series of 8 points with no discernible endpoint.
2013-180089.oxy | 89 | P26 | 592 | 3500 | 2.966 | 0.516 | 0.516 | 1188 | 137.94 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 2.9608 | 2.9608 | 0 | 4 | 2.6 | OXY: Niskin top valve was completely open. Very messy titration curve led to a very ambiguous endpoint.
2013-180089.oxy | 89 | P26 | 593 | 3000 | 2.585 | 0.43 | 0.4273 | 1180 | 131.75 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 2.5801 | 2.5637 | 0.016386 |  | 2.7 | Endpoint was extremely late
2013-180089.oxy | 89 | P26 | 594 | 2500 | 1.942 | 0.331 | 0.332 | 1150 | 134.5 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 1.9382 | 1.9442 | 0.0059429 |  | 2.9
2013-180089.oxy | 89 | P26 | 595 | 2000 | 1.266 | 0.231 | 0.231 | 1175 | 142.95 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 1.2634 | 1.2634 | 0 | 3 | 3.1 | OXY: Messy curve near endpoint. Endpoint ambiguous.
2013-180089.oxy | 89 | P26 | 596 | 1500 | 0.597 | 0.106 | 0.106 | 1189 | 135.58 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.59621 | 0.59621 | 0 |  | 3.5
2013-180089.oxy | 89 | P26 | 597 | 1250 | 0.522 | 0.092 | 0.0918 | 1174 | 133.59 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.52145 | 0.52025 | 0.0011968 | 3 | 3.7 | OXY-1: Niskin spigot was open prior to sampling.
2013-180089.oxy | 89 | P26 | 597 | 1250 | 0.54 | 0.095 | 0.095 | 1178 | 137.21 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.52496 | 0.52496 | 0 |  | 4 | Incorrect flask ID. Should be 1178
2013-180089.oxy | 89 | P26 | 598 | 1000 | 0.505 | 0.092 | 0.0919 | 1134 | 138.21 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.50376 | 0.50318 | 0.0005781 |  | 4.1
2013-180089.oxy | 89 | P26 | 599 | 800 | 0.528 | 0.089 | 0.0888 | 1146 | 127.78 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.52676 | 0.5255 | 0.0012521 |  | 4.5
2013-180089.oxy | 89 | P26 | 600 | 600 | 0.661 | 0.115 | 0.1152 | 1141 | 133.44 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.65983 | 0.66103 | 0.0011982 |  | 4.9
2013-180089.oxy | 89 | P26 | 601 | 400 | 0.8 | 0.146 | 0.145 | 1182 | 141.23 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.79824 | 0.79258 | 0.0056556 |  | 5.1
2013-180089.oxy | 89 | P26 | 602 | 300 | 1.41 | 0.245 | 0.2445 | 1124 | 136.39 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 1.4071 | 1.4041 | 0.0029296 |  | 5.1
2013-180089.oxy | 89 | P26 | 603 | 250 | 1.821 | 0.3 | 0.3 | 1130 | 129.88 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 1.8174 | 1.8174 | 0 |  | 5.1
2013-180089.oxy | 89 | P26 | 603 | 250 | 1.823 | 0.299 | 0.299 | 1126 | 129.29 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 1.8196 | 1.8196 | 0 |  | 5.3
2013-180089.oxy | 89 | P26 | 604 | 200 | 2.671 | 0.445 | 0.445 | 1119 | 132.01 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 2.6658 | 2.6658 | 0 |  | 5.2
2013-180089.oxy | 89 | P26 | 605 | 175 | 3.405 | 0.601 | 0.6 | 1114 | 140.12 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 3.3986 | 3.3929 | 0.0057011 |  | 5.2
2013-180089.oxy | 89 | P26 | 606 | 150 | 4.631 | 0.776 | 0.776 | 1106 | 133.37 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 4.6222 | 4.6222 | 0 |  | 5.6
2013-180089.oxy | 89 | P26 | 607 | 125 | 6.627 | 1.106 | 1.106 | 1138 | 133.08 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 6.6148 | 6.6148 | 0 |  | 6.1
2013-180089.oxy | 89 | P26 | 607 | 125 | 6.409 | 1.052 | 1.052 | 1158 | 130.88 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | -99 | -99 | 0 | 5 | 5.9 | OXY-1: Bubble on lens. There was no endpoint. Just a short straight line.
2013-180089.oxy | 89 | P26 | 608 | 100 | 6.998 | 1.206 | 1.2059 | 1144 | 137.4 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 6.9853 | 6.9848 | 0.00058156 |  | 6.2
2013-180089.oxy | 89 | P26 | 609 | 75 | 6.984 | 1.135 | 1.136 | 1136 | 129.66 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 6.9709 | 6.9771 | 0.0061682 | 3 | 6.5 | OXY: Endpoint is ambiguous due to a messy titration
2013-180089.oxy | 89 | P26 | 610 | 50 | 7.069 | 1.179 | 1.179 | 1128 | 133.02 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 7.0566 | 7.0566 | 0 | 3 | 8.6 | OXY: Niskin bottle spigot was leaking prior to sampling
2013-180089.oxy | 89 | P26 | 611 | 25 | 6.01 | 1.024 | 1.0237 | 1123 | 135.78 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 5.9987 | 5.9969 | 0.0017658 |  | 14.1
2013-180089.oxy | 89 | P26 | 612 | 10 | 5.933 | 0.947 | 0.9472 | 1110 | 127.26 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 5.9226 | 5.9239 | 0.0012573 | 3 | 14.9 | OXY: Niskin bottle spigot was leaking prior to sampling
2013-180089.oxy | 89 | P26 | 613 | 5 | 5.941 | 1.061 | 1.0607 | 1142 | 142.24 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 5.9301 | 5.9284 | 0.0016845 |  | 14.9
2013-180105.oxy | 105 | P4 | 647 | 1318 | 0.454 | 0.089 | 0.0892 | 1094 | 148.09 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.45352 | 0.4546 | 0.001078 | 3 | 4.8 | OXY: Niskin bottle top valve was leaking. New O-ring was installed after this cast.
2013-180105.oxy | 105 | P4 | 648 | 1250 | 0.402 | 0.073 | 0.0728 | 468 | 135.7 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.40132 | 0.40014 | 0.0011779 |  | 4.5
2013-180105.oxy | 105 | P4 | 649 | 1000 | 0.212 | 0.043 | 0.0428 | 2033 | 143.99 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.21152 | 0.21041 | 0.0011091 |  | 5.1
2013-180105.oxy | 105 | P4 | 649 | 1000 | 0.209 | 0.039 | 0.039 | 712 | 130.66 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.20895 | 0.20895 | 0 |  | 5.2
2013-180105.oxy | 105 | P4 | 650 | 800 | 0.226 | 0.043 | 0.043 | 2090 | 135.3 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.22531 | 0.22531 | 0 |  | 5.5
2013-180105.oxy | 105 | P4 | 651 | 600 | 0.495 | 0.088 | 0.0875 | 510 | 134.44 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 0.49432 | 0.49135 | 0.0029728 | 3 | 5.9 | OXY: Niskin bottle was leaking in an unspecified location. A new O-ring was installed prior to next cast.
2013-180105.oxy | 105 | P4 | 652 | 400 | 1.029 | 0.184 | 0.1844 | 715 | 139.32 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 1.0272 | 1.0295 | 0.0022937 |  | 6.5
2013-180105.oxy | 105 | P4 | 653 | 300 | 2.164 | 0.395 | 0.3947 | 2049 | 144.21 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 2.1602 | 2.1586 | 0.0016611 |  | 6.7
2013-180105.oxy | 105 | P4 | 654 | 250 | 2.669 | 0.445 | 0.4447 | 611 | 132.11 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 2.6638 | 2.6619 | 0.0018156 |  | 7 | OXY: Niskin bottle spigot was leaking prior to sampling. A new O-ring was installed prior to next cast.
2013-180105.oxy | 105 | P4 | 655 | 250 | 2.626 | 0.479 | 0.4796 | 1053 | 144.45 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 2.6209 | 2.6243 | 0.0033167 |  | 7.1
2013-180105.oxy | 105 | P4 | 656 | 200 | 3.032 | 0.521 | 0.5212 | 3209 | 136.29 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 3.0265 | 3.0276 | 0.0011727 |  | 7.5
2013-180105.oxy | 105 | P4 | 656 | 200 | 3.031 | 0.544 | 0.5442 | 3038 | 142.3 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 3.0259 | 3.027 | 0.0011225 |  | 7.6
2013-180105.oxy | 105 | P4 | 657 | 100 | 5.576 | 0.953 | 0.9525 | 3111 | 136.14 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 5.5658 | 5.5629 | 0.0029351 |  | 7.8
2013-180105.oxy | 105 | P4 | 658 | 50 | 6.593 | 1.197 | 1.1972 | 2014 | 144.65 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 6.5806 | 6.5817 | 0.001104 |  | 9.2
2013-180105.oxy | 105 | P4 | 659 | 25 | 6.261 | 1.027 | 1.0275 | 621 | 130.79 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 6.2494 | 6.2525 | 0.003057 |  | 14.8
2013-180105.oxy | 105 | P4 | 660 | 36 | 6.221 | 1.13 | 1.1303 | 3142 | 144.69 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 6.2091 | 6.2107 | 0.0016555 |  | 12.6
2013-180105.oxy | 105 | P4 | 661 | 36 | 6.229 | 1.12 | 1.1196 | 2086 | 143.22 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 6.2179 | 6.2157 | 0.0022304 |  | 12.6
2013-180105.oxy | 105 | P4 | 662 | 5 | 5.677 | 0.974 | 0.9732 | 631 | 136.68 | 0.0027 | 0.7173 | 10.0271 | 0.010025 | 0.14066 | 5.6663 | 5.6616 | 0.0046773 |  | 17.5

## Raw AutoOxy Data
File Name | LOC:EVENT_NUMBER | LOC:STATION | Sample_number | Depth | Oxygen:Dissolved | Thiosulfate_titer | Flask_Number | Flask_Volume | Blank | Standard
2013-180002.oxy | 2 | P1 | 8 | 111 | 1.124 | 0.209 | 3070 | 145.33 | 0.0027 | 0.716
2013-180002.oxy | 2 | P1 | 8 | 111 | 1.066 | 0.191 | 3076 | 139.72 | 0.0027 | 0.716
2013-180002.oxy | 2 | P1 | 10 | 5 | 6.769 | 1.21 | 1019 | 142.45 | 0.0027 | 0.716
2013-180002.oxy | 2 | P1 | 10 | 5 | 6.761 | 1.139 | 620 | 134.32 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 25 | 105 | 1.324 | 0.221 | 3029 | 130.81 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 25 | 105 | 1.325 | 0.227 | 639 | 134.23 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 26 | 100 | 1.3 | 0.227 | 632 | 136.83 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 27 | 75 | 1.83 | 0.318 | 196 | 136.95 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 28 | 75 | 1.793 | 0.327 | 2040 | 143.75 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 29 | 50 | 2.244 | 0.41 | 3121 | 144.4 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 30 | 25 | 3.47 | 0.599 | 569 | 137.06 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 31 | 10 | 5.796 | 1.041 | 1050 | 143.02 | 0.0027 | 0.716
2013-180003.oxy | 5 | P2 | 32 | 5 | 7.36 | 1.341 | 1054 | 145.21 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 79 | 1318 | 0.114 | 0.026 | 1094 | 148.09 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 80 | 1250 | 0.408 | 0.074 | 468 | 135.7 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 81 | 1250 | 0.373 | 0.072 | 2033 | 143.99 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 81 | 1250 | 0.399 | 0.07 | 712 | 130.66 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 82 | 1000 | 0.232 | 0.044 | 2090 | 135.3 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 83 | 800 | 0.209 | 0.04 | 510 | 134.44 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 84 | 600 | 0.4 | 0.076 | 3038 | 142.3 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 85 | 400 | 1.027 | 0.19 | 2049 | 144.21 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 86 | 300 | 1.547 | 0.26 | 611 | 132.11 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 87 | 250 | 1.889 | 0.346 | 1053 | 144.45 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 88 | 200 | 2.474 | 0.426 | 3209 | 136.29 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 88 | 200 | 2.477 | 0.436 | 715 | 139.32 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 89 | 175 | 2.959 | 0.508 | 3111 | 136.14 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 90 | 150 | 3.214 | 0.586 | 2014 | 144.65 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 91 | 125 | 3.639 | 0.599 | 621 | 130.79 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 92 | 100 | 4.916 | 0.894 | 3142 | 144.69 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 93 | 75 | 6.263 | 1.126 | 2086 | 143.22 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 93 | 75 | 6.298 | 1.08 | 631 | 136.68 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 94 | 50 | 6.518 | 1.197 | 2065 | 146.28 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 95 | 25 | 6.273 | 1.108 | 3009 | 140.73 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 96 | 10 | 5.83 | 0.97 | 2055 | 132.61 | 0.0027 | 0.716
2013-180009.oxy | 9 | P4 | 97 | 5 | 5.76 | 0.982 | 3047 | 135.84 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 144 | 2000 | 1.398 | 0.255 | 3151 | 143.15 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 144 | 2000 | 1.382 | 0.244 | 3207 | 138.49 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 145 | 1500 | 0.653 | 0.112 | 3187 | 131.48 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 146 | 1250 | 0.384 | 0.068 | 3193 | 131.63 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 147 | 1000 | 0.279 | 0.051 | 3196 | 132.41 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 148 | 800 | 0.211 | 0.04 | 3189 | 133.59 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 149 | 600 | 0.387 | 0.07 | 3191 | 134.95 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 150 | 400 | 1.127 | 0.19 | 3188 | 131.61 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 151 | 300 | 2.093 | 0.36 | 3184 | 135.87 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 152 | 250 | 2.71 | 0.467 | 3198 | 136.51 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 152 | 250 | 2.682 | 0.459 | 3182 | 135.6 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 153 | 200 | 2.956 | 0.541 | 3177 | 145.09 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 154 | 175 | 3.204 | 0.578 | 3178 | 143.12 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 155 | 150 | 3.48 | 0.639 | 3166 | 145.75 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 156 | 125 | 4.158 | 0.758 | 3160 | 144.9 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 157 | 100 | 4.805 | 0.864 | 3171 | 143.06 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 158 | 75 | 6.511 | 1.085 | 3157 | 132.87 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 159 | 50 | 6.175 | 1.042 | 3200 | 134.5 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 160 | 25 | 6.308 | 1.086 | 3201 | 137.21 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 160 | 25 | 6.275 | 1.135 | 3206 | 144.07 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 161 | 10 | 5.786 | 1.058 | 3204 | 145.59 | 0.0027 | 0.716
2013-180023.oxy | 23 | P8 | 162 | 5 | 5.775 | 1.052 | 3181 | 145.05 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 203 | 3276 | 2.324 | 0.401 | 1177 | 136.45 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 204 | 3000 | 2.284 | 0.393 | 1127 | 136.05 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 205 | 3000 | 2.284 | 0.386 | 1192 | 133.66 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 205 | 3000 | 1.957 | 0.342 | 1188 | 137.94 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 206 | 3000 | 1.846 | 0.323 | 1188 | 137.94 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 207 | 2000 | 1.299 | 0.223 | 1150 | 134.5 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 208 | 1500 | 0.605 | 0.113 | 1175 | 142.95 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 209 | 1250 | 0.332 | 0.061 | 1189 | 135.58 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 210 | 1000 | 0.355 | 0.064 | 1174 | 133.59 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 210 | 1000 | 0.258 | 0.049 | 1178 | 137.21 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 211 | 800 | 0.285 | 0.054 | 1134 | 138.21 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 212 | 600 | 0.603 | 0.101 | 1146 | 127.78 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 213 | 400 | 1.393 | 0.237 | 1141 | 133.44 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 214 | 300 | 2.437 | 0.435 | 1182 | 141.23 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 215 | 250 | 2.854 | 0.491 | 1124 | 136.39 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 216 | 200 | 3.764 | 0.615 | 1130 | 129.88 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 217 | 175 | 3.769 | 0.613 | 1126 | 129.29 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 218 | 150 | 4.758 | 0.789 | 1119 | 132.01 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 218 | 150 | 4.736 | 0.834 | 1114 | 140.12 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 219 | 125 | 5.64 | 0.944 | 1106 | 133.37 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 220 | 100 | 6.261 | 1.007 | 1105 | 128.27 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 221 | 75 | 6.746 | 1.107 | 1158 | 130.88 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 222 | 50 | 6.794 | 1.171 | 1144 | 137.4 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 222 | 50 | 6.786 | 1.103 | 1136 | 129.66 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 223 | 25 | 6.576 | 1.097 | 1128 | 133.02 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 224 | 10 | 5.921 | 1.009 | 1123 | 135.78 | 0.0027 | 0.716
2013-180031.oxy | 31 | P12 | 225 | 5 | 5.902 | 0.942 | 1110 | 127.26 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 293 | 3682 | 2.747 | 0.469 | 1446 | 135.3 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 294 | 3500 | 2.665 | 0.442 | 1434 | 131.4 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 295 | 3500 | 2.669 | 0.491 | 1442 | 145.71 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 296 | 3000 | 2.421 | 0.408 | 1445 | 133.37 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 297 | 2500 | 1.915 | 0.318 | 1440 | 130.98 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 298 | 2000 | 1.233 | 0.21 | 1413 | 133.23 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 298 | 2000 | 1.237 | 0.225 | 1407 | 142.36 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 299 | 1500 | 0.539 | 0.094 | 1433 | 132.4 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 300 | 1250 | 0.328 | 0.061 | 1427 | 136.98 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 301 | 1000 | 0.273 | 0.05 | 1432 | 132.53 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 302 | 800 | 0.335 | 0.061 | 1422 | 134.16 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 303 | 600 | 0.458 | 0.087 | 1417 | 143.36 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 304 | 400 | 1.265 | 0.215 | 1443 | 133.03 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 305 | 300 | 2.098 | 0.372 | 1411 | 140.05 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 305 | 300 | 2.084 | 0.355 | 1402 | 134.53 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 306 | 250 | 2.642 | 0.464 | 1428 | 139.09 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 307 | 200 | 3.055 | 0.51 | 1429 | 132.42 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 308 | 175 | 3.782 | 0.628 | 1430 | 131.98 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 308 | 175 | 3.785 | 0.671 | 1409 | 140.83 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 309 | 150 | 4.925 | 0.855 | 1406 | 138.18 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 310 | 125 | 5.719 | 0.95 | 1426 | 132.36 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 311 | 100 | 6.771 | 1.225 | 1424 | 144.17 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 312 | 75 | 6.79 | 1.14 | 1416 | 133.88 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 312 | 75 | 6.793 | 1.194 | 1421 | 140.09 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 313 | 50 | 6.531 | 1.099 | 1414 | 134.15 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 314 | 25 | 6.198 | 1.019 | 1408 | 131.07 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 315 | 10 | 5.909 | 1.089 | 1404 | 146.73 | 0.0027 | 0.716
2013-180042.oxy | 42 | P16 | 316 | 5 | 5.898 | 0.968 | 1401 | 130.83 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 409 | 4024 | 2.999 | 0.516 | 1177 | 136.45 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 410 | 3500 | 2.831 | 0.486 | 1127 | 136.05 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 411 | 3500 | 2.823 | 0.476 | 1192 | 133.66 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 411 | 3500 | 2.821 | 0.491 | 1188 | 137.94 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 412 | 3000 | 2.457 | 0.409 | 1180 | 131.75 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 413 | 2500 | 1.942 | 0.331 | 1150 | 134.5 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 414 | 2000 | 1.21 | 0.221 | 1175 | 142.95 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 415 | 1500 | 0.515 | 0.092 | 1189 | 135.58 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 416 | 1250 | 0.313 | 0.057 | 1174 | 133.59 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 417 | 1000 | 0.345 | 0.064 | 1178 | 137.21 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 418 | 800 | 0.441 | 0.081 | 1134 | 138.21 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 418 | 800 | 0.44 | 0.075 | 1146 | 127.78 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 419 | 600 | 0.679 | 0.118 | 1141 | 133.44 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 420 | 400 | 1.457 | 0.262 | 1182 | 141.23 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 421 | 300 | 2.249 | 0.388 | 1124 | 136.39 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 422 | 250 | 2.919 | 0.478 | 1130 | 129.88 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 423 | 200 | 3.67 | 0.597 | 1126 | 129.29 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 423 | 200 | 3.672 | 0.61 | 1119 | 132.01 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 424 | 175 | 4.244 | 0.748 | 1114 | 140.12 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 425 | 150 | 5.435 | 0.91 | 1106 | 133.37 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 426 | 125 | 6.705 | 1.119 | 1138 | 133.08 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 427 | 100 | 6.838 | 1.122 | 1158 | 130.88 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 428 | 75 | 7.016 | 1.209 | 1144 | 137.4 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 429 | 50 | 7.182 | 1.167 | 1136 | 129.66 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 430 | 25 | 6.967 | 1.162 | 1128 | 133.02 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 431 | 10 | 5.833 | 0.994 | 1123 | 135.78 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 431 | 10 | 5.839 | 0.932 | 1110 | 127.26 | 0.0027 | 0.716
2013-180059.oxy | 59 | P20 | 432 | 5 | 5.84 | 1.043 | 1142 | 142.24 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 590 | 4318 | 3.204 | 0.551 | 1177 | 136.45 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 591 | 4000 | 3.043 | 0.522 | 1127 | 136.05 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 591 | 4000 | 3.051 | 0.514 | 1192 | 133.66 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 592 | 3500 | 2.966 | 0.516 | 1188 | 137.94 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 593 | 3000 | 2.585 | 0.43 | 1180 | 131.75 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 594 | 2500 | 1.942 | 0.331 | 1150 | 134.5 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 595 | 2000 | 1.266 | 0.231 | 1175 | 142.95 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 596 | 1500 | 0.597 | 0.106 | 1189 | 135.58 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 597 | 1250 | 0.522 | 0.092 | 1174 | 133.59 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 597 | 1250 | 0.54 | 0.095 | 1174 | 133.59 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 598 | 1000 | 0.505 | 0.092 | 1134 | 138.21 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 599 | 800 | 0.528 | 0.089 | 1146 | 127.78 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 600 | 600 | 0.661 | 0.115 | 1141 | 133.44 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 601 | 400 | 0.8 | 0.146 | 1182 | 141.23 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 602 | 300 | 1.41 | 0.245 | 1124 | 136.39 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 603 | 250 | 1.821 | 0.3 | 1130 | 129.88 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 603 | 250 | 1.823 | 0.299 | 1126 | 129.29 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 604 | 200 | 2.671 | 0.445 | 1119 | 132.01 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 605 | 175 | 3.405 | 0.601 | 1114 | 140.12 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 606 | 150 | 4.631 | 0.776 | 1106 | 133.37 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 607 | 125 | 6.627 | 1.106 | 1138 | 133.08 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 607 | 125 | 6.409 | 1.052 | 1158 | 130.88 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 608 | 100 | 6.998 | 1.206 | 1144 | 137.4 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 609 | 75 | 6.984 | 1.135 | 1136 | 129.66 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 610 | 50 | 7.069 | 1.179 | 1128 | 133.02 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 611 | 25 | 6.01 | 1.024 | 1123 | 135.78 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 612 | 10 | 5.933 | 0.947 | 1110 | 127.26 | 0.0027 | 0.716
2013-180089.oxy | 89 | P26 | 613 | 5 | 5.941 | 1.061 | 1142 | 142.24 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 647 | 1318 | 0.454 | 0.089 | 1094 | 148.09 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 648 | 1250 | 0.402 | 0.073 | 468 | 135.7 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 649 | 1000 | 0.212 | 0.043 | 2033 | 143.99 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 649 | 1000 | 0.209 | 0.039 | 712 | 130.66 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 650 | 800 | 0.226 | 0.043 | 2090 | 135.3 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 651 | 600 | 0.495 | 0.088 | 510 | 134.44 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 652 | 400 | 1.029 | 0.184 | 715 | 139.32 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 653 | 300 | 2.164 | 0.395 | 2049 | 144.21 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 654 | 250 | 2.669 | 0.445 | 611 | 132.11 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 655 | 250 | 2.626 | 0.479 | 1053 | 144.45 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 656 | 200 | 3.032 | 0.521 | 3209 | 136.29 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 656 | 200 | 3.031 | 0.544 | 3038 | 142.3 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 657 | 100 | 5.576 | 0.953 | 3111 | 136.14 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 658 | 50 | 6.593 | 1.197 | 2014 | 144.65 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 659 | 25 | 6.261 | 1.027 | 621 | 130.79 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 660 | 36 | 6.221 | 1.13 | 3142 | 144.69 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 661 | 36 | 6.229 | 1.12 | 2086 | 143.22 | 0.0027 | 0.716
2013-180105.oxy | 105 | P4 | 662 | 5 | 5.677 | 0.974 | 631 | 136.68 | 0.0027 | 0.716