## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2013-35 (NOTE: Full processing report files under 2013-36)

Agency: Ocean Networks Canada / NRCAN

Location: WCVI

Project: Venus

Party Chief: Dewey R.

Platform: J.P. Tully

Date: April 29, 2013 – May 6, 2013

Cruise: 2013-36

Agency: Ocean Networks Canada / NRCAN

Location: WCVI

Project: Neptune

Party Chief: Kulin I.

Platform: J.P. Tully

Date: May 9, 2013 – May 28, 2013

Processed by: Germaine Gatien

Date of Processing: 17 January 2014 -

Number of original HEX files: 42 (3 from 2013-35 and 39 from 2013-36)

Number of CTD files: 42

Number of bottle casts: 39* (2 from 2013-35 and 37 from 2013-36)

*For event #3 a CHE file will be archived in case needed later. Only sediments were collected.

For event #43 bottles were closed but not sampled, so no CHE file will be archived.

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#1438), a SeaPoint Fluorometer (#2228) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4601), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (no serial # available).

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The data from this cruise were processed together with those from cruise 2013-36 which followed it. The equipment was the same for the two legs of this Ocean Networks Canada / NRCAN work. The dissolved oxygen calibration sampling from the 2 cruises were combined. There were 3 casts from Leg 1 and 39 from Leg 2. The comments that follow are based on the whole cruise unless there is a note indicating otherwise.

The Daily Science Log was in good order, though there is no mention of the altimeter in the equipment list. There are only a few rosette log sheets, but there are DO analysis logs which provide the information needed since that was the only rosette sampling from the cruise, other than sediment sampling for 1 cast.

The descent rate of the CTD reached a maximum of 2.8m/s and was frequently >2m/s for short periods during deeper casts. These high rates are associated with rapid deceleration to low values, ~0.5m/s, so are likely due to swell. Such rapid deceleration results in corruption of data and is hard on the cable. The problem may be related to the way the ship is positioned during CTD deployment, as the use of the LARS system is perceived to remove the need to minimize ship motion.

There was no salinity calibration sampling but cruises 2013-12 and 2013-13 which bracket this cruise, and had the same equipment, both showed CTD salinity to be within 0.001 of bottle samples, on average.

There was dissolved oxygen sampling but the comparison between bottles and SBE dissolved oxygen was very noisy, particularly in Effingham Inlet where large DO gradients challenge the sensor and lead to significant errors associated with incomplete flushing of Niskin bottles. Based on casts outside of inlets there was a reasonable fit that resembled other recent fits for this sensor.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.20 mL/L from 0 to 150db

±0.06 mL/L from 150 to 600db

±0.04 mL/L below 600db

The PAR values are generally low because most work was done at night.

An SBE pH sensor was attached to the CTD. Users are warned that the data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The altimeter was missing from the equipment list, but otherwise the Daily Science Log Book was in good order.

There were only a few rosette log sheets, including those from 2013-35.

There are dissolved oxygen analysis log sheets for all casts, but the event numbers were restarted from 1 for Leg 2 on those sheets whereas the file names for Leg 2 start from 4 as do the entries in the Daily Science Log. During post-cruise oxygen analysis processing, a correction to the event numbers on the analysis sheets was entered for events 4 to 16. After that the analysis sheet event numbers do match the file names and Daily Log entries.

The sample numbers for each cast were started from “1” rather than the usual unique sample numbers for all samples taken.

The dissolved oxygen data were obtained in QF spreadsheet format from the post-analysis processor.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments and the dates for two calibrations were incomplete, so those were fixed. There were no changes through the cruise.

The adjusted file was saved as 2013-36-ctd.xmlcon.

There is some doubt about which dissolved oxygen sensor was used. The log gives serial #1438 and that DO sensor was on CTD #0585 in April. But it was on a different CTD on the Ricker in late June. The configuration files were incorrect for several cruises starting in late June. Processing was run with the assumption that the sensor was #1438.

##### Conversion of Full Files from Raw Data

All files were converted using 2013-36-ctd.xmlcon.

A few casts were examined and all expected channels are present. The descent rate was kept fairly high and the noise level varied from very steady to moderately noisy.

The two temperature and conductivity channels are fairly close during the downcasts and, as usual, much farther apart during the upcasts.

This fluorometer has produced suspicious data since late May 2013, but an initial examination shows that the dark values look reasonable, so the problem may have arisen after this cruise. There are no extracted CHL available to determine if maximum values are reasonable.

Altimetry has spikes near the bottom but the signal looks useful.

The DO, pH, PAR and transmissivity profiles look normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2013-36-ctd.xmlcon.

The files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. There were a few spikes in both salinity channels for cast #2, but there was no sampling from that cast, so no editing is needed.

There was bottle firing but no sampling from casts #2 and #43, so those will not be processed further. Cast #11 had no bottles fired and hence no ROS file.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, the DO analysis spreadsheet was examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-36-bot-hdr.txt; it may need further editing to reflect problems found during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-36oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2013-36oxy.csv.

That file was converted into individual *.OXY files.

The OXY files were merged with CST files in 4 steps.

Then the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The first attempt produced many outliers. Some were traced to errors in the assignment of sample numbers to Niskin bottles. The DO Analysis Log Sheets generally had the information recorded correctly but in a few cases there are mistakes; for those, it is not completely clear which bottle was sampled when 2 were fired at the same depth. The depths are recorded clearly.

Even after those corrections there were many significant outliers:

- Event #6 – DO value from 200m out of line; DO value and draw temperature suggests bottle closed well above 200m. Flag 5 with comment “Bottle did not close at intended depth.”

- Event #7 - 200m and 50m – DO way out of line but no evidence that bottle did not close at those depths. Likely this was a case of sampling from the wrong bottles as the 50m sample looks close to the CTD at 200m and the 200m bottle is close to the CTD at 420m. Both will be replaced with pad values and flag 5 with comment “Probable mis-sample.” The draw temperatures look ok though and the bottles appear to have closed as intended.

- Event #16 - Sample from 30db looks much lower than CTD DO and lower than expected at this level. But vertical DO gradient is high, so Flag 4 and comment “Outlier in comparison to CTD DO and nearby samples from similar depth.” This could be an analysis problem or poor flushing.

- Event #17 –There are 2 major outliers – one has already been flagged due to a bubble and the other had 1 replicate rejected, but it looks like the one kept is the one out of line. So the sample with the bubble previously flagged 4 will be changed to 5. The replicates will be switched – rejecting one and using the other with flag 46.

- Event #29 – The 30m bottle value is much higher than the CTD and nearby casts at 30m, but large vertical gradients make it hard to say it is definitely bad, so flagged 4.

- Event #30 – Sample #3 at 30m. Bottle lower than CTD; replicates are good. Possibly poor flushing of Niskin in very high DO gradient. No flag attached.

- Event #31 – Sample #5. The 30m bottle is much lower than the CTD; could be due to incomplete flushing in a very high gradient. No flag attached.

- Event #32 – Samples #3 and #4 – It looks like the samples labelled #3 and #4 Rep 1 somehow got switched. There are erasures in the analysis log. Sample #3 has a value close to #4 REP 2, so it was averaged with #4 REP 2 and that value was entered for sample #4 with flag 46. The sample originally called #4 Rep 1 was entered as the value for sample #3 and flagged 4.

- Event #34 – Error in ADDSAMP file. Fixed.

There are two common sources of error in these comparisons, especially in Effingham Inlet. The CTD sensor has a slow response that may not fully equilibrate when the vertical gradient is very high as is the case in the Effingham casts. There is also a problem that Niskin bottles do not completely flush especially when conditions are calm, again a factor in Effingham. So the CTD may read low and the bottle values can be low as well. A good match may not be a sign of good data and errors of either sign are both likely.

After these changes were made the OXY files were recreated and the merge process repeated. COMPARE was then rerun. Fits were made that included all casts, then just Effingham Inlet casts and finally, all casts except Effingham Inlet. Outliers were excluded, first based on large differences and then based on residuals. Trend lines were added that included an offset and then with a zero offset. Allowing an offset produced higher r2 values and looked better especially at low values. SeaBird recommend using a zero offset because they don’t trust low value titrations. When there are few low value measurements this often produces better results, but when we have many values <1 we can identify outliers at the low end and generally allowing an offset leads to a better fit even well away from the origin.

The fits found with an offset were:
CTD DO Corrected = CTD DO * 1.0673 +0.0351 (R2 = 0.9423) All casts, excluding outliers

CTD DO Corrected = CTD DO * 1.0722 +0.0300 (R2 = 0.9670) Excluding Effingham Inlet & outliers

CTD DO Corrected = CTD DO * 1.0589 +0.0275 (R2 = 0.9211) Effingham Inlet excluding outliers

The Effingham Inlet fit is very noisy with a high percentage of outliers and the slope of the fit is much lower than all other recent fits for this sensor. There are several sources of error for the inlet. In the presence of a very high DO gradient poor flushing of the Niskin bottle will lead to low values for the bottle DO so there will be outliers on the positive side for CTD-Bottle which is where most of the outliers are. On the other hand the slow response of the CTD sensor might lead to those values being low in high DO gradient areas, producing errors on the negative side of the fit. For Effingham there are many such outliers for high DO>7mL/L. Most of them come from the top 10m where the local DO gradient tends to be fairly low in Effingham Inlet, so the flushing issue is likely less significant. Variability is high in the CTD sensor near the surface. So the outliers could be due to slow CTD sensor response or just a poor match between bottles and CTD in a rapidly changing environment.

File 2013-36-recal1.ccf was prepared to apply the correction:

CTD DO Corrected = CTD DO * 1.0722 +0.03

This was applied to the SAM and MRGCLN2 files and COMPARE was then rerun to ensure this correction was appropriate and it was.

This correction produces an increase in CTD DO of about 10% at 1mL/L and 7.6% at 8mL/L; if the offset had been set to 0, the correction would be 7.9% throughout.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no significant outliers were found. No additional flags appear justified.

The comparison with bottles suggests that the DO sensor has been identified correctly in the configuration file since the fit is similar to other recent uses of this sensor.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts to see what setting brought the vertical offset between the downcast and upcast DO traces into the best agreement with the temperature traces. Since it was last recalibrated and serviced settings of 4.2s, 4.5s, and 4.8s have been used, so values between 4.5 and 5s were tested. Only one cast produced clear results with 4.8s looking best.

ALIGNCTD was used to advance the DO Voltage by 4.8s relative to the pressure.

##### CELLTM

Tests for CELLTM do not work very well in casts with many stops for bottles. The results varied somewhat and given the noisy upcast data it is difficult to choose which is best, but for the 2 casts with no stops for bottles the best choice was (α = 0.02, β=7) for the primary and (α = 0.03, β=9) for the secondary.

CELLTM was run using (α = 0.02, β=7) and (α = 0.03, β=9) for the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 deeper casts to examine differences between sensor pairs. For comparison data are included from one cast of 2013-12 which used the same equipment shortly before this cruise and one cast from 2013-13 which followed it. The differences are a little out of line with those seen during the other two cruises, but the 2013-36 traces were very noisy. In any case none of the differences are very high. Salinity differences at 1000db are within 0.0016.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2013-12-0059 | 360 | +0.0001 | -0.00007 | -0.0008 | Medium, Noisy |
| 2013-12-0010 | 350 | -0.0001 | -0.00003XN | -0.003 XN | High, Very noisy |
|  | 1000 | -0.0001 | -0.00014 N | -0.0016 | High, Very noisy |
| 2013-12-0033 | 350 | -0.0003XN | -0.0002 XN | -0.002 | High, Noisy |
|  | 1000 | -0.0003N | -0.00016 | -0.0015 XN | High, Noisy |
| 2013-12-0034 | 350 | -0.0003XN | -0.00025 XN | -0.003 XN | High, XNoisy |
|  | 1000 | -0.0003N | -0.00014 N | -0.0012 XN | High, XNoisy |
| 2013-13-0077 | 340 | -0.0001 | -0.00006 | -0.0006 XN | High, Steady |

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run. No problems were noted - the speeds look fine and the fluorescence never went off-scale.

Surface check was run and shows an average surface pressure for the cruise was 3.2db which is typical for the Tully. One file was examined that has data recorded at very low pressure with the pumps on and the conductivity is very low.

The cross-reference check was compared with the log book and errors were found in the station names for events #12 and 33. Those were fixed. There were some small discrepancies in time and positions between the log and file headers; some were close, so it is likely that the headers are fine.

Separate track plots were made for each leg; those were added to the end of the report.

There were no water depth readings in most headers, and one was wrong, but the data was available from the log book, so those were entered in the MRG and CLN files using a text editor.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for a selection of casts and the headers look appropriate.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were limited to casts with no stops for bottles. The shift of +65 records which was used for the data from 2013-12 produced good results on cast #11. The shift of +70 records which was used for 2013-13 produced reasonable results but not quite as good as +65. For the other casts checked it was hard to judge what was best.

SHIFT was run on the pH:SBE channel using a setting of +65 records.

Conductivity

During 2013-01 the shift settings chosen for the conductivity channels were -0.5 records for the primary and +1.5 records for the secondary. During 2013-16 with very shallow sampling, the choice of -0.5s looked best for both channels, while during 2013-12 the choice made was -0.5 for the primary and -0.3s for the secondary.

Tests were run on 4 casts using a variety of shifts and the best results were the same as for 2013-01.

SHIFT was run using -0.5 records for the primary conductivity and +1.5 records for the secondary.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. At times the upcast vs downcast DO traces seem closer than those for temperature, but often the opposite is true. The upcast temperature was very noisy making it very hard to judge the alignment. Overall the alignment looks ok.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late April 2011 and were used for 1 cruise in 2011 with no salinity sampling, and during 2013-01 and 2013-12 with good sampling and 2013-16 with poor sampling. Both salinity channels were found to be within 0.001 of bottles, on average, during 2013-01 and 2013-12. The 2013-01 cruise included many deep samples where low salinity gradients ensure that incomplete flushing does not have a large effect on comparisons. During 2013-13 which occurred after this cruise both salinity channels were found to be within 0.001.

Dissolved Oxygen

This sensor was recalibrated in March 2011 and had been used on 4 cruises previous to this one. For the first, DO voltage was calibrated rather than concentration. The other 3 cruises showed a slow drift in calibration. The offset was sometimes set to 0 in fits and sometimes not. For comparison, fits were made with 0 offset for all of them and those showed that the DO sensor was reading low by 7%, 7% and 7.9%.

Pressure

The sensor was recalibrated in April 2011 and was used for a 2011 cruise and 3 previous cruises. The factory offset was used for all those cruises plus one that followed 2013-36.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. For some casts there was no local climatology available. All temperature and salinity profiles fell within the local climatology except for some temperatures above the maximum near the bottom for 4 casts in about 1000m of water and 1 case of low salinity below 170m. These look like real variations rather than indications of instrumental problems. Going a little outside the 3 standard deviation criterion is not unusual so close to shore.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration - There were no post-cruise calibrations available at the time of processing.

##### CHANNEL CHOICE STUDY

There are no salinity calibration samples, but cruises before and after this one show both sensors within 0.001 of bottles. The differences between channels were not large. The primary salinity was chosen for archiving for cruise 2013-12 before this cruise and for most casts for 2013-13 which occurred after this one. An examination of a few casts showed the secondary salinity to be noisier than the primary, so the primary will be selected for editing.

##### DETAILED EDITING

The primary temperature and salinity channels were edited and all casts needed editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the cast. In a few cases some bad salinity points were removed and temperature left in place.

There were some corrupted data, especially salinity, that were associated with rapid deceleration of the CTD even though the descent rate was not <0.5m/s. This has been noted on other cruises when the maximum descent rate becomes very high. It is recommended that the descent rate be kept close to 1m/s, and typically for deep casts the maximum is 1.5m/s in rough weather. For this cruise it got as high as 2.8m/s.

All edited files were saved with extension EDT.

##### Initial Recalibration

The pressure does not appear to require recalibration.

Based on other cruises the salinity does not require recalibration.

Based on COMPARE the SBE Dissolved Oxygen was recalibrated by applying the following equation:

CTD DO Corrected = CTD DO * 1.0722+ 0.03

CALIBRATE was run using file 2013-36-recal1.ccf to apply the correction to the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles – bottles excluded in the previous comparison were excluded and there was no CTD data available sufficiently close to some bottles. The differences show that the CTD DO is higher than the bottles by an average of 0.014mL/L and a standard deviation of 0.045mL/L. The differences are fairly flat with DO or pressure, with differences slightly higher at low DO or low pressure. Given the scatter the slope of the fit does not look significant. However, because it looks like subtracting a constant value might improve this fit, a test comparison was done using an initial recalibration with zero offset (slope = 1.079) to see if that improves the comparison between recalibrated downcast DO and upcast bottle samples. It did not work as well, producing values too low at the low end of the DO range and too high at the high end of the range. (See 2013-36-dox-comp3b.xlsx)

The differences are likely explained mostly by poor flushing so that the bottles contain water from lower in the water column than the pressure indicated in the CHE files.

No further recalibration will be applied.

(See 2013-36-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined and no problems were found.

Profile plots were examined on screen and no problems were found.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH and PAR data are nominal and

unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.20 mL/L from 0 to 150db

±0.06 mL/L from 150 to 600db

±0.04 mL/L below 600db

There was no salinity calibration sampling during this cruise, but cruises

before and after this one had the same equipment and salinity was found

to be within 0.001 of bottle samples.

The maximum descent rate of the CTD was extremely high leading to very high

deceleration rates which leads to some data corruption and loss of data.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2013-36-proc.doc.

(Note: Cruises 2013-35 and 2013-36 were processed together.)

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. For most offshore casts the saturation was between 100% and 110% with a few values between 110% and 130% early in the cruise and in the outer part of Effingham Inlet. Saturations were >130% in Saanich Inlet and the inner part of Effingham Inlet and for a few near-shore casts late in the cruise. Those are typical values for the time of year.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, ph:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with different units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data. The format for the Dissolved Oxygen samples was F7.2 rather than the usual F8.3 based on the advice of the analyst.

There were many cases where bottles from which there were no DO samples, but something else may have been sampled by other groups, so the lines were left in the CHE files.

Standards check was run on all files and no errors were found other than the titrated DO format.

A header check was run on the final files. No problems were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

The CHE and CTD files for events #1 - #3 were renamed as 2013-35 and the chief scientists name and cruise number were adjusted in the files. A note was made in the headers to inform users that these files were processed together with those from 2013-36.

# Particulars

2. No bottles fired.

3. Bottles fired but only sediment samples taken. Created CHE file in case needed.

8. Missing file – CTD data likely not archived. Bottles sampled.

11. No bottles fired. No CHE file.

31. Lot of variability in salinity at the surface.

35. Problems with the winch power.

38. Winch problem power off very often.

43. Deep basin in Saanich to test winch. No bottles sampled. No CHE file created.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 5Apr2011 | Factory |  |  |
| Conductivity | 3321 | 30Mar2011 | Factory |  |  |
| Secondary Temp. | 4700 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 1766 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 2Apr2011 | Factory |  |  |
| PAR | 4601 | Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | Dec2010 |  |  |  |
| SBE Fluorometer | 2228 | n/a |  |  |  |
| Pressure Sensor | 77511 | 22Apr2011 | Factory |  |  |
| Altimeter | 16504 | 16Mar2011 |  |  |  |