## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April | ### Updated channel names & formats in TOB files. GG |
| ### 22 Nov 2021 | ### Corrected Salinity:Bottle formats S.H. |
| ### 20 Aug 2020 | ### HPLC Data added. S.H. |
| 14 July 2016 | Comment added to CHE and CTD files re fluorometer problem. See document Fluor-2228-problem.docx. |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |
| 11 Feb 2014 | Corrected order of bottle data for casts during which Niskin bottles were fired out of order. Niskin bottle #s were incorrect in raw data. See more detail in note after PARTICULARS section near end of report. GG |
| 14 Jan 2014 | Added & corrected some CHL data; replaced negative NH4 values with 0 values. GG |

## PROCESSING NOTES

Cruise: 2013-38

Agency: OSD

Location: WCVI

Project: La Perouse / WCVI

Party Chief: Yelland D.

Platform: John P. Tully

Date: 28 May 2013 – 7 June 2013

Processed by: Germaine Gatien

Date of Processing: 29 August 2013 – 23 December 2013

Number of original HEX files: 90	Number of CTD files: 89 (1 cast was aborted)

Number of bottle files: 40		Number of bottle casts processed: 37 (3 fired with no sampling)

Number of original TSG files:   13	Number of processed TSG files:   13

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#2228) with a 3X cable on the primary pump, a Biospherical QSP-200 PAR sensor (#4601) and an altimeter.

A thermosalinograph (SeaBird 21 S/N 3363) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor and a flow meter.

The data logging computer was #3.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list. There was no mention of the TSG, though it was clear from details in the log it was in use.  The Sampling log was inaccurate. There were many cases where it is indicated that samples were drawn, mostly extracted CHL, but a search of the analyst’s spreadsheet shows no indication the samples ever were taken. There is no note of a change in plan in the Sampling Log, Daily Science Log or cruise report. These errors waste time as a search is made for missing results. A simple line drawn through the X on the Sampling Log would be sufficient.

Files from the first cast in Saanich Inlet were sent to IOS for initial processing. This is a useful step as it can ensure reasonable data are being logged on all channels and that correct data are being displayed at sea. No errors were found in the configuration file, though it was later found that the fluorescence gain had been entered wrong. It was noted that there were unusually large differences between temperature and salinity channels.

Both salinity channels were found to be within 0.0010 of bottle samples.

There were unusual excursions in salinity that appear to be due to varying alignment of T and C sensors. Such variations are always seen in CTD data, but not to the extent seen in these data. These features are not due to shed wakes overwhelming the CTD since there is no effect on temperature. The problem appears to be associated with high deceleration rates of the CTD. The problem was worst for the deepest casts, perhaps because the descent rate variations are greatest for those casts, with typical values changing from -1.5m/s to -0.7m/s within 3m, so within 2 or 3 seconds. This problem may be related to the use of the LARS system. The problem was worse in the secondary salinity, so primary temperature and salinity were selected for archiving. Small excursions could generally be corrected by interpolation, but for many of these features, the salinity points were deleted. The primary temperature data looked ok, so were not removed. While these alignment variations can be significant, editing plus metre-averaging of data should minimize the effect.

During 2013-50 the CTD tech noted that the plumbing geometry was different on the primary and secondary circuits; possibly the secondary flow was not good. This could lead to further alignment problems explaining why the secondary salinity was so much worse than the primary.

Photos of the CTD set up are recommended to enable investigation later.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.6mL/L from 0 to 30db

±0.3mL/L from 30db to 150db

±0.2mL/L from 150db to 500db

±0.03mL/L below 500db

The CTD SeaBird Fluorescence values were about 34% of the extract chlorophyll values, on average.

There was some concern during the cruise that the TSG time differed from SCS time by about 4 minutes. Checks suggest that if there is an error it is a lag so that the positions are correct for the time stated. There is some confusion about the various sources of time used for log entries and data downloads.

The TSG data quality is low in many parts of the record because the flow rate was often very low. Where the flow rate is <0.6 the temperature, salinity and fluorescence data have been removed, but time, position and flow rate data were left in the files. The low flow rate appears to affect even the intake temperature.

More underway loop samples would be very useful. We get useful comparisons between TSG and CTD data, but the depths are not matched exactly, and the TSG system may operate differently in motion.

The TSG salinity data are noisy in patches and no pattern could be found that would help explain this. Clogged filters may be involved. Some editing was applied to salinity, but often it is impossible to distinguish good points from bad. The TSG salinity was reasonably close to CTD salinity while stopped and the calibration is probably reliable, but in noisy patches it can be low by up to 1psu.

The TSG fluorometer values are a little higher than those from the CTD fluorometer at about 40 to 50% of rosette extracted CHL samples.

TSG Intake temperature is available for files #5-13 but not for the first 4 files. While an estimate can usually be made of intake temperature based on comparisons of lab temperature with CTD temperature, the frequent low flow rate makes that approach inappropriate for this cruise. Even when there was an intake temperature available, there is some evidence that it is not reliable when flow rates are low.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The files from the initial cast in Saanich Inlet files were sent to IOS for initial processing. No errors were found in the configuration file, but it was noted that the differences between temperature and conductivity channels were higher than usual. This proved to be a problem throughout the cruise.

Later, it was discovered that the fluorescence gain had been set wrong; it should be 3X for all casts as that was the cable in use on the ship.

##### Preliminary Steps

The Log Book and rosette log sheets as well as various analysis logs. A few problems with Niskin bottles were noted in the log, but there appeared to be no systematic problems. Also noted was that the configuration file had the wrong gain for the fluorometer until after event #29.

Care will be needed in processing cast #32 because acquisition started too early – during the initial drop.

During events #1, 12 and 108 some bottles were fired but there was no sampling, so there is no need for CHE files.

Extracted chlorophyll, NH4, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The file creation date was added to the names of those files to avoid confusion in case some changes need to be made later. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. The only problem is that the fluorometer gain entered was wrong early in the cruise. The cable in use was 3X throughout. A file with that gain was saved as 2013-38-ctd.xmlcon. No errors were found in the calibration constants.

Since this is the first cruise using this DO sensor since it was last calibrated at the factory, the hysteresis parameter E should be fine-tuned. There was insufficient sampling of DO at depths below 2000m to enable the tests for E. Though it was likely not a critical issue for these data, a preliminary comparison of titrated DO and SBE sensor DO was run for cruise 2013-17, which followed 2013-38, had the same DO sensor and had appropriate sampling. The results showed only slight differences between corrections above and below the DO minimum and given the scatter in the comparisons the difference appeared to be insignificant. No adjustment will be made to the nominal value of E.

The PAR sensor was not always mounted. Based on log notes cast lists were prepared with and without PAR so that it will be easy to remove PAR as appropriate.

PAR on: 1-29, 51-55, 71-78, 83, 87-91, 108-109, 115-150

PAR off: 31-50, 57-70, 80-81, 84-86, 94-107, 111-114

##### Hysteresis Study

As noted in the previous section, hysteresis tests were not done as the deepest casts were to 2000db. Tests had been run for the cruise that immediately followed this one, 2013-17.

##### BOTTLE FILE PREPARATION

The ROS files were converted using file 2013-38-ctd.xmlcon.

Those files were put through CLEAN to add event numbers (*.BOT).

Header Check was run on the BOT files and no problems were found.

Temperature and salinity were plotted for all BOT files. The secondary salinity in cast #1 was cleaned lightly around 20m using CTDEDIT. The edited file was copied to BOT.

Cast #108 looks odd but there was no sampling for that cast.

Casts #1, 12 and 108 had no sampling according to the log book, so they were removed from the list to be processed.

The BOT files were then averaged to enable an ADDSAMP file to be prepared. Sample numbers were added to the ADDSAMP file based on rosette log records. A few problems were encountered:

- Cast #32 – The sampling log shows 2 bottles fired and a note “MISFIRE ON ONE”. There are 4 firings in the BOT file including two at 250m. The sampling log also has a note that Niskin #1 misfired, but sample #84 and #85 are entered for Niskin #1 and Niskin #2. That was the plan, but which bottle was actually sampled? The CHL and Nutrient values show that the sample could not have come from 250m, so must be from Niskin #3 or #4. The Daily Science log says it was a CTD cast only and that 2 bottles were tripped at 250m for ammonia standards. So those 2 bottles do not need to be processed. The samples #84 are probably from the 5m bottle, but SBE fluorescence values are very similar at those depths. For other nearby casts the intended samples at the CHL maximum were not taken, so it will be assumed that sample #84 is from 5m.  A flag 3 will be attached to the nutrients (the CHL is already flagged 3) with a note explaining that it could be from 25m, but appears to be from 5m.

- Cast #43 – complex – the sampling log says Niskin #2 didn’t fire – it did fire, but not at the planned depth. Niskin #12 did not fire as indicated on the log.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-38-bot-hdr.txt.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2013-38.xls. Event numbers were missing, so those were added. A simplified version of the spreadsheet was saved as 2013-38chl.csv which was then converted to individual CHL files. Loop data were removed from that file but were added to file 2013-38-loops.xlsx.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-38oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2013-38oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2013-38SAL.xls. The file was simplified and saved as 2013-38sal.csv. That file was converted to individual SAL files. Loop data were removed, but were first added to file 2013-38-loops.csv. The salinity data were analyzed 13 to 24 days after collection.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-38nuts.xls which included a report on precision.  The file was simplified, reordered on sample numbers and saved as 2013-38-nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2013-38_NH4.xls which included a report on precision. The file was simplified and saved as 2013-38-NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

Data from the SAMAVG files were exported to a spreadsheet and a few casts were checked against rosette sheets to ensure all expected bottle data were present. There were some cases where sampling was indicated on the sampling log sheet, but no sample was found. The spreadsheets were checked and the samples were not present, and it appears they were never actually gathered.

##### Compare

Salinity

Compare was run with pressure as reference channel. Salinity samples were analyzed within 2 to 3 weeks of collection.

There were 3 major outliers. The first involved a sample with the wrong sample #. There was a note in the salinity spreadsheet indicating that two samples had the same ID and one sample ID was missing. COMPARE made it obvious which of the samples had been mislabelled. This was fixed and COMPARE was run again.

The other outliers were:

- Sample #267, cast #84   ~750db – The bottle is higher than the CTD by about 0.005. The CTD data are not particularly noisy at firing time, but the descent rate does show that the sea state was rough. This could be a case of either evaporation of sample or incomplete flushing of the Niskin, but most likely is just due to shed wakes and variability during the stop. No flag was added.

- Sample #305, cast #100, ~125m: flagged 3 due to problem with cap. Higher than CTD by 0.02 so likely some evaporation though poor flushing could also explain this. Flagged changed to 4.

- Sample #306, cast #100, ~100m: there was no flag on this sample. Lower than CTD by 0.02. The local gradient is fairly high and there was a lot of variability during the stop, though it was low around firing time. No flag is justified.

When bottles above 500db are excluded plus one outlier, the primary salinity was found to be low by an average of 0.0005 and the secondary high by 0.0010. Standard deviations were ~0.0008 for both channels. Exclusion of a few more outliers would reduce the slope of the fits and decrease the average difference for the primary but increase it for the secondary. The outliers are almost all in the direction that suggests either some evaporation of samples or incomplete flushing. The latter is probably the explanation for shallow bottles since small vertical offsets can account for errors of this size. The samples were analyzed promptly so evaporation is unlikely.

There is some hint of time-dependence but there are not enough casts to make this convincing, especially given the noise level in the comparison. Tests of non-linearity are also inconclusive.

For full details for the COMPARE run see file 2013-38-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The fit of differences against DO concentration is close to linear if an offset is allowed. When low DO values are allowed and outliers are removed based on residuals:

DOX_BOT = 1.0471 * DOX_CTD +0.0172

It makes very little difference whether DO values <0.5mL/L or <1mL/L are included or not, if an offset is allowed. If the offset is set to 0, then it does make a more significant difference since the fit does not do well fitting values <2mL/L. Whether those bottle values are reliable is the question. For this cruise there are many low values and they fall into a tight group, so they do not look unreliable.

The comparison from 2013-17 which followed this cruise shows a fit of:

DOX_BOT = 1.0518 * DOX_CTD +0.0157

However, the values near the origin were too noisy to justify using the offset. The fit selected for recalibration of the DO was

DOX_BOT = 1.0542 * DOX_CTD

If the same sort of fit had been used for 2013-38 the slope would have been 1.0507. This suggests a little drift occurred, but the fit from 2013-17 is not sufficiently robust to support that conclusion.

There were 4 outliers clearly out of line:

- Cast #33, sample #98 – no flag or comment. There is a temperature inversion below 50db so that the temperature is about the same at 50db and 100db, but there is no reversal seen in the CTD DO traces during either downcast or upcast. The bottle DO value is very close to the 50db bottle. It is possible this was a miss-sample, but also possible that both bottle and CTD are ok, but not well matched due to local variability. The nutrients are normal – so no evidence of bottle closing late. The difference is very large and the DO value very high for 75m. Flagged 3 since could be either bottle or CTD sensor causing difference.

- Cast #54, sample #195 – flagged 3 – analyst changed this to 4.

- Cast #68, sample #200 – flagged 4 – this confirms there is a problem but is not so far out of line that it needs replacing with a pad value. Comment amended.

- Cast #68, sample #201– flagged 4 – this confirms there is a problem but is not so far out of line that it needs replacing with a pad value. Comment amended.

- Cast #125, sample #358 - 5m sample is an outlier, but this is likely to be a problem with the CTD and/or mismatch between bottle and CTD. There is no evidence for a problem with the bottle.

For more details see 2013-38-dox-comp1.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined. An error was found in cast #125 – the titrated oxygen values had been misassigned for the last two bottles. This was fixed; the error does not have a significant effect on COMPARE, so that step was not run again. No other problems were found. No further problems were detected.

Fluorescence

COMPARE was run using the SeaPoint fluorescence and the Extracted Chlorophyll from bottles.

A linear fit through the origin shows the CTD fluorescence being 34% of the extracted CHL.

A plot of the ratio of Fluorescence to extracted CHL against CHL shows that the ratio is, as usual, high for very low fluorescence (between 1 and 2) and then the ratio drops to values between 0.1 and 0.7.

For more detail see file 2013-38-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2013-38-ctd-new.con.

A few casts were examined and all expected channels are present. The descent rate was kept high, on average, but is frequently noisy enough that there is likely to be shed wake corruption.

The two temperature and conductivity channels are closer during the downcasts than the upcasts. There are some spikes and odd excursions in both channels. There are larger differences and more spikes in the top 200m than we are used to seeing. This will be studied in greater detail before deciding which T/C sensor pair to archive.

Cast #16 contains very little data; the log says the cast was aborted so it will not be processed.

Altimetry looks useful; dissolved oxygen, fluorescence, PAR and transmissivity look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 4 casts to determine the offset between the DO channel and the primary temperature. It is very hard to judge because the temperature is noisy and the results differ between shallow and deep waters, but overall a 4.0s advance looked best.

ALIGNCTD was used to advance the DO Voltage by 4.0s relative to the pressure.

##### CELLTM

The upcast data are noisy so the usual tests for CELLTM settings were a little hard to interpret. Tests were run on a few casts using a variety of settings. The differences among settings are slight.

CELLTM was run using (α = 0.02, β=7) for the primary and (α = 0.03, β=7) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are always very noisy, but these data are much noisier than usual. So these estimates are extremely rough.

The differences are not unusually high, just very noisy. For comparison, the results for one cast in 2012 are included since the same sensors were used (before the most recent servicing).

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2012-13-0017 | 1000 | -0.0004 VN | 0.00009 | 0.0014 | Noisy, High |
| “ | 1800 | -0.0006 | 0.0001 | 0.0018 | Noisy, High |
| 2013-38-0040 | 1000 | ~0 | 0.0002 | 0.0015 | High, very noisy |
| 2013-38-0080 | 1000 | -0.0006 | 0.00001 | 0.0015 | High, very noisy |
|  | 1900 | -0.0005 | 0.00001 | 0.0015 | High, very noisy |
| 2013-38-0111 | 1000 | ~0 | 0.00001 | 0.0015 | High, moderate |
|  | 1900 | -0.0002 | 0.00001 | 0.0018 | High, moderate |

The very noisy differences are not associated with high average differences. To investigate this further, plots were made of cast #111 since the descent rate was reasonably steady; differences in temperature were displayed with the 2 temperature traces and descent rate of the CTD.

In general, when differences are high due to misalignment of the temperature sensors, the sign stays the same unless the gradient changes sign and the differences are greatest where the temperature gradient is high. That does look like what is seen in the top 50m of the cast. However, the differences are varying a lot and when the traces are examined in detail there are sections where alignment does not explain what is seen. The secondary temperature appears to be drifting with a different shape from the primary. It is impossible to say one is right and one wrong, but the secondary looks unusual. There may be a combination of misalignment and some irregular flow due to pump function or plumbing configuration.

This would suggest that the primary sensors are the better choice for archiving.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

For event #32 the data from the initial drop to 10m for soaking were logged for this cast, so those data were removed from the file.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are no off-scale or negative fluorescence values.

The lowest pressure recorded was from the upcast of event #8. With pumps on and pressure ~0.2db, the salinity was about 6, so it looks as though the pressure calibration was fine.

The average surface pressure as found by the Surface Check routine was 2.5db, which is reasonable.

The mixed-layer was shallow so surface salinity bottles are not useful for calibration purposes.

The cross-reference check was compared with the log book and a few files had errors in the station name or the format of station names. Those were corrected. There are also differences between the log book and the headers. So some checks were made using the TSG files.

A few casts were selected where the log book entry for a CTD cast was significantly different from the header entry. The time was found in the TSG file that matches the time in the log and the position was recorded. The same was done for the header time. In each case the TSG position agrees with the position for that time in the log and header.

Plots were examined of TSG files at the relevant times to check on variations in latitude and longitude to make a judgment about whether the ship was stopped or not. Three casts were checked:

- Event #25 – scan #143 of TSG file #3 matches the header and scan #161 matches the log. In this case there is some movement of the ship after the header time. The longitude settles down at about the log time, but latitude continues to vary. This is the only one of the cases where there is so much movement.

- Event #99 – scan #75 of TSG file #8 matches the header and scan #109 matches the log. Both readings come from a time when the ship was stopped.

- Event #117 – scan #142 of TSG file #11 matches the header and scan #156 matches the log. Both times and positions are from a time when the ship was stopped, though there is some drift.

So the positions are all consistent with the time recorded. The header times come from just after the ship stopped in two cases and it is not clear in the 3rd. The log entries come from 9, 17 and 7 minutes later for events #25, 98 and 117 respectively.

It looks like the choice of times is not critical – the position will be right for time. At worst the header time might be slightly too early but the ship would have been moving very slowly. Given drift during the casts we never know the exact position for a cast, so any error due to this time variation is not likely to be significant.

The header entries are found very shortly after the ship stopped and the log entries are later. Data from 2013-17 were examined and the log times are consistently a few minutes after the header times and positions are very close. Since the same ship and CTD were in use, it looks like there is no problem with the header times. It appears that there was more variation in the pattern of noting log times. During 2013-17 when the same equipment was used there were only minor differences between the header and log with the log times later by only a minute or two. Header times will not be adjusted.

Another error found is that event #98 was not a CTD cast but a NET cast. The CTD cast should have been saved as #99 according to the log. The file names and event numbers were changed.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and MRGCLN2 files were exported to spreadsheets. Some casts did not get within 15m of the bottom so there are no header entries. Plots were made of a selection of casts and no problems were found in the altimetry headers of the CLN files.

The water depth header entries were checked against log entries; discrepancies of <2m were ignored and others were investigated. Using a estimate based on maximum depth sampled plus minimum altimetry, it was decided whether the log or header entry best represented water depth when the CTD was at the bottom. The header was changed were the log entry appeared a better choice.

The same changes were made in the SAMAVG files. The maximum pressure of the SAMAVG files were also checked to see if there was any sampling near the bottom. For 2 casts that had only surface samples, the altimeter heading was removed because the information suggests that the bottle was fired close to the bottom.

##### 13. Shift

Fluorescence

SHIFT was run on all casts to advance the SeaPoint fluorescence by +24 records.

After this step a few plots were checked to ensure that this setting worked well and it did.

Conductivity

An initial test on one cast showed that a negative shift is needed for the primary and positive for the secondary. Tests were run on a group of casts.

For the primary conductivity the best results were with -0.7 records, though a lot of noise remains.

For the secondary conductivity there was even more noise. The best result was with a setting of +0.3 records, but there were only slight improvements.

Two runs of SHIFT were used to apply those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The alignment looks ok.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The conductivity sensors were both recalibrated in late January 2013 and this cruise was their first use. To provide some guidance a quick comparison was done for the cruise that followed and the primary salinity was found to be low by an average of 0.0018 and secondary was very close to the bottles. The pressure dependence was very low especially in the primary. There were more outliers in the secondary than in the primary. Analysis was done within 3 to 6 weeks of collection.

2. Dissolved Oxygen

The DO sensor was recalibrated in February 2013and this was the first cruise to use it since then. A preliminary comparison of 2013-17 which followed this cruise showed the slope of the correction of the CTD DO was ~1.054 when the offset was set to 0. With an offset and excluding some outliers based on residuals the fit is:

DOX_BOT = 1.0518 * DOX_CTD +0.0157

3. Pressure

The pressure sensor was recalibrated in March 2013 and this is its first use since then.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. These are not appropriate for the casts closest to shore. The temperature data were all within the climatology except for some low temperatures for 1 cast in the northern Strait of Georgia at about 100-125db. Salinity values were outside the climatology near the surface for a few casts near shore and they were low below 200m for 4 casts in inlets. The climatology is not well suited to either of those areas. The salinity was a little high around 70db at the 3 offshore casts on the LBP line

None of these excursions suggest a systemic error, so are not evidence of calibration errors.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

From this cruise the bottle comparison shows the primary salinity to be low by an average of 0.0005 and the secondary high by 0.0010. The analysis was run within 2 to 3 weeks after collection. Observations discussed in section 10 suggest there may be some problems with the secondary channels. The primary channels will be selected for archiving.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. The descent rate was often noisy.

As noted earlier there are many odd excursions in salinity that do not appear to be related to shed wake corruption or stray spikes. These unstable features look like they are caused by misalignment of temperature and conductivity, possibly related to high deceleration rates of the CTD. The problem was worst for the deepest casts, perhaps because the descent rate variations are greatest for those casts, with typical values changing from -1.5m/s to -0.7m/s within 3m, so within 2 or 3 seconds. In studying these features it was discovered that many such excursions had already been removed at the WILDEDIT phase.

In some cases these features are too large and complex to be corrected by cleaning, so the salinity values were removed where there are large excursions. Temperature points were not removed for those features. Such editing was applied to casts: 33-49, 54-63, 71-89, 96-113 and 151.

For future cruises an eye should be kept on whether there are significant excursions in salinity that are not seen in temperature; these sometimes are two-sided excursions first towards lower salinity values followed by a section in the opposite direction. If such excursions are noted and the CTD descent rate is varying rapidly, then it may be a good idea to slow down a little, though that does run the risk of data loss due to shed wakes. On the other hand, this is not something that has been noted before, so the cause may be something to do with how the CTD is mounted so that it is more prone to the oscillating wake shedding making the CTD swing or the flow rate vary. If there have been any recent changes in the way the equipment has been mounted, this could be considered as a cause of this problem, but the CTD technician does not think there have been such changes. He has expressed concern about CTD descent rates sometimes being too high.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure looks ok.

The primary salinity was found to be lower than bottles by an average of ~0.0005 below 500m, so no recalibration is required.

SBE Dissolved Oxygen data was recalibrated using file 2013-38-recal1.ccf to apply the fit described in section 11:

DOX_BOT = 1.0471 * DOX_CTD +0.0172

CALIBRATE was run on the MRGCLN2, SAM and EDT files using that file.

COMPARE was rerun to check that the DO correction was done correctly and it was. The average difference after removing a few outliers shows the CTD DO to be within 0.0003 of the bottles on average, though the standard deviation is 0.02mL/L. (See file 2013-38-sal-comp2.xls.)

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. ALIGNCTD corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. When the differences were plotted against DO concentration, the fit looked quite flat. When outliers were excluded based on residuals, the average difference between titrated DO samples and CTD DO sensor was 0.00003mL/L but the standard deviation was 0.08mL/L. No further calibration is required.

A plot of differences against pressure was used to make a final study of outliers and to make a rough estimate of the accuracy of downcast CTD dissolved oxygen.

##### Special Fluorometer Processing

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The files were bin-averaged (0.25db bins) and put through HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved.

A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channels in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and a little more editing was done to 1 cast.

There were unstable features in some other casts, but they were mostly in areas where small instabilities are expected and no simple editing would make them stable.

##### Final CTD File Steps (REMOVE and HEADEDIT)

A check was made of the cast lists with and without PAR sensor by plotting PAR. A few errors were found and corrected.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

Profile and T-S plots were produced at this point to check for errors. Transmissivity was very low for one cast at about 15m but the fluorescence values and DO gradient were both very high at the same level, so this looks reasonable. There were small unstable features that were investigated at the editing stage – these may be real.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section, to fix the project name and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, fluorescence and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.6mL/L from 0 to 30db

±0.3mL/L from 30db to 150db

±0.2mL/L from 150db to 500db

±0.03mL/L below 500db

For details on the processing see processing report: 2013-38-proc.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged widely from 90% to almost 180%. However, most values were between 95% and 110%. Exceptions were casts that very close to shore, casts in the Strait of Georgia, casts in the LB/B/C line at sites with depths between 100 and 150m and a few casts in Queen Charlotte Sound. The higher DO saturations are generally associated with higher fluorescence and a check of a few casts with particularly high saturation show that the DO bottle and CTD DO are close. So the high saturation is likely due to high productivity and not poor calibration of the DO sensor.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. The only problems noted are that some samples shown on Sampling Log sheets are missing. There were no notes anywhere about a decision to not take a lot of the CHL samples and some NH4 samples. A simple line through the X on the prepared sheets would have been sufficient. As it is, the spreadsheets, cruise report and DO Analysis log sheets were all checked to ensure that data had not been lost.

A header check was run and no errors were.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found.

##### 25. Thermosalinograph Data

Data were provided in 13 hex files.

The external thermistor was not connected for the first 3 files.

There were 2 loop salinity samples taken while the ship was underway.

There were problems at sea with the TSG and the operator thought the XMLCON file might be corrupted.

The following note was included with the raw TSG data:

Started logging on SCS June 1, 2013, approx 10:30 local.  It seems that the NMEA time on the TSG is almost 3 minutes fast. At first, the TSG computer was on local time, reset to UTC at 17:56 UTC (check previous files, probably local time).

No explanation for NMEA time error, compared to GPS time (where it's getting it) and the error is there.  GPS taken as correct.

SCS time stamps are from a corrected clock on Science Server, high confidence that it is correct.  It's unknown, however, whether the positions associated with the 'incorrect' time are also incorrect (a lag on all GPS data?).  This needs to be checked by correlating positions and times while steaming.

a.) Checking calibrations

There were many stops to clean filters – hence the large number of files.

The configuration files for the 13 casts are identical and the parameters are correct.

However, there is some confusion about this since there is a comment in the header of file #5 saying a new con file was being tested. And while the cruise was underway there was a communication from sea asking for a new con file because the operator thought the one they had must be corrupted.

There is one difference between the con file sent while the cruise was underway and the ones associated with the hex files, and that is that the “NMEA time added” box was not ticked in the latter.

The header of file #5 says a new con file was tried, but there are no differences among the con files provided; so either the XMLCON file was changed after acquisition, or the file was overwritten but the header comment not changed or they found the one already in use was best.

There was one other con file in the folder which is missing the external sensors.

Perhaps some con files were changed after acquisition, but they all include the external sensors.

The correct file was saved as 2013-38-tsg.con.

b.) Conversion of Files

The files were converted to CNV files using configuration file 2013-38-tsg.con. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Rough checks were made of times in the files because of some comments about problems. Comparing to CTD times and positions indicates that times are roughly right in files #2 and 13. File #1 is harder to judge. The TSG was started while the ship was in Saanich Inlet. It did not stay there long, so this is presumably after the cast at Saanich Inlet. That cast ended at 02:35 and the TSG file starts at 03:36. Did the ship immediately leave Saanich Inlet after the cast, or spend an hour in the inlet? If it left immediately, then there is a 1 hour error in the TSG file. If it stayed around for an hour, and turned the TSG on only towards the end of that hour, then the time could be correct. There are no other CTD casts that occur during the time of file #1. There is no SCS data from that time period. The note mentioned above suggested that the times were local and that does not seem to be right for file #1. The chief scientist cannot remember how quickly the ship left the inlet, so it will be assumed the time is correct.

Based on the note mentioned above, a comparison was made of SCS and TSG times/positions. It showed little difference between them and between them and the log book entries when the ship was stopped. While moving there are significant differences. For example:

June 1 - underway after event #58 - 2013/06/01 09:00:05

TSG (File #5)    	48 33.4896       126 34.0260

SCS		          	48 33.4224       126 34.1518

Looking through the SCS file the position that best matched the TSG position was found about 40s later.

06/01/2013 09:00:46.     48 33.4893       126 34.0261

It was noted at sea that the TSG time was about 4 minutes later than the SCS time. This may indicate that there was a lag in both time and positions in the TSG system, so that the recorded information is accurate. The TSG only records every 30s, so the TSG clock may be running fast relative to the SCS clock, but the difference is not significant.

Time-series plots were produced. There are large sections with low or zero flow. There is no intake temperature until file #5. Salinity is often noisy in patches, with no obvious link to flow rate, temperature or fluorescence. There is one spot where the noise settled down after the flow was stopped briefly. That was probably done to clear the filter, so clogged filters might be implicated in salinity noise.

File #1 - just a short drop-out which has little effect on the data.

File #2 – steady decrease in flow rate ending at about 0.56. There is a drop-off in fluorescence values as the flow reaches about 0.7, but this is in an area where we expect a drop-off as the ship moved offshore on the LB line.

File #3 – There is a low flow rate (0.5-0.6) until very close to the end. The fluorescence shows a large drop between LB9 and LB10, but the CTD fluorescence shows the same drop.

File #4 – The flow drops off suddenly around station A2 and the variable traces change suddenly too.

File #5 – The flow went to zero while ship was stopped for cable repairs at LD11. It was off for about 6 hours. As the ship headed landwards fluorescence jumped upwards at about LD06.

File #6 – No problems with flow but salinity is noisy. There seems no pattern to the noise – not correlated with particular ship heading or other variables.

File #7 – Flow ok except for one spike – salinity noisy.

File #8 – Some spikes and noise in salinity.

File #9 – Noisy salinity at end.

File #10 – Noisy salinity at end.

File #11 – Flow is ok, but there are huge differences in temperature and variability in salinity, but this is in inlets, so not unreasonable. Very small difference in depth of sample would have large effects, and even at the same depth probably lot of variability.

File #12 – Flow rate drops in second half and all data flatten somewhat.

File #13 – Flow zero second half – data bad.

The track plot looks fine and was added to the end of this report. There is gap between files #1 and 2.

c.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2013-38-ctd-tsg-comp.xls.

70 CTD casts with data from 4m overlapped with TSG data. There were 13 TSG files, 1 of which had no overlap with CTD casts. 12 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. Normally these differences are small, but as noted above, there are some differences in time between the log and CTD headers. A plot was made of the differences in the two times versus the difference in positions between CTD and TSG data. Outliers are more common when the time differences are higher, but not all time differences correspond to outliers; in particular the largest time difference has only a slight difference in position. There is no evidence that the time in the CTD headers and TSG records are significantly out of sync, but a small error can lead to large differences if the ship is drifting. Four outliers identified were excluded from the studies that follow.

The differences in latitude and longitude were all <0.0004° and the median differences in both are 0.00002° when the outliers are excluded. But even those differences are not huge. This shows both the times and positions are reliable for both systems

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Rosette samples and TSG and CTD data

For these comparisons the 4 outliers identified in the previous section were excluded as were CTD casts 8 to 21 (near-shore), 128 to 142 (inlets) and the final 149-151 casts (in the Strait of Georgia.) This left some sections of data where differences were reasonably steady. The casts with very low flow rates were among the excluded casts. The flow does vary considerably, so plots will be made to see how they affect the data

- T1 vs T2 The intake temperature was not available until file #5. The average difference between the intake and lab temperature was 0.56C° and the median was 0.27 C° when all data are included. A plot of differences against cast # shows that there was a quiet section between casts #80 and #103. When only those casts are included the average difference is 0.26C° and the median 0.25C°, which are both close to the median for the whole data set.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

1. Intake Temperature The difference between the TSG intake temperature and the CTD temperature was noisy and was only available from cast #54 onwards. For the quiet section between casts #80 and #103 the lab temperature is higher than the CTD temperature by an average of 0.005C° and a median value of 0.004C°. When restricted to flow >0.8 and excluding outliers, the median is also 0.004C°.

2. LAB TEMP Using the same data as in the analysis of the intake temperature the lab temperature is found to be higher than the CTD temperature by 0.266C° When more data is included from before the intake temperature was available the results are similar with the average being 0.276C° and the median only slightly lower at 0.266C°. This is consistent with the differences between the two TSG temperature channels discussed in the previous section. When only casts with flow rate >0.8 are included, the median difference is a little lower at ~0.255C°, and when only flow rates <0.6 are included the median difference is 0.314C°. This is surprising since it would seem that the water would spend longer in the loop so it could warm more towards the ambient temperatures of the ship, and the water is likely drawn from a little deeper where the water is cooler. Does something go wrong with either the intake thermistor or the intake itself when the flow rate is very low? Could the water be warmed somewhat before it is measured by the thermistor?

These observations tell us about the TSG when it is stopped. There are no cases where intake temperature was available during a CTD cast with low flow rate except when that rate was zero. So a few files were examined in detail to see how the differences vary while underway. The traces are extremely noisy but in a few quiet sections with a high flow rate the lab temperature was found to be higher than the intake temperature by about 0.22C°. A few other sections were found with similar temperatures but low flow rate, and the differences were lower at about 0.20C°. So for both stopped and underway observations, there is less heating in the loop when the flow rate is low.

3. SALINITY  TSG salinity data are extremely noisy, especially in the inlets. Using the same data as used in the analysis of the intake temperature, the TSG salinity is lower than the CTD salinity by an average of 0.108 and a median value of 0.0.062. When a larger group of casts are included based on low salinity variability, TSG salinity is lower than the CTD by an average of 0.078 and a median of 0.035. When only casts with flow >0.8 are included the TSG values are low by a median of 0.04, but when only flows <0.6 are included it is low by 0.023. So the lowest differences are associated with lower flow rates explaining why the differences are lower when more data are included. Higher flow rates may lead to water being drawn from higher in the water column. The noise level and variable flow make estimates of salinity error very rough.

4. FLUORESCENCE Again, using the data from part 1 of this section, the TSG fluorescence is 1.12 times higher, on average, than the SeaPoint fluorescence from the CTD with a median ratio of 1.01. If more data are included the TSG fluorescence is relatively higher because earlier casts have a higher ratio of TSG to CTD fluorescence, and those were not included in the cast #80-103 analysis. The maximum ratio is 2.3. There does not appear to be a relationship between the ratio and the CTD fluorescence. There does appear to be a relationship to flow rate with the TSG closest to the CTD values when the flow rate is high. That could be due to flow rate affecting the depth from which the TSG water comes. However, the data with low flow rate is limited and all the data are very noisy.

(See 2013-38-ctd-tsg-comp.xls.)

Loop Bottle - TSG Comparisons There were 2 loop samples for salinity and 1 for extracted chlorophyll.  These are the only comparisons possible for the TSG when it is underway. The TSG salinity was lower than the CTD salinity by 0.055 and 0.050. The TSG fluorescence is 1.155ug/L while the extracted chlorophyll sample is 1.53ug/L; the ratio FL/CHL is about 0.76.

(See 2013-385-loops-tsg-comp.xlsx.)

- Surface rosette CHL vs TSG fluorescence

There were no near-surface salinity samples during this cruise. CHL samples were available.

A comparison was made between the TSG fluorescence and the Niskin bottle CHL data. There were 23 cases of samples taken during TSG acquisition. The match is not expected to be great because the times for the TSG data are from the start of casts, not the time of the rosette sampling, and the TSG samples are likely from higher in the water column than the rosette samples. The TSG fluorescence is found to be close to the CHL on average, but this is because it is much higher for CHL<0.2ug/L but only about half the CHL value for CHL>0.2ug/L. It has often been noted that fluorometers read relatively higher at low CHL values.

(See 2013-38-rosette-tsg-comp.xls.)

- Calibration History

The TSG pressure, temperature and conductivity were recalibrated in March 2012 and it was used during 2013-01. It may have been used on other cruises from which we have not bent sent data.

There were no data available from the intake thermistor for that cruise and there were many problems with the flow rate.

The heating in the loop while stopped was estimated to be about 0.28 Cº by comparing lab temperature with CTD temperature. The value is similar to others found in the past at that time of year. The heating is dependent on intake temperature.

CHL was low for 2013-01 with no values >1ug/L. When the lowest values are excluded the TSG fluorometer values were about 10-15% higher than those from the CTD fluorometer. When all data were included the TSG was high by about 20%.

The TSG salinity was unreliable for much of the cruise. It was slightly lower than loop salinity when the ship was stopped and slightly higher underway. There was insufficient data to enable recalibration.

Conclusions

1. The TSG clock appears to have worked well. There are some odd differences between log times/position and NMEA times/positions in the CTD files, but this does not appear to have affected the TSG data.

2. The flow rate drifted downwards during some files, with values <0.6 in files 2, 3, 4 and 12. There are patches of zero flow in file #5 and 13. There are a few other very brief spikes to zero; probably the system was turned off on purpose to clean the filter.

3. The temperature in the loop increases by about 0.25Cº based on comparisons during stops for CTD casts, but this will vary according to intake temperature and flow rate. Underway estimates of heating in the loop are very rough due to very noisy traces. When the flow rate was high the lab temperatures are higher than intake temperatures by ~0.16Cº in an area with high intake temperatures to ~0.25Cº in cooler waters. These values are somewhat lower than those found when stopped, perhaps due to water coming from a little shallower while underway.

4. The TSG intake temperature appears to be higher than the CTD temperature by ~0.004Cº when outliers are excluded. Given some mismatch in time and depth of sampling between the two sources, recalibration is not justified. There are no external temperature data available until file #5, so we would normally use the estimate of heating in the loop to create a proxy for intake temperature. However, the varying flow rate makes this inappropriate. The Temperature:Intake channel will be removed from files 1-4.

5. The differences between the TSG Salinity and that of the CTD are very noisy, but the TSG appears to be low by a median of ~0.04 when the flow rate is high and by about 0.02 when it is low. During 2013-01 the salinity was found to be within 0.005 of CTD salinity. Underway loop samples suggest that the TSG salinity is low by 0.05 or 0.06, but there were many problems with noise in the data set. No recalibration is justified, but if the sensors are recalibrated, this can be revisited.

6. The TSG fluorescence is about 50% of the CHL samples from the rosette. The TSG fluorescence is about 75% of the one underway CHL sample (~1.5ug/L.) The TSG fluorescence data are higher than the CTD fluorescence by up to 20%. The CTD fluorescence was found to be about 34% of the rosette CHL samples, which would imply that the TSG fluorescence is about 40% of the rosette CHL. More loop samples would be useful.

f.) Editing

The ATC files were copied to *.EDT.

CTDEDIT was used to clean salinity, though in many sections the salinity was too noisy to edit.

There were some interruptions in the flow in the loop. Some of these were probably deliberate, short interruptions for the purpose of cleaning the filter, and for one long interruption the ship was stopped for cable repairs. Examination in CTDEDIT was also used to determine which records had flow <0.6. Then Ultredit was used to pad the temperature, fluorescence and salinity values for those records. (This is hard to do using CTDEDIT because we want to leave the record in place with the time and position, but we need to remove data from more channels than can be done in a single pass through the editor.)

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

The flow channel was not removed since it explains the gaps and some users may want to remove even more data based on this variable.

The intake temperature channel was also removed from files #1 to 4.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

PAR on: 1-29, 51-55, 71-78, 83, 87-91, 108-110, 115-150

PAR off: 30-50, 57-70, 80-81, 84-86, 94-107, 111-114

1. All bottles closed – no sampling.

12. One Niskin closed for bulk sampling – no sample #s assigned.

17. Rosette didn’t release #16.

29. Con file changed after this cast – according to the log the fluorometer gain should be 3X, not 30X.

32. Acquisition started too early – during initial drop.

96. Sounder and altimeter problems – stop at 1240db went up 10m and then down again slowly to 	1281db.

98/99 – CTD file saved as event #98 but should be 99.

108. Niskins 8, 9, 11, 12 closed as test – no sampling.

Corrections to CHE files – 11 February 2013 – G. Gatien

Errors found in 2013-38 bottle files. These were cases where bottles were fired out of order – generally that was not clearly indicated on the rosette sheet, though it can be figured out. The ROS files do not reflect this out-of-order firing. No reason was found for the errors in Niskin Bottle #. The channel name was correct but the data are clearly wrong.

There are 7 casts where bottle data are associated with the correct sample number, but the wrong CTD data and Niskin bottle #: 4, 8, 34, 35, 41, 76, 100. The sample numbers and bottle data were shifted to the correct lines. The nutrient profiles are more believable after this step and the CHL samples are at the depths listed on the sample labels.

Cast #48 also has an inconsistency but of a different sort – the rosette log indicated that there were 2 bottles fired at 5m but only 1 was actually fired. There is no note about this on the rosette sheet, but in the CTD Daily Science Log it is stated that bottle #8 was not fired. The CHL sample was probably taken from Niskin #7 but left with the sample number originally intended. A note was added to the header to explain this. So this is appears to be an error in the rosette log, not in the data file.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 2280 | 29Jan2013 | Factory |  |  |
| Secondary Temp. | 2668 | 1Feb2013 | Factory |  |  |
| Secondary Cond. | 2754 | 30Jan2013 | Factory |  |  |
| Transmissometer | 1396DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 16Feb2013 | Factory |  |  |
| SeaBird Fluorometer | 2228 |  |  |  |  |
| PAR | 4601 | 16Mar2011 | IOS |  |  |
| Pressure Sensor | 63507 | 15Apr2011 | Factory |  |  |
| Altimeter | 1204 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2013-38

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 7Mar12 | Factory |  |  |
| Conductivity | 3363 | 7Mar12 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | Aug12 | IOS |  |  |
| Temperature:Secondary | 0603 | 7Mar12 | “ |  |  |
| Flow meter | ? | n/a |  |  |  |