## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 14 July 2016 | Comment added to CHE and CTD files re fluorometer problem. See document Fluor-2228-problem.docx. |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2013-50

Agency: OSD

Location: WCVI, Douglas Channel, Effingham Inlet

Project: WCVI & Northern Inlets Moorings

Party Chief: Juhasz T.

Platform: John P. Tully

Date: 27 June 2013 – 10 July 2013

Processed by: Germaine Gatien

Date of Processing: 5 November 2013 – 17 January 2014

Number of original HEX files: 25

Number of CTD files:  25	Number of bottle casts:	25

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#2228) with a 3X cable on the primary pump, a Biospherical QSP-200 PAR sensor (#4615) and an altimeter.

The deck unit was a Seabird model 11.

The data logging computer was #2.

All casts were run with the LARS mid-ship station.

The oxygen kit was #1 (IOS Dosimat Model 665)

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS rosette with 24 10L Seabird/Homemade bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had an equipment list. There is no information about the TSG and no TSG data were provided, so it is assumed it was not in use. However, it is not unusual to have TSG data mislaid, so it is helpful to have that section crossed out in the log when it is not in use.

The Daily Science Log Book had many comments entered, but the writing was sometimes hard to read and because of the placement of comments it was not always clear to which events they referred. In one case the comment about a Niskin being fired accidentally was listed by one event in the Daily Log and by another in the Rosette sheets. It is a common problem on mooring cruises that the comments for the mooring retrievals take a lot of space on the blank pages. If there is room on the blank page but not right beside the comment, be sure to use the event # or a symbol to relate the comment to the main entry. If you have comments to enter and don’t have space it is recommended that you pick a page near the back of the book and make the comment with the Event #. Leave a brief note on the main page saying that there are comments at the back of the book.

There were some inconsistencies between entries in the rosette log and Daily Science log and in some cases the sampling indicated on the rosette log does not appear to have occurred; the samples were either not delivered for analysis or were mislabelled. In other cases no sampling is indicated but clearly did occur. Many samples were shown as duplicates but were not reported as such by the analyst, so it is assumed that duplicates were not taken. Some sample numbers were used twice. To fix that a leading 9 was added to sample numbers for event #15. There were also errors in the bottle labels for event #15, so that they differed by 1 from those shown on the rosette log sheet. For event #15 there is confusion about whether nutrients came from even or odd-numbered Niskin bottles. It has been assumed that the sample numbers were correct, but since they were incorrect for salinity this may not be so. All these problems lead to less confidence than usual in the assignment of bottle analyses to particular bottles in the CHE files, particularly for event #15. Sampling was simpler later in the cruise and no problems were noted in the rosette log after event #23.

The most serious problems occurred early in the cruise probably because of a change in plans on how to number samples when two bottles were fired at each level. The practice of preparing the rosette sheets and labels early makes it very difficult to keep good records when changes are made. When changes have to be made to the sampling plan it would be wise to use sample numbers that are completely different from the original plan and take time to prepare a new rosette log sheet.

Another perennial problem with mooring cruises is the confusion in entries as to which time zone is in use. PDT was used early in the cruise and UTC later. The rosette logs were clearer with both times listed.

An initial attempt to convert the raw files failed because most hex files had errors in the headers. Some had carriage returns that were fixed using a repair routine. Most had errors that were fixed by using a text editor. There were long lines of 0 entries, header entries without a *, missing END statements and cases where the depth was 2 lines below the appropriate header.

There were many comments about problems with Niskin bottles – leaky spigots and misfires. There was also a comment about differences between the plumbing on the primary and secondary CTD pumps.

Salinity samples were analyzed within 13-26 days of collection. There were strong near-surface salinity gradients. The comparison between CTD salinity and bottles showed both salinity channels to be within 0.001 of the bottles, on average, but there was a lot of scatter in the comparison. In many cases it appears that there was incomplete flushing of Niskin bottles and other outliers may be due to high local variability or noise in the CTD data.

The analyst recommended that the dissolved oxygen samples be trusted only to the 2nd decimal place due to relatively low precision.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.25mL/L from 0 to 100db

±0.1mL/L from 100 to400db

±0.05mL/L below 400db

As usual CTD fluorescence values are nominal, but users should be aware that the dark values from this fluorometer have been unusually high since late May 2013 and increasing with time. The fits of CHL versus fluorescence for this sensor have also shifted from previous years, with fluorescence generally lower than expected. There are few deep casts and little chlorophyll sampling from this cruise, but the trend to high dark values appears to have continued.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure, conductivity and DO sensors were obtained.

Next the calibration constants were checked for all instruments. The configuration file did not change during the cruise. The fluorometer cable gain is not entered in the log book. The configuration file indicates that is was a 10X cable which was what was in use for the 2nd half of cruise 2013-17, so that makes sense.

It is assumed that the PAR was on for all casts since there is no note of a change and no casts were deeper than 700m.

##### BOTTLE FILE PREPARATION

Hysteresis tests were run during 2013-17 and no adjustment of factor E in the DO calibration was found necessary.

The ROS files were converted using file 2013-50-ctd.xmlcon.

Those files were put through CLEAN to add event numbers (*.BOT).

Temperature and salinity channels for the BOT files were plotted and no significant outliers were found.

Header Check was run on the BOT files and the only problem noted was in fluorescence; as noted earlier it has high values at depth.

The BOT files were then averaged to enable an ADDSAMP file to be prepared. Sample numbers were added to the ADDSAMP file based on rosette log records. A few adjustments were made in this process based on log notes and Chief Scientist’s notes:

- Event #18 – Niskin #1 was fired but no sample taken. The line was removed from the ADDSAMP file.

- Event #35 – Niskin #2 was fired but no sample taken. The line was removed from the ADDSAMP file.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged

to produce SAMAVG files.

The addsamp.csv file was sorted on Event_Number and Sample_Number and then converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-50-bot-hdr.txt; it may need further editing to reflect problems found during processing. The * in the spreadsheet file names that follow stands for the date the file was created. These names are updated as corrections are made.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2013-50chl*.xls. The file included comments and flags and an event-number column. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2013-50chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-50oxy*.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2013-50oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2013-50SAL*.xls. The file was simplified and saved as 2013-50sal.csv. The salinity data were analyzed within 13 to 26 days of collection. The simplified spreadsheet was converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-50nuts*.xls which included a report on precisions. There were no station names or event numbers in the file, only sample numbers. So the other details were added based on the rosette log information, though there were inconsistencies in records, so this was considered a “first guess”.  The file was simplified, reordered on sample numbers and saved as 2013-50nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

Note: The rosette log indicates there was an NH4 sample taken, but that was likely part of the carbon sampling for which delivery of results is always delayed. It can be added later.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s.

Those files were merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The MRG files were exported to a spreadsheet for an initial check that all samples had been added and that the assignments of event numbers and station names were correct. There were huge problems, most notably with cast #15, but also with a few other samples. There are many inconsistencies between the various logs and labels for 2013-50, and the writing is often difficult to read.

Some problems were expected because there was a duplication of some sample numbers. The analyst who prepared the salinity and dissolved oxygen spreadsheets added a leading 9 to the sample #s from event #15. That is a good solution so was applied to the nutrients. There was no CHL sampling from that event. But the problem goes beyond the duplication of sample numbers. There is inconsistent information available about sampling from event #15, including:

Rosette Log –Samples from every bottle, sample #s 24-43 (these were renamed as 924-943) and dups from #32. Two bottles fired at each depth. Nuts shown as coming from odd numbered Niskin bottles and DO from even bottles. All Nuts are duplicates. DO sample #37 is duplicate.

Daily Science Log – 20 bottles with sample #s 25-34 which would only make sense if the same sample number were given to the 2 bottles fired at a single depth. That has been done in the distant past, and perhaps someone believed that was the plan. The original sample # entries had been erased and replaced on the rosette log – presumably the rosette sheet matched the CTD Log entries before those changes.

Oxygen analysis log – samples from even numbered Niskin bottles with odd numbered sample #s 25-43 and duplicate from sample #37. So this agrees with the rosette log.

Another case of confused records is event #6. The salinity analysis sheet indicates that two samples were taken from station SC01 (event #6), one of which had no sample number but was labelled “bottle 13”. The rosette sheet indicates that there was sampling from every Niskin bottle at that station. So either many bottles have been mislaid or plans changed.

Pulling together the information from analysis spreadsheets is very difficult given so many discrepancies. The salinity spreadsheet is the easiest to work with since we now have merged files with CTD pressure and many salinity samples to compare with the CTD salinity. For event #15 it appears that:

- The bottle with the sample # missing does appear to have been assigned to the right spot based on comparison of bottle and CTD salinity (they differ by almost 3psu, but that is not terribly unusual right at the surface.)

- The raw samples said to be from DC09 were numbered 25-44 with duplicates for 33. So the duplicate has a sample number higher by 1 than that indicated in the rosette log. Looking at the salinity values we expect them to be close in pairs 924/925, 926/927 etc. But while the CTD salinity values are close, the corresponding bottle samples are not. The sample given as #925 is likely the missing sample #924. And sample #944 is similar to #943 and similar to the two surface CTD salinity values for event #15, so they appear again to have sample numbers higher by 1 than the rosette log sample numbers. The analyst noted that #44 was not on the rosette log, but thought it might have been added. But no extra bottle was fired. If we just assume the sample numbers on the salinity bottle labels were all too high by 1, the values make sense for the depths sampled. Otherwise the pairs from a single depth don’t make sense. The salinity duplicate analysis of samples from repeat Niskin bottles is not affected by this error since the bottles were paired appropriately in that part of the spreadsheet.

- There may be an error in the nutrient spreadsheet like that in salinity. The rosette log indicates that nutrients were taken from odd-numbered Niskin bottles, but the sample numbers show the opposite. There is no way to distinguish these. Since there is a special note that the oxygen samples came from even-numbered Niskin bottles and it is usual during this cruise to take nutrients and oxygen from different bottles when 2 were fired at a single depth, it will be assumed that the nutrient sampling is shown correctly on the rosette sheet.

A few other errors were discovered in the event numbers of the dissolved oxygen spreadsheet. Those were corrected as well as the problems with sample numbers established in event #15 and the merge process was repeated.

##### Compare

Dissolved Oxygen –

During 2013-17, which preceded this cruise, there were problems with the dissolved oxygen data, but it was believed that this was related to its being mounted on the secondary pump. For this cruise the log indicates that the sensor was on the primary pump.

COMPARE was run and the initial plot of differences between the titrated DO and SBE DO versus SBE DO has many outliers, with many cases of the SBE DO reading higher than the bottles. The data come from 3 distinct areas, Northern inlets around Douglas Channel, Effingham Inlet and 2 offshore casts west of Haida Gwaii and Brooks Peninsula. So it is not a surprise that plotting these casts in groups leads to different results. Most of the low DO values come from Brooks Peninsula and the western part of Effingham Inlet. The fits using just Douglas Channel and Brooks Peninsula are reasonably tight and look normal. The fit in Effingham Inlet looks random. When outliers are gradually reduced and fits forced through the origin, the following fits are found:

All casts

DO_Bottle = DOX_CTD *1.0545	R2=0.73 using 75% of data

Douglas Channel

DO_Bottle = DOX_CTD *1.0548 	R2=0.51 using 90% of data

Brooks Peninsula

DO_Bottle = DOX_CTD *1.0593 	R2=0.75 using 86% of data

Douglas Channel + Brooks Peninsula

DO_Bottle = DOX_CTD *1.0545 	R2=0.59 using 89% of data

Effingham Inlet

DO_Bottle = DOX_CTD *1.0526 	R2=0.63 using 58% of data

The value found that includes all casts produces results in good agreement with other areas, but almost half the data were excluded from the Effingham casts. The fit is close to those used for the previous 2 cruises that used this sensor:

DO_Bottle = DOX_CTD *1.0471 + 0.0172	 (2013-38)

DO_Bottle = DOX_CTD *1.054		 (2013-17)

DO_Bottle = DOX_CTD *1.0518 + 0.0157	 (2013-17 allowing offset)

For 2013-17 the offset did not look appropriate as there was a lot of scatter in the fit.

SeaBird recommends forcing the fit through the origin because of the unreliability of titrations below 1mL/L. But sometimes when there is a good supply of low DO bottles, including them in the fit produces a more convincing fit even above 1mL/L, as seen above for 2013-38. There are real sources of offsets, so if the low titration values look reliable – not much scatter - it seems wise to use them. So a fit was tried with all casts, using the same data as previously, but allowing an offset. The result was:

DO_Bottle = DOX_CTD *1.0504 +0.0165	R2=0.74 using 75% of data

The offset is close to that found in 2013-38 but the slope is higher than 2013-38 and lower than 2013-17. There is not much difference between the fits, so the zero offset will be used as recommended.

A plot against pressure shows that most of the severe outliers are from the surface area and a plot against file pair number shows that the cases of the SBE DO reading higher than bottles are mainly from Effingham Inlet where gradients are high. The SBE DO sensor is generally expected to read low and calibration drift over time leads to even lower values. High values are mostly due to poor flushing of Niskin bottles in high DO gradient waters. Such flushing problems are likely worse in calm waters. Where there is a sub-surface DO maximum there could be high values due to slow DO sensor response, which may explain why a lot of outliers come from 10m. There are other outliers due to analysis problems and noisy SBE DO data, though we expect those to be more random in sign.

For more details see 2013-50-dox-comp1.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no further problems were detected.

Salinity –

Compare was run with pressure as reference channel. There were many outliers in the top 6db. There are many potential causes for these differences. There have been problems in the past with delayed analysis leading to bottle values that are too high, but this would affect all levels and analysis was done fairly promptly for this cruise. If the problem is due to poor flushing of Niskin bottles in the presence of a high salinity gradient, then we would expect the CTD to read lower than the bottles, and that is true in the majority of cases. If flushing is complete, then the vertical separation between the CTD and Niskin would lead to the CTD reading too high, which could be significant in a large salinity gradient. The near-surface gradients are generally high in the inlets. Finally there may be rapid changes in salinity near the surface which may account for some of the outliers in either direction.

For cast #26 the bottle value was ~29.4 and the CTD value ~20.6. The bottle value corresponds to CTD salinity at ~12db of the downcast and ~10db of the upcast. The bottle value is the average of duplicates that were very close, so it looks reliable. There does not appear to have been much change between the downcast and upcast, and the standard deviation in the CTD data during the 10s window centred on firing time is very small. This leaves the explanation that the Niskin did not flush well.

The only case of a difference >0.2 that is below 6db is from 20db. Two bottles were fired and differ by ~0.5 and the primary and secondary CTD salinity differ by ~0.6, suggesting a lot of variability.

The differences are very noisy even when values above 100db are excluded. There is clear pressure dependence. A fairly flat fit against pressure was achieved by excluding points above 100db and differences >0.005. The deep points with differences >0.005 included 3 bottom bottles and 2 from higher in the water column. It is not unusual to see bottom bottles as outliers, which may be due to gradients related to bottom currents or sediments in samples. The average difference using this small selection was -0.0005 for the primary CTD salinity and +0.0006 for the secondary. The standard deviation in both sets was ~0.002.  The median differences were -0.0003 and +0.0009. These results are close to those found during 2013-38 in Late May. During 2013-17 in early June the CTD read relatively lower but analysis was not as quick so bottles may have been a little high due to evaporation.

For full details for the COMPARE run see file 2013-50-sal-comp1.xlsx.

Fluorescence

COMPARE was run using the SeaPoint fluorescence and the Extracted Chlorophyll from bottles. There were only 3 casts with CHL sampling and the range of CHL was from ~0 to 5.7ug/L.

For extracted CHL values <0.1ug/L the ratio of CTD Fluorescence to Extracted CHL is extremely high but falls rapidly as CHL increases, reaching a value of 1 when CHL was 4ug/L and 0.9 for CHL=5.7ug/L. The trend would suggest that the fluorometer would read too low for higher CHL. This resembles other comparisons since May 2013, but differs from those previous to that time. There are concerns about the performance of this fluorometer.

For more details see file 2013-50-fl-chl-comp.xlsx.

##### Conversion of Full Files from Raw Data

An initial attempt to convert the files failed because all but 4 hex files had errors in the headers. Some had carriage returns that were fixed using a repair routine. Most had errors that could only be fixed by editing. There were long lines of 0 entries, header entries without a *, missing END statement, depth on one line and the depth entry 2 lines below.

After repairs all files were converted using 2013-50-ctd.xmlcon.

A few casts were examined and all expected channels are present.

The fluorescence data look bad. The dark values are unusually high at about 1.6ug/L at ~650db. A revisit to the data of 2013-38 and 2013-17 shows that the dark value steadily rose through May and June ending with values of >1.5ug/L at 2000 to 4000db. So it looks like there was something wrong with this sensor. There are also notable differences between upcasts and downcasts, which were also seen in 2013-17. The SBE fluorescence is always stated to be nominal, but for this cruise it must be treated as less reliable than usual.

Since the fluorescence cable choice was not entered in the equipment list, tests were done to see if some of the odd fluorescence behaviour might be due to an error in the gain used in the configuration file, 10X. Tests were run on cast #30 using all possible gains to see if reasonable values could be obtained. That cast was deep (~700m) and there are extracted chlorophyll samples available from the top 50m. The extracted CHL, dark values and the SeaPoint fluorescence obtained using the different gains were:

| Gain | CTD FL Dark value in ug/L | CHL at 20db | Upcast CTD FL at 20db | CHL at 10db | Upcast CTD FL at 10db | CHL at 2db | Upcast CTD FL at 2db |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1X | 15 | 0.3 | 25.2 | 1.6 | 28 | 4.0 | 39 |
| 3X | 5 | 0.3 | 8.4 | 1.6 | 9.4 | 4.0 | 13 |
| 10X | 1.5 | 0.3 | 2.5 | 1.6 | 2.8 | 4.0 | 3.9 |
| 30X | 0.5 | 0.3 | 0.9 | 1.6 | 0.9 | 4.0 | 1.3 |

The 1X and 3X cables produce very high dark values and they produce fluorescence that is much higher than extracted CHL for both low and high CHL, which is not what we expect from this type of sensor. Normally at CHL~0.3ug/L we get a ratio of fluorescence to CHL that is between 3 and 20 and for CHL~1.6ug/L we expect a ratio >1 and at 4ug/L the ratio is usually close to 1.

The 30X cable produces a dark value that looks better than the 10X though still rather high. At 20db the fluorescence is believable from either the 10X or 30X cable, with the 30X perhaps a little low. At 10db and 2db the 30X results definitely look too low. The 10X looks much more like results we have seen in the past from this type of sensor, but the evidence is not conclusive given the high variability in such data. We can at least say that there is no clear evidence that the gain was entered wrong in the configuration file. Moreover, the 30X cable would be an odd choice to make in inlets in June. Fluorescence must be treated with caution – the values do not look reliable but the shape may be useful.

The descent rate is sometimes noisy but usually quite steady.

The pairs of temperature and conductivity channels are most often in reasonable correspondence during downcasts, on average, but there were many excursions of one relative to the other. The upcasts are noisy, as usual. Some casts had quite large differences between sensors and it is not immediately obvious which pair are producing better results.

Dissolved oxygen looks noisier than usual. Similar problems were noted during 2013-17 and were associated with problems in secondary channels. PAR and transmissivity look normal.

The altimetry signal is often very noisy at the bottom, but there is likely enough good data to enable an estimate.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Because there were many stops for bottles during all casts and because noisy temperature may be leading to noise in the dissolved oxygen, the usual tests were inconclusive in establishing the best choice for advancing the DO channel. For 2013-17 a setting of +4.5s was used, and that value produces reasonable results for 2013-50.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

The upcast data are very noisy making the usual tests for CELLTM settings hard to interpret and there were no casts without stops for bottles. During 2013-17 and 2013-38 the best choice for both sensors was found to be (α = 0.0245, β=9.5).

CELLTM was run using (α = 0.0245, β=9.5) for both conductivity channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

There were no deep casts for this cruise. But given some concerns about a repair done between events #23 and 24, the deepest casts before and after the repair were examined. Included in the table below are values for the two previous cruises 2013-38 and 2013-17. The values highlighted in grey were from a group of casts that stood out as odd during 2013-17.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2013-38-0111 | 1000 | -0.0006 | 0.0002 | 0.0015 | High, moderate |
|  | 1900 | -0.0005 | 0.00001 | 0.0015 | “ |
| 2013-17-0017 | 1000 | 0.005 | 0.00021 | -0.0018 | High, F.Noisy |
|  | 1800 | 0.0017 | 0.00016 | 0.0004 | “ |
| 2013-17-0036 | 1000 | 0.0001 | 0.00013 | 0.0014 | High, XNoisy |
| “ | 1800 | -0.0001 | 0.00013 | 0.0018 | “ |
| “ | 2800 | -0.0005 | 0.00014 | 0.0022 | “ |
| 2013-17-0065 | 360 | ~0 VN | 0.00006 | 0.0007 | High, VNoisy |
| “ | 650 | -0.0003 | 0.00011 | 0.0011 | “ |
| “ | 1000 | -0.0001 | 0.00012 | 0.0015 | “ |
| “ | 1800 | -0.0002 | 0.00015 | 0.0018 | “ |
| “ | 2800 | -0.0004 | 0.00015 | 0.0023 | High, XNoisy |
| 2013-17-0086 | 360 | -0.0008 | 0.00002 | 0.0011 | High, FNoisy |
|  | 650 | -0.0003 | 0.00011 | 0.0016 | “ |
| “ | 1000 | 0.0003 | 0.00016 | 0.0015 | “ |
| “ | 1800 | 0.0001 | 0.00015 | 0.0019 | “ |
| “ | 2800 | -0.0003 | 0.00016 | 0.0023 | “ |
| “ | 2800 | -0.0007 | 0.00015 | 0.0027 | “ |
| 2013-50-0006 | 340 | -0.0009 | +0.00001 | 0.0010 | Mod, F Steady |
| 2013-50-0019 | 360 | -0.0005 | 0 | 0.0005 | High, Steady |
| 2013-50-0021 | 360 | -0.0020 | -0.00013 | 0.0003 | High, FNoisy |
| 2013-50-0027 | 360 | -0.0013 | +0.00006 | 0.0006 | High,Steady |
|  | 460 | -0.0012 | +0.00012 | 0.0008 | “ |
| 2013-50-0030 | 360 | -0.0014 | -0.00007 | 0.0007 | High, Steady |
| “ | 650 | -0.0005 | +0.00005 | 0.0014 | “ |

The differences during 2013-50 look slightly different before and after the repair before cast #24, though the variability ~360m makes that difficult to judge. The differences at 650m are reasonably close to differences at that level for 2013-17-0086.

##### Initial Assessment of T-S data

When this equipment was used during 2013-38, problems were found in both salinity channels with unstable excursions while temperature appeared to be unaffected. There was also shed wake corruption during that cruise, but that showed up in both temperature and salinity and correspond to low descent rates. The problem appeared to be related to rapid deceleration of the CTD somehow affecting alignment. While there were problems in the primary, those in the secondary were much worse. This was not found to be a significant problem during 2013-17, which was the next cruise using this equipment. One possible explanation is how the ship is aligned in rough seas. With the LARS system the ship’s crew are less likely to point into the waves and so large decelerations may result.

During 2013-17 both salinity channels were found to be very noisy but the most unusual feature was that the difference between the two temperature channels was unusually noisy. This was even seen in a cast with a very steady descent rate. There was a group of casts that were much worse than others.  Dissolved oxygen data were considered of lower quality than usual and the casts with the large differences between temperature channels also produced dissolved oxygen that was out of line in the comparison with titrated calibration samples. So it was thought that the temperature was the main problem with the dissolved oxygen data, possibly due to a problem with the secondary plumbing.

Mid-way through this cruise a leaky TC duct was replaced. An investigation was made to see if that accounted for some of the problems noted in the previous paragraph. Comparisons were made of the pairs of traces before and after the repair, and no obvious change was seen. It was also noted that the plumbing arrangement differed for the two pumps which may be a more important issue.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

For cast #6 an initial drop to 10db was included in the acquisition file. A text editor was used to remove those records so that DELETE will select the best data for the downcast.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  No obvious problems were found but there appear to be some spikes, most likely all due to temperature spikes.

The cross-reference check was compared with the log book and the only error found was 1 wrong station name which was corrected in the headers of profile and bottle files. Times are confused in the log with PDT early in the cruise and later UTC. The notes about this are unclear. There are clear entries in the rosette log sheets and the NMEA headers are correct.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made to check that the readings were reasonable. There were many spikes right at the bottom and frequently there was shoaling during the stop at the bottom, so while the readings look reasonable, there is no clear “right” reading. There was no altimeter header entry for the bottle file from event #24 so one was added since it is fairly clear from the plot. There were a few errors in water depth which were corrected in the profile and bottle files based on log entries.

During this cruise pressure readings were taken on deck before and after the cast. A few of the readings were difficult to read, so those were excluded.

| Event # | Reading before Cast (db) | Reading after Cast (db) | Difference (db) |
| --- | --- | --- | --- |
| 6 | 0.3 | 0.11 | 0.19 |
| 15 | 0.14 | 0.13 | 0.01 |
| 19 | 0.78 | 0.14 | 0.64 |
| 23 | 0.11 | -0.16 | 0.27 |
| 24 | 0.3 | 0.12 | 0.08 |
| 27 | 0.35 | 0.18 | 0.17 |
| 28 | 0.36 | 0.1 | 0.26 |
| 30 | 0.24 | 0.16 | 0.1 |
| 31 | 0.23 | 0.16 | 0.07 |
| 32 | 0.19 | 0.13 | 0.06 |
| 33 | 0.25 | 0.23 | 0.02 |
| 34 | 0.24 | 0.12 | 0.12 |
| 35 | 0.25 | 0.13 | 0.12 |
| 36 | 0.25 | 0.1 | 0.15 |
| 38 | 0.29 | 0.14 | 0.15 |
| 39 | 0.2 | 0.09 | 0.11 |
| 41 | 0.23 | 0.13 | 0.1 |
| median | 0.25 | 0.13 | 0.12 |
| stdev | 0.14 | 0.08 | 0.14 |

The readings at the end of casts are more consistent. The differences become smaller and vary less as the cruise goes on. These values would suggest that the pressure is reading very slightly high, but well within the 1db initial accuracy of the sensor.

The surface report gave an average surface reading of 0.5db. This is very shallow for the Tully, but the examination of a few casts shows that the pumps were not turned on until the CTD was at about 2db. This suggests the pumps were turned on just as the rosette was fully underwater. There are very low initial conductivity readings. The first conductivity readings are very close to 0.  For two upcasts the pumps were on until the sensor was very close to the surface and the conductivity shoots to 0 at ~+0.2db for one and at ~0db for the other. Recalibration of pressure is not justified.

##### Shift

Fluorescence

The tests for fluorescence alignment for this cruise are very difficult because all casts have stops for bottles on the upcasts. However, the usual setting of +24 records does do a good job of making the offset between upcast and downcast more similar to that for temperature.

SHIFT was run on all casts to advance the SeaPoint fluorescence by +24 records.

Conductivity

Tests were run on 4 casts to determine how to align the conductivity channels relative to temperature, so as to minimize noise in the salinity channels.

A shift to the primary conductivity of -0.5 records looked best overall and for the secondary values a shift of +1.4 records was best. The differences between settings tested were slight but all did improve the data.

SHIFT was run on all casts using -0.5 records for the primary conductivity and +1.4 records for the secondary conductivity.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel; further alignment does not appear necessary, though it is hard to judge with so many stops during the upcast so that the response is quite different from the downcast.

##### DELETE

DELETE was run on all files.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings:

The DEL files were copied to EDT files in case some do not need editing.

##### DETAILED EDITING

At this stage it must be decided whether primary or secondary channels will be edited, and hence archived. The bottle comparison shows that both salinity channels are close to the bottles. The secondary salinity looked slightly noisier and the primary channels were selected for most casts on the previous cruise, so the primary channels were selected for this cruise.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems (chiefly miss-alignment of T and C) and likely to affect the bin-averaged values and many records corrupted by shed wakes including some surface records.

Editing was required for all casts.

Most casts from Effingham Inlet have some bad pad values in the temperature channel, with 99.0000 for bad values rather than -99.0000. CTDEDIT would not work with these files. A text editor was used to fix the bad pad values and the channel limits were adjusted to reasonable values. Then CTDEDIT worked properly. This problem might be related to the conversion problems described in section 5.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The temperature and conductivity sensors were used for 2 cruises since the last factory calibration. For 2013-38, with analysis done within 13-24 days of collection, the primary was found to be low by 0.0005 and the secondary high by 0.0010. For 2013-17 with analysis within 22-40 days, the primary was found to be low by 0.0017 and the secondary low by 0.0005. The 2013-17 results may be influenced by delays in analysis leading to evaporation of samples.

2. Dissolved Oxygen

The DO sensor was used during 2013-38 and 2013-17. For the first the fit against bottles was found to have slope 1.0471 and offset +0.0172.  For 2013-17 the offset was forced to be 0 and the slope was found to be 1.054. Those are close indicating only slight drift.

3. Pressure

The sensor on CTD #443 has been used on 2 other cruises since factory calibration and the factory offset was found to be appropriate for both.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. There is no local climatology for Douglas Channel and Effingham Inlet, but there is a “Marsden” climatology that covers 1° X 1° squares. Many of the deep salinity values in Douglas Channel are below the minimum and the deep temperatures are occasionally high. This is not surprising as there are few data from the channel area and the values are expected to differ markedly from offshore areas. The offshore casts fell within the climatology as did the Effingham Inlet casts with the exception of one case of low salinity at depth.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### Initial Recalibration

No recalibration was considered necessary for pressure and salinity.

Dissolved oxygen was using the results of the comparison described in section 4:

DO_Corrected = DOX_CTD *1.0545

CALIBRATE was run using file 2013-50-recal1.ccf to apply that correction to the dissolved oxygen channel in the SAM and MRGCLN2 files. COMPARE was rerun to check that the salinity was recalibrated appropriately and it was; the average difference in the fit when the same points were used as in the first fit, was +0.0002mL/L. The recalibration was then applied to the EDT files.

(See file 2013-50-dox-comp2.xls.)

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to correct for errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The fit looks excellent below 100db. Above that there are more differences >0 than <0 which is likely due to poor flushing of Niskin bottles as discussed earlier. The SBE Dissolved Oxygen is likely within 0.25mL/L even in the top 100db.

Based on the final comparison a very rough estimate was made of the accuracy of the downcast Oxygen:Dissolved:SBE data:

±0.25mL/L from 0 to 100db

±0.1mL/L from 100 to400db

±0.05mL/L below 400db

(See 2013-50-dox-comp3.xls.)

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

A median filter, fixed size=11, was applied to the fluorescence channels in the COR1 files to reduce spikiness. Before and after plots of a few casts showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number,  Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and

unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.25 mL/L from 0 to 100db

±0.10 mL/L from 100 to 400db

±0.05 mL/L below 400db

The two CTD salinity channels were generally within 0.001 of each other.

The comparison with bottles had a lot of scatter, but the average

differences showed that both CTD salinity channels were well within

0.001 of bottle values. During a May 2013 cruise with more deep sampling,

both salinity channels were found to be within 0.001 of bottle samples.

SBE Fluorescence data from this sensor are considered unreliable. Dark values

are much higher than expected from this cruise and other cruises in May

and June 2013 and appear to be increasing with time. Fluorescence values have

also been lower than usual relative to bottles, particularly at higher

chlorophyll values. The data are provided since the shape may be useful.

For details on the processing see the report: 2013-50-proc.doc.

The cross-reference list was produced no errors were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values for the offshore casts were between 100% and 120%. For the Douglas Channel area saturation ranged from 105% to 120%. In Effingham Inlet saturation was higher, ranging from 160% at the head of the inlet to ~130% at the mouth. The offshore values look reasonable for the season and calm conditions.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number to add “Mid-ship” to the instrument location section and to add a comment about quality flags and analysis methods and a few notes about the CTD data. Note that the formats for the titrated dissolved oxygen data were changed from the usual F8.3 & F6.1 to F7.2 & F5.0 based on the analyst’s recommendation.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no errors were found, though there are many samples not indicated on the rosette sheets, as noted earlier.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were.

The track plot was produced on screen and no errors were found.

Header check was run and no problems were found.

A cross-reference listing was produced for the CHE files and no errors were found.

##### Sensor History

The sensor history was updated.

# Particulars – comments from logs

General

- Many of the comments were difficult to read and it was not always clear which event they referred to, so there may be errors in the entries below.

- There were many comments on problems with the Niskin bottles. These are reported in the analysis spreadsheets and comments on bottle samples in CHE files.

- The times are sometimes PDT, sometimes UTC. Either is fine if it is clear, but in many cases it was not. Some entries were changed later but  don’t indicate either PDT or UTC.

6. Initial drop to ~20m with pumps off; brought back to surface and dropped again with pumps on.

18. Error in event # and station name in original file – corrected in some files. Check converted files.

19. Apparent misfire - Niskin #7, Sample 45.

23. Tubing on secondary TC duct was replaced after this cast due to leak.

23. Niskin #1 was not fired – sample numbers were adjusted.

24. Niskin 7 looks like mis-trip.

25. Station name DC04 and DC02 both entered in log – looks like DC02 is correct based on position.

27. Bottle 19 leaky spigot, bottle 7 misfire, bottle 6 sticky spigot, bottle 3 top seal leaks

28. Water mixing to bottom ∆T fluctuates almost to 120m.

31. Problems entering depth in SeaSave – entered in header files after acquisition. Converted files ok.

31. Altimeter read 4.6 at the bottom but there were spikes in Trans, Temp and Cond.

35. Niskin #2 tripped by accident – no samples taken and sample numbers offset for other bottles.

39. May have almost touched bottom then moved up to 2.8m above bottom. (some spikes in ctd data but no evidence of touch-down.)

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 2280 | 29Jan2013 | Factory |  |  |
| Secondary Temp. | 2668 | 1Feb2013 | Factory |  |  |
| Secondary Cond. | 2754 | 30Jan2013 | Factory |  |  |
| Transmissometer | 1396DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 16Feb2013 | Factory |  |  |
| SeaBird Fluorometer | 2228 |  |  |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Pressure Sensor | 63507 | 15Apr2011 | Factory |  |  |
| Altimeter | 1204 |  |  |  |  |