## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2013-55

Agency: Ocean Sciences Division

Location: Strait of Georgia

Project: Aquaculture Environmental Studies

Party Chief: Sutherland T.

Platform: Vector

Date: July 2, 2013 –July 9, 2013

Processed by: Germaine Gatien

Date of Processing:  October 2013 –May 2, 2014

Number of original HEX files excluding tests: 82

Number of CTD files: 82		Number of bottle casts:	32

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), an SBE 43 DO sensor (#0047), a SeaPoint Fluorometer (#2356) with a 10X cable, a Biospherical QSP-400 PAR sensor (#4601), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (no serial # available).

The CTD deck unit was an SBE 11+, serial number 0425. The logging computer was #2.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and Dissolved Oxygen analysis logs were in good order except that the dissolved oxygen sensor was incorrectly identified. Since the configuration file had the same error, a lot of processing had already been done when the error was discovered.

It is essential that visual checks are made when entering sensor serial numbers in the log book. This is most critical for the externally-mounted sensors. As a backup it is recommended that dated photographs be taken and saved every time sensors are changed on a CTD.

Salinity was not recalibrated, but there may be a significant error. There was a post-cruise calibration that indicated that the primary salinity was high by ~0.0036 and the secondary low by ~0.0032, but there was a cruise between this one and the return to the factory. We know that there was a shift in at least one conductivity sensor (temperature differences did not change significantly) between this cruise and the one that followed based on looking at differences in the salinity channels. But there is no way to establish which pair changed, and it is even possible that both shifted in opposite directions.  If the primary did not shift, then for this cruise the secondary is low by about 0.003. If the secondary did not shift, then for this cruise it is high by ~0.003. It seems odd that both sensors were either significantly high or low before the shift given what looks like quite steady performance over the years. There is likely a problem with the comparison with bottles for this cruise due to incomplete flushing of Niskin bottles. Many of the comparisons from other cruises using these sensors were also weak due to either delayed processing or flushing problems, but taken as a group there is no sign of a problem. Since there was no way to be sure which sensor shifted or if both shifted at the same time in opposite directions, no recalibration was applied to salinity.  Salinity accuracy should be considered ±0.004.

The differences between the SBE dissolved oxygen sensor values and titrated samples were noisy but the correction is close to one used for the previous cruise.

The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real. The pH channel was removed from event #100 because the vial was left on during deployment and the data look bad.

PAR, PAR:Reference, Fluorescence, pH and Transmissivity data are nominal and unedited except that some records were removed in the course of editing the temperature and salinity. The PAR:Reference values look much too high at times but at other times they correspond well to PAR values at the surface.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5  mL/L from 0 to 15db

±0.15  mL/L from 15 to 75db

±0.05 mL/L below 75db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen and salinity.

There were many notes in the log – those of special concern are:

100 – pH data should be removed because the vial was left on the sensor.

110 – header error

115 – may be a problem in DO data 27-33 m – check for flow problem.

132 – Niskin vents open so rerun as 133. The CTD data should be fine.

Nutrients, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The calibration constants were checked for all instruments and no errors were found, though the date of one calibration was incomplete, so that was added. There were no changes through the cruise. The same equipment was listed for 2013-12 and 2013-13.

These data were partly processed when it was discovered that the DO sensor was likely entered incorrectly in both the Daily Science Log book and the configuration file used at sea. This discovery came as a result of finding the same DO sensor (1438) listed for Ricker cruise 2013-43 that overlapped with 2013-13 and cruise 2013-55. When the comparison of 2013-13 titrated dissolved oxygen samples with the DO:SBE data looked unusual, it became obvious that the error must be with the configuration file for CTD #0585 used on the Vector rather than CTD #0506 on the Ricker. Tests were run on 2013-13 seeing if DO S/N 1438 should have been 1483, but that was no better. Finally, the CTD tech found a note indicating that #0047 was in use on the Vector. Doing the 2013-13 bottle comparison with that assumption led to normal results. So it will be assumed that sensor #0047 was used during 2013-13, -55 and -14. Comparisons with bottles will be run for each to ensure they look ok.

Photographs of the CTD equipment would have been very helpful and that is recommended in future.

The corrected configuration file was saved as 2013-55-ctd.xmlcon.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

##### Conversion of Full Files from Raw Data

All files were converted using 2013-55-ctd.xmlcon.

The header error in file #110 was fixed. There was no rosette sampling for that cast.

A few casts were examined and all expected channels are present. The descent rate was mostly very quiet. The primary salinity looked bad during part of the upcast of event #136. The data looks ok during the bottle stops; there was no salinity sampling during the poor section. The other variables look ok during this section. The problem was not seen in any of the other plots checked.

The two temperature and conductivity channels are fairly close during the downcasts and, as usual, farther apart during the upcasts. The descent rate is generally steady but sometimes low.

Altimetry is sometimes noisy at the bottom. Fluorescence, transmissivity, pH, PAR, SPAR and transmissivity look normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2013-55-ctd.xmlcon.

The files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files to check for outliers. CTDEDIT was used to examine and edit a few salinity points in casts #20, 73 and 95. The output files were copied to *.BOT.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records. There were a few adjustments needed where a bottle was not sampled.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine.

The SAM files were bin-averaged on sample number.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-55-bot-hdr.txt.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-55oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2013-55oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2013-55SAL.xls. The analysis was done within 9 to 16 days of collection. The files were simplified and saved as 2013-55sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-55nuts.xls. This includes a precision study.

The file was simplified, reordered on sample numbers and the file was saved as 2013-55-nuts.csv.

The file was converted to individual NUT files.

Note that the nutrient spreadsheet includes bucket samples which do not correspond to CTD bottle firings.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

T

here were many apparent bottles fired during event #86, but only the first 10 were intended and only the first 9 were actually sampled. So all firings except bottles 1-9 were removed from the MRG file.

The MRG files were put through CLEAN to remove SeaBird Headers and comments from the secondary file.

An initial run of Header Check was examined. It was found that the intended salinity sample from cast #107 was not present but there is a note in the log that it was intended but forgotten.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts to see what setting brought the vertical offset between the downcast and upcast DO traces into the best agreement with the temperature traces. When this sensor was used during 2013-13 tests found that an advance of +2s to +3s worked best and for 2013-14 results varied from +3.5s to +4s were best, on average.

For this cruise the best choice is especially hard to judge because many casts are shallow and there were frequent stops for sampling which complicates interpretation, but a value of +4.8s looked reasonable. Since this is higher than found on cruise before and after, further tests were done, but a low setting produced poor results. The DO gradient was quite low, and descent rate quiet, so this may be because a better comparison was possible with these data. Whatever the reason, using a value of +3s looks poor.

ALIGNCTD was used to advance the DO Voltage by 4.8s relative to the pressure.

##### CELLTM

As for 2013-13 most settings tested improve the data. For 2013-13 a setting of (α = 0.0245, β=9.5) looked reasonable overall for both sensors and that setting works well on these data.

CELLTM was run using (α = 0.0245, β=9.5) for the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The usual examination of differences between temperature, conductivity and salinity channels was not run because none of the casts are deep enough to get reliable results.

##### Compare

Salinity

Compare was run with pressure as reference channel.

Plots of differences versus pressure show a lot of scatter. This is expected in such shallow sampling. Above 50m most samples have higher salinity than the CTD. The analysis was done quickly so evaporation of samples is not likely to be significant. Incomplete flushing of Niskin bottles could account for the differences as the water sampled may well come from a little deeper than the water sampled by the CTD. The very steady descent rates show that conditions were calm which may exaggerate this effect. This becomes less serious below 50m where the salinity gradients are much lower.

When only the 9 samples below 50db are included, the primary salinity is higher than bottles by an average of 0.0006 and the secondary low by 0.0045, but the standard deviations for both are ~0.03. The standard deviations in the salinity channels during a stop are often higher than we would usually consider a cut-off for such comparisons. The only bottle below 100db (228db) suggests the primary is high by 0.0006 and the secondary is low by 0.00007. The standard deviations in the salinity are very low for that bottle stop.  And finally if the 3 deepest casts are included the primary appears to be high by 0.0007, but one of the values has a high standard deviation in the CTD data. If that one is excluded the primary is even closer to bottles. The secondary is low by 0.0007 using the 3 bottles. A sample from 66db during event #84 is clearly an outlier; it was taken from the bottom bottle and such samples tend to show up as outliers.

This comparison is too weak to reach a conclusion about the calibration, though the deepest bottles suggest that both CTD channels were within 0.001 of bottles. In all comparisons the primary salinity is higher than the secondary, with a median difference of 0.0014.

The only problem sample is one that was labelled incorrectly. Either the station name or sample number was wrong or both. Sample #206 is found during cast #108 at station Mal55, but the label says Mal34. The rosette log for cast #108 does not indicate that a salinity sample was taken during that cast. The bottle value differs from the CTD salinity at the relevant depth by 0.07. The only missing sample from any other cast is sample #201 from cast #107 and the CTD salinity there is lower by ~1.0. Given no convincing evidence about the source of this sample, the value was replaced with a pad value. The SAL file was removed since that was the only sample for that file. This will not be added to any bottle file.

For full details for the COMPARE run see file 2013-55-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There was a lot of scatter but no major outliers. The samples flagged by the analyst were checked and only 1 looked a little out of line (Sample #66, Event #47) but it was a 5m sample and there is a lot of scatter in near-surface samples when there is often a steep DO gradient.  There is no justification to change any of the flags.

Minor outliers were excluded based on residuals. SeaBird recommend using a fit with offset=0. While that is not always found to produce a good fit, it is the only choice for these data since there are no samples with values <2mL/L. The fit found was

CTD DO Corrected = CTD DO * 1.0672

That correction is a little smaller than found for 2013-13 when the slope was 1.075, and for 2013-14 when it was 1.084. The differences between bottle and CTD DO are theoretically due only to calibration drift, but other factors likely have some effect on the fits. We often find that post-cruise factory calibrations indicate smaller drift than those seen in our bottle comparisons. Poor flushing of Niskin bottles will lead to lower bottle values than those seen by the sensor which would make the correction smaller than they should be, not larger. On the other hand, slow response of the DO sensor may make the sensor values slightly lower than the bottles where flushing is efficient and the CTD is physically lower than the Niskin; both those effects might lead to corrections slightly larger. As DO ranges and gradients vary, the corrections suggested by bottle comparisons will vary too. But what explains significantly higher corrections is a puzzle. There may be other problems that are only seen in motion that are not accounted for. Whatever, the cause, the bottle comparison is all we have to go by.

For this cruise the salinity comparison suggests poor flushing of Niskin bottles above 60m and that would definitely affect the fit, since the CTD DO would be compared to samples from deeper in the water column, especially near the surface. That would reduce the slope of the DO correction. The range and gradients of DO for this cruise are closer to that of 2013-13 than 2013-14. There were also rough seas during 2013-14 which might reduce flushing problems and thus increase the correction slope.

The correction:

CTD DO Corrected = CTD DO * 1.0672

was applied to the MRGCLN2 and SAM files. COMPARE was rerun to ensure the correction was appropriate and it was.

For full details of the comparisons see files 2013-55-dox-comp1.xls and 2013-55-dox-comp2.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined and significant outliers were in high gradient areas. No additional flags appear justified.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  The minimum pressure was 1.2db at the end of a cast.

The PAR channel appears to be off-scale at the beginning of the file, though this data will likely be removed by DELETE. Salinity has some unbelievably high values.

The speeds look fine. The fluorescence did not go off-scale.

The surface check was run and shows an average surface pressure for the cruise was 2.5db which is reasonable for the Vector.

The cross-reference check was compared with the log book and two errors in station names were found and corrected in the headers of casts #114 and 115 (including bottle file for #115).

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for all casts since there were some problems noted at sea. Despite many spikes near the bottom of casts, all the header entries look appropriate (within 2m of the log book) except for cast #131. As noted at sea the trace looks strange for #131, so the altimetry header was removed from the file. This was not a bottle file.

The depths were checked against the log book. There were many small discrepancies which were ignored if <2m. The drifting near-shore accounts for the differences. For the others, the altimetry and maximum pressure sampled were checked to see which was more realistic. Changes were made to match the log entry, as needed, to both the CLN files and the SAM files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Conductivity

During 2013-13 the shift settings chosen for the conductivity channels were -0.5 records for the primary and -0.3 records for the secondary. For 2013-14 the choices were -0.5 records for both.

The casts were not very suitable for the usual tests as there was little noise in the salinity; the shifts showed only small effects, but there did seem to be slight improvement using -0.5 and -0.3 records.

SHIFT was run using -0.5 records for the primary conductivity and -0.3 records for the secondary.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests to establish the best setting for the alignment were inconclusive. The casts are shallow and stops frequent and there are few distinguishing features. Making this judgment even harder is the noisy upcast temperature and high variability in the area. On a few 2012 cruises using this sensor, a setting of +65 records was found best, for 2013-16 a value of +75 records was selected since it generally does a better job near the surface and most casts were very shallow. For 2013-12 +65 looked best.

For this cruise a setting of +70 looks like a reasonable choice, but the evidence is weak.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but the upcasts were so different from downcasts that it was very hard to judge the alignment. No further alignment was applied.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Conductivity:

The conductivity sensors were recalibrated in late March 2011 and early April 2011 and were used for 1 cruise in 2011 and 6 times in 2013 previous to this cruise. There was good calibration sampling during 2013-01 and both salinity channels were within 0.001 of bottles. One other cruise had only bottom bottles. Two cruises had a lot of scatter with evidence of poor flushing near the surface, but showed salinity within 0.001 of bottles when outliers were excluded. The sensors were also used for 2013-14 which followed this cruise. It was noted that the differences between the two conductivity channels changed sign between 2013-55 and 2013-14, and the primary salinity shifted from being lower than the secondary by about 0.001 to being higher by about 0.006. Both CTD salinity channels were lower than bottles, but due to delay in analysis the comparison was not trusted; there was a lot of scatter which suggests evaporation might be an issue. Thus it was thought likely that the primary was much closer to the bottles than it appeared and the secondary salinity might be higher than the bottle salinity.

Dissolved Oxygen

The sensor was used on only one other cruise since its last recalibration in March 2011. During 2013-13 the fit was forced through the origin with a slope of 1.075. A preliminary analysis of cruise 2013-14 which occurred after this cruise shows a slope of 1.084.

Pressure

The sensor was recalibrated in April 2011 and was used for 1 cruise in 2011 and many during 2013. The factory offset was used for all those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature and salinity profiles fell within the local climatology where it was available except for a few spikes that will be removed in editing.

Repeat Casts – There were no repeat casts and there is too much variability for comparison of nearby casts to have any value in assessing repeatability of the sensors.

Post-Cruise Calibration – There was a post-cruise calibration in January 2014 that showed the primary salinity to be too low by about 0.0036 when it arrived back at the factory and the secondary to be too high by 0.0032. This implies a difference between sensors of ~0.0068. Cruise 2013-14 in September 2013 used these sensors. Since there was a large change in differences between the two salinity channels between 2013-55 and 2013-14, the post-cruise calibration cannot be assumed to be relevant to 2013-55.

##### CHANNEL CHOICE STUDY

COMPARE does not help in choosing which sensor pair to select for archiving. For all previous cruises since the last calibration, the primary channels were selected, though the differences were considered slight. During 2013-13 the two salinity channels were close and had a similar level of noise. The primary channels were selected except for 5 casts in which the primary salinity had some bad sections. The two salinity channels were very close in value during stops. But during the cruise that followed this one there were large salinity differences and based on post-cruise calibrations it appears that one channel was reading low and the other high by about the same amount. For this cruise the secondary T-S plots are a little smoother so they were chosen. There may have been some minor problem affecting alignment of the primary conductivity.

##### DETAILED EDITING

The secondary temperature and salinity channels were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the cast.

All files required some editing.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure does not appear to require recalibration.

Salinity recalibration is problematic. We know that something changed between this cruise and 2013-14 so we cannot assume that the post-cruise calibration provides information relevant to this cruise. It is most likely that the shift was due to a change in only one conductivity sensor. If the primary did not shift, then for this cruise the primary is low by about 0.003 and the secondary by ~0.003. If the secondary did not shift, then for this cruise the primary is high by ~0.004 and the secondary is high by ~0.003. It seems odd that both sensors were either significantly high or low before the shift given what looks like quite steady performance over the years. Of course many of the comparisons were weak, but taken as a group there is no sign of a problem. Since there is no way to be sure which sensor shifted or if both shifted at the same time in opposite directions, no recalibration will be applied to salinity.

SBE Dissolved Oxygen was recalibrated by applying the following equation:

CTD DO Corrected = CTD DO Raw* 1.0672

CALIBRATE was run using file 2013-55-recal1.ccf to apply the correction to the EDT files.

COMPARE was rerun on the recalibrated data and the results were good.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. When the differences were plotted against pressure the CTD looked low by for the 5m bottles and a little high for the 10m bottles. The fit of differences against CTD DO was very noisy, especially for higher DO values. No further calibration looks likely to improve the fits. See 2013-55-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. There were a few small unstable features in the T-S plots but some are expected in this region. There is no evidence that these are due to instrumental problems.

Profile plots were examined at this point.

It was noted that the PAR:Reference values look much too high at times, but at other times they correspond well to PAR at the surface.

When all casts were plotted together, the dissolved oxygen profile for cast #130 stood out as odd. The lowest values found during the whole cruise were seen at mid-depth. Bottle values confirm that these are bad CTD oxygen data. The primary temperature also looks a little odd, so the problem was likely due to a blockage in the primary pump system. There was a large fluorescence spike at the bottom of the cast and after that the data look fine. Unfortunately, the upcast trace is too noisy to be a good substitute, so the DO should be removed from the downcast file. This problem will not affect the CHE file.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE,  Altimeter, Status:Pump, Descent_Rate and Flag

For cast #100 the pH:SBE channel was also removed because the vial was left on the sensor and the data look bad.

For cast #130 the Oxygen:Dissolved:SBE channel was also removed.

Cast #115 was checked to see if there was a problem with the dissolved oxygen data but it looks very similar to other casts in that area. In the log it was noted that there was a “bump” in DO near the bottom, but a slight rise in values near the bottom is typical of other nearby casts. However, the upcast DO looks very bad, so that might be what was noticed.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH:SBE, PAR and PAR:Reference data are

nominal and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5  mL/L from 0 to 15db

±0.15  mL/L from 15 to 75db

±0.05 mL/L below 75db

Salinity has not been recalibrated, but there may be a significant error.

Incomplete flushing of Niskin bottles makes interpretation of the bottle

comparison uncertain. If only bottles below 50db are included both

salinity channels appeared to be close to bottles. There was a large

change in at least one of the conductivity sensor calibrations between

this cruise and the one that followed, so the post-cruise calibration

cannot be applied to these data.

Salinity is considered ±0.004.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2013-55-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. There was a lot of variability with values between 90% and 160% in Malaspina Inlet and between 60% and 90% in Gorge Harbour. The high values are in areas with high near-surface temperature gradients and likely little mixing, while the low values are in areas where gradients are low and mixing likely significant. These variations are not unusual given the season and location. These values do not suggest a problem with dissolved oxygen processing.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #100, the  pH:SBE channel was also removed because the vial had been left on. The data look bad.

DERIVE was run to add Oxygen:Dissolved:DO and Oxygen:Dissolved channels with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

Standards check was run on all files and no errors were found.

A header check was run on the final files and a few problems were found. This step was repeated after corrections until all problems were resolved.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

1. Cast not archived.

16. 30m off station because a log boom was in the way.

19. Altimeter not clear on depth.30. Two bottles fired at 40m48. Depth kept increasing –s topped about 4m off bottom.

70. Started 1m lower than normal at surface.

86. Problem with Seasave – all bottles were fired after the 5m bottle.

89. Seasave ok.

89-107. Time-series – every 30 minutes at Station Mal34.

100. Vial was left on pH meter.

105. Latitude field from NMEA data was blank. Restarted Seasave and all was well.

107. Intended salinity sample not taken.

109. Altimeter unclear.

110. Error in station name – should be GH16, not GH18.

115. Strange bump in DO trace from 27-23m upcast. Blockage in flow.

131. Altimeter indicated 16m off bottom at 75m which was the depth of this station in 2012. Altimeter cleaned after cast since trace looked odd.

132. All Niskin vents open, so cast rerun as 133. No bottle file prepared as no sampling.

133. Returned to 5m at end because forgot to fire 5m bottle.

136. Tide line went by ship as rosette was @40-0m.

139. Stopped at 100m for a few minutes due to discrepancies in sounder depths between bridge and lab. Bridge sounder believed inaccurate.

141. Some sediment washed into water from sieve just prior to the CTD being put into water.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 5Apr2011 | Factory |  |  |
| Conductivity | 3321 | 30Mar2011 | Factory |  |  |
| Secondary Temp. | 4700 | 1Apr2011 | Factory |  |  |
| Secondary Cond. | 1766 | 29Mar2011 | Factory |  |  |
| Transmissometer | 1185DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 0047 | 1Apr2011 | Factory |  |  |
| PAR | 4601 | Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | Dec2010 |  |  |  |
| SBE Fluorometer | 2356 | n/a |  |  |  |
| Pressure Sensor | 77511 | 22Apr2011 | Factory |  |  |
| Altimeter | 16504 | 16Mar2011 |  |  |  |