## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 11 Feb. 2025 | Correction to turbidity units. G.G. |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2013-57

Agency: Oceans Network Canada, Univ. of Victoria, UBC, Fisheries and Oceans: Ocean Sciences Division

Location: Strait of Georgia, WCVI

Project: Dynamics of Hypoxia

Party Chief: Dewey R.

Platform: R/V Falkor

Date: August 19, 2013 – August 31, 2013

Processed by: Germaine Gatien

Date of Processing: 10 February 2014 – 14 April 2014

Number of original HEX files:  43	Number of casts processed: 39 (4 test files not processed)

Number of bottle casts:	32		Number of bottle casts processed: 31 (1 test file not processed)

# INSTRUMENT SUMMARY

The CTD used for this cruise was provided by the R/V Falkor; full documentation was available including pre-cruise calibration parameters. For further information about the ship see the web site: http://www.schmidtocean.org/

A Seabird 911+ (#1087) was mounted below a rosette that contained 24 2L Ocean Test Equipment bottles. Attached to the CTD were a Wetlabs CSTAR transmissometer (#1476DR), an SBE 43 DO sensor (#2331), a WetLabs Fluorometer (FLNTURTD-2578), and an altimeter (#59543). The calibration bottles were 1m above the CTD. The DO sensor was pumped and the fluorometer was not pumped.

An SBE 45 Thermosalinograph was mounted with a fluorometer (FLS 2702) and operated with a remote temperature probe (Seabird SBE 38).

The deck unit was a Seabird model 11, serial number 0920.

The SeaSave version was V7.21f.

The oxygen kit was an Aanderaa Optode 3835 S/N 224.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The file names and event numbers do not match those in the log book. There was a test cast that was saved as event #1 but was called event #0 in the log book. There are frequent notes in the log mentioning that the file names are out of sync with the log, but this was not noted on the rosette sheets so that the analysts were unaware of the problem. The computer file names were used and analysis results adjusted to match. One other inconsistency was that the casts at stations RAD08 and MB16 were recorded in the log in the wrong order. The CTD file names were correct. Notes were made on the rosette sheets to indicate the associated CTD file name.

The CTD log had an equipment list, but no list of personnel other than the Chief Scientist. Notes in the log mostly pertain to a description of the profiles, with no notes about any problems encountered with the equipment other than the issue with event numbers.

The sampling logs and dissolved oxygen analysis logs were generally in good order.

There were many spikes in the CTD data. The obvious ones were removed by program WILDEDIT or by using a graphical editor. There are likely smaller spikes that were not detected, so the quality of the data is likely lower than usual for this type of CTD.

Salinity samples were analyzed within 77-91days of collection which is considered too long for good results. The comparison between CTD salinity and bottles showed both salinity channels to be lower than the bottles by 0.007 and 0.008. Delayed analysis leads to high bottle salinity values due to evaporation of samples, so the CTD salinity is likely more accurate than the comparison suggests. No recalibration was applied to CTD salinity.

The differences between the two temperature sensors are larger than usual, ~0.0015C° below 1000db, but the conductivity differences must have been compensating since the salinity differences are small, ~0.0013. Without a reliable salinity comparison it is impossible to determine which temperature sensor was providing better data. In deep water with low temperature gradients there is no obvious difference in the shape of the traces to suggest a problem with either sensor.

The transmissometer parameters in the configuration file were wrong and produced values that were >100% per metre in deep water. Using the calibration sheet from the factory the slope and offset were recalculated which produced reasonable values. The sensor was last calibrated in 2011; recalibration is long overdue and is relatively easy to do.

There were many problems in the titration of dissolved oxygen samples including particles in the samples and high temperatures in the lab which likely accounted for the many bubbles reported. The analyst recommended that only 2 decimal places be reported for the titrated DO based on the precision study. Dissolved oxygen was recalibrated based on the comparison with bottles.

The Oxygen:Dissolved:SBE data in the downcast files are considered, roughly, to be:

±1 mL/L from 0 to 30db

±0.6 mL/L from 30 to 100db

±0.2 mL/L from 100 to 400db

±0.04 mL/L below 400db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The file names are non-standard for IOS format. The names include what looks like event numbers but they do not match those in the log book. There was a test cast that was saved as event #1 but was not given an event number in the log book. There are frequent notes in the log mentioning that the file names are out of sync with the log. The chief scientist advised that the file names should not be changed to match the log.

The Falkor’s Seabird 911+ was mounted below a 24-bottle rosette. The CTD was lowered to 10 metres and allowed to soak there. One Marine Technician turned on the pumps at this depth. The CTD was then returned to the surface. The other Marine Technician turned on the pumps at this depth.

The downcast was started at 0.33 m/s but usually increased to 1 m/s at depth. Bottles were fired on the upcast.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as a report from Susan Allan who had done partial processing of the data. Calibration sheets were available for the sensors.

Nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts.

The cruise summary sheet was completed.

There is no history available for the sensors as they do not belong to OSD.

Next the calibration constants in the configuration file were checked for all instruments. The parameters were all correct except for a minor error in the date of the DO sensor calibration. There were 2 sensors listed in the configuration file for which calibration information was not available. These are for a 2nd oxygen sensor and of interest to particular researchers. They will be converted but will not be prepared for the IOS Data Library as we have no information on how to process these data.

A new folder was set up in the raw section for CTD data with standard name formats. The file names were changed to standard format using Bulk Rename Utility. The file names differ from the event numbers in the sampling log and CTD log, but the computer file names are to be used. In some cases there is a note in the Daily Science log indicating the CTD file name as well as the log entry which differ by 1. There are 2 sets of BL and HDR files for event #34 so the pair with the right time will be used.

A return was made to the beginning of processing when it was realized that the transmissivity data are bad. The shape is ok, but values >100% / 25cm were found regularly. Values of ~93% are considered the highest possible in clear water and we see values of ~90%/m in at 4000db at Station Papa and at 2000m just beyond the shelf break during La Perouse cruises. The slope and offset entered in the configuration file do not look right given the factory calibration provided, but it was possible that there was a later calibration. The chief scientist checked and could find no indication that had been done. Given that the offset was 0, perhaps the slope used was derived empirically on an assumption that 100% is correct at depth.

Tests were done to see if using slope 19.1833 and offset -0.0575 produced reasonable data. The results look reasonable with a maximum value of ~90%. So conversion will be done with these parameters. The new configuration file was saved as 2013-57-ctd.xmlcon while the original was saved as 2013-57-ctd-orig.xmlcon.

##### BOTTLE FILE PREPARATION

Hysteresis test results were not available for the DO sensor and there is only 1 cast deep enough for this to be a concern. There is likely not enough calibration data to do the tests.

The ROS files were converted using file 2013-57-ctd.xmlcon using the default value of E (0.036).

The Userpoly0 and Userpoly1 channels were included.

File 2013-57-0001 comes from a test cast, but was converted since samples were taken from 1 bottle. So it might prove useful. But there is no sample number and it is not stated which Niskin bottle was sampled when 24 were fired. No CHE file will be prepared.

Those files were put through CLEAN to add event numbers (*.BOT).

Header Check was run on the BOT files and no problems were found.

Station names are missing from events #2 to 7. Those were added based on log records.

Water depths were missing but were available from the log book for all but 2 casts, so those were added to the headers of the BOT files.

Temperature and salinity were plotted for all BOT files. Many outliers were found, some of which are due to closing bottles on the fly in rough weather. There were so many spikes that it was decided to open every file in CTDEDIT. Editing was applied to remove spikes from the following files:

Cast #11 – Both salinity & temperature channels, pressure, DO

Cast #16 – Primary Salinity

Cast #20 – Primary Temperature and Primary Salinity

Cast #25 – Pressure

Cast #27 – Pressure, Primary and Secondary Salinity

Cast #28 – Pressure, Primary and Secondary Salinity, DO and Secondary Temperature

Cast #29 – Primary Salinity

Cast #35 – Pressure and Primary Salinity

The BOT files were then averaged to enable an ADDSAMP file to be prepared. Sample numbers were added to the ADDSAMP file based on rosette log records. A few problems were noted in this process:

- Event #11 has only 23 bottles on the list and one is said not to have fired. But 24 bottles were fired. Perhaps Niskin #22 did not close, but it was fired. Niskin #24 was fired at the surface, but apparently was not sampled.

- Stations RAD08 and MB16 are not entered in the logs in time order and have the wrong event numbers, but the CTD file names are assigned correctly. So the event number for MB16 was originally #23 but should be #25 and RAD08 was originally #24 and should stay as #24.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was sorted on Event_Number and Sample_Number and then converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-57-bot-hdr.txt; it may need further editing to reflect problems found during processing. The * in the spreadsheet file names that follow stands for the date the file was created. These names are updated as corrections are made.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-57oxy*.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2013-57oxy.csv. The event numbers were corrected. That file was converted into individual *.OXY files. The comment for event #11 was changed to indicate that a bottle had not closed instead of the original comment that it had not fired. There were many problems with the analysis including many bubbles and particles. The temperature in the ship lab was too high. The post-cruise analyst recommended that this data be trusted only to the second decimal place.

SALINITY

Salinity analysis was provided in spreadsheet QF2013-57SAL*.xls. The file was simplified and saved as 2013-57sal.csv. The event numbers were corrected. The salinity data were analyzed within 77 to 91 days of collection. The simplified spreadsheet was converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-57nuts*.xls which included a report on precisions. The file was simplified, reordered on sample numbers and saved as 2013-57-nuts.csv. The event numbers were corrected. The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps. After the 3th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number, the usual method used. The output files were named MRGCLN1s.

Those files were merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The MRG files were exported to a spreadsheet for an initial check that all samples had been added correctly. No errors were found.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When outliers were excluded based on residuals a plot of differences between extracted DO samples and CTD DO versus CTD DO leads to the fit:

DO_Bottle = DOX_CTD *1.1104 - 0.0249

The outliers were investigated. Many had high standard deviations in the CTD data. The most extreme outliers were:

- Event 6 - 5 and 10m. It looks like samples #37 and #38 were reversed. There is a note on the DO analysis log sheet about them being reversed in the AutoOxy file, but it looks like this was not corrected. The OXY file was adjusted to run another comparison and a note was sent to the post-cruise DO analyst who confirmed the reversal should be corrected.

- Event 18 – 5m – sample #156 flagged 36 by analyst – both samples appear out of line and there was a flag by one of them. This was changed to 46.

- Event 27 – 200m – a spike in pressure for the first record at 200m made the COMPARE routine fail. A text editor was used to remove that record. No other spikes were found in pressure or DO.

- Event 28 – 20m – spikes in the DO made the DO value at 20m bad. A text editor was used to replace bad values with pad values in the DO channel and both conductivity channels.

The analyst found a few other errors. As more outliers were removed based on residuals, the slope changed slightly while the offset got steadily larger.

A fit that looks reasonable was:

DO_Bottle = DOX_CTD *1.1110 - 0.0324 (1)

The outliers are mostly associated with high standard deviations in the CTD data or are near odd features in DO that are likely not well resolved by the SBE sensor. If the offset is forced to zero the fit is:

DO_Bottle = DOX_CTD *1.1012 (2)

The fit through the origin looks poor even when values with DO <2mg/L are excluded.

To test for hysteresis a plot was made with bottles from above 2000db only so that the deep bottles show up in red and the shallow in green. There are only 3 samples deeper than 2000db and they do not stand out from the red points.

For more details see 2013-57-dox-comp1.xls.

Plots of CTD salinity versus titrated DO and CTD DO were made and no further outliers were found.

Salinity

Compare was run with pressure as the reference channel.

Because there were few salinity samples, the test cast (event #1) was included in the comparison.

There are 3 major outliers in both the primary and secondary salinity:

- Event #1 – This was a test cast and 24 bottles were fired at one level. It is not recorded which Niskin was sampled so it was assumed it was bottle #1. Choosing another bottle does not help as there is little variation in the CTD salinity.

- Event #35 - The bottles at 10m and 5m were fired on the fly, so it is not surprising that they are out of line. The 10m bottle has a very high standard deviation in the CTD salinity and it is fairly high for the 5m sample.

There is no need to flag these samples since event #1 is not to be archived and the outliers in event #35 are more likely due to the bottles being fired on the fly than analysis problems.

There is some pressure dependence with the largest differences at depth, but there is a lot of scatter in the differences. The pressure dependence stays the same even if the 2 deepest samples are excluded and is higher in the secondary channel..

The standard deviations during bottle stops are slightly higher in the primary salinity. Whether outliers were determined based on standard deviations during bottle stops or based on differences significantly different from the average, the results were similar. The average differences are as follows:

| Pressure Range | Primary Sal – Bottle | Secondary Sal - Bottle |
| --- | --- | --- |
| All | -0.0071 | -0.0081 |
| 0-500db | -0.0066 | -0.0075 |
| Below 500db | -0.0085 | -0.0099 |

The delay in analysis of bottle samples likely explains some of the difference and the scatter in the results as evaporation rates may be random. The errors that can arise due to poor flushing of Niskin bottles are unlikely in conditions where the bottles were moving considerably during bottle stops. Also if there were flushing problems, we would expect that the samples would read relatively high in high-gradient zones, so the differences should be more negative near the surface. The opposite is true. So flushing doesn’t seem to be a significant factor.

The differences between the two salinity channels are higher below 500db, presumably because there is more pressure dependence in the secondary channel. So unless the primary temperature/salinity pair looks considerably noisier than the secondary, it looks like the better choice for editing and archiving, but this conclusion is based on very few data with a lot of scatter.

A plot of salinity differences versus salinity suggests that there might be some non-linearity in the salinometer, but both the quantity and quality of the sample data makes such a conclusion very weak. The same salinometer has been in use for other cruises that do not show any non-linearity.

For full details for the COMPARE run see file 2013-57-sal-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2013-57-ctd.xmlcon.

Cast #1 was a test cast with different line lengths so it requires a different configuration file; since it was only a test, it was not processed.

A few casts were examined and all expected channels are present.

The temperature and conductivity channels are close on the downcast but, as usual, differ in a somewhat random fashion on the upcast. However, during stops for bottles, the differences look like those during the downcasts. While moving upwards the secondary seems slightly noisier during the primary in the few features that were examined in detail.

There are many spikes in the data. A test run of WILDEDIT removed those that had been noted, though there may be many others thus far undetected. The secondary T and C channels have more spikes after WILDEDIT than the primary do.

The altimetry is noisy at the bottom, but if the algorithm fails it should be possible to make a reasonable estimate by examining plots.

The turbidity trace looks similar to transmissivity but has much more hysteresis.

The Rinko traces look like temperature and the mirror image of dissolved oxygen, so presumably need some further processing steps. These channels will be kept for the use of researchers.

The stops for bottles were at least 30s.

Susan Allan noted that the upcast was very different from the downcast for event #2. An upcast file will be prepared for her use, but this will not be placed in the IOS Data Archive.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests are inconclusive on these data. The very noisy descent rate for most parts of most casts means that the oxygen sensor occasionally has time to “catch up” to temperature. At other times it lags. Comparing the vertical offset between the temperature and DO traces proved extremely difficult, since there are few common features between downcasts and upcasts. A few areas were found where there were a few such features and the best results varied from 4s and 6s. Another approach is to just look at the downcast and try to make features in temperature and DO match. With that criterion 4s looked better, though occasionally 5s looked best. Overall 4s looks most convincing, though no choice will suit the entire profile due to variations in gradients and temperature.

ALIGNCTD was used to advance the DO Voltage by 4s relative to the pressure.

##### CELLTM

The data are extremely noisy making the usual tests for CELLTM settings hard to interpret, especially for the secondary sensors. Tests were run on 3 casts with 5 different settings. The best results for the primary sensors was with the setting (α = 0.03, β=9) for 2 casts and (α = 0.0245, β=9.5) for the other with little difference between those 2 choices. The cast with the best upcast data (no bottle stops) was the one with the settings (α = 0.0245, β=9.5). For the secondary sensors the setting (α = 0.03, β=9) was best overall, though again (α = 0.0245, β=9.5) was similar.

CELLTM was run using (α = 0.0245, β=9.5) for the primary and (α = 0.03, β=9) for the secondary conductivity channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity, dissolved oxygen concentration and descent rate.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

The descent rate is extremely noisy most of the time with many complete reversals of direction.

At this point cast #2 was studied in detail to be sure there was no instrumental problem that made the upcast so different from the downcast. The vertical offset between downcast and upcast temperature traces varies from 0 to about 40m. There is a 40m deep well-mixed layer in the upcast that does not appear in the downcast. The temperature and salinity values in that mixed layer are seen only at about 4m in the downcast. The drift during the cast was calculated based on latitude and longitude recorded at the start, bottom and end and there was little drift during the downcast, but a drift of 0.16km during the upcast. There is more noise in the T-S plot for the upcast but otherwise it looks very close to that from the downcast. The pumps were on throughout. The transmissivity for the downcast climbs quickly while the upcast shows lower values in the upcast mixed-layer. Fluorescence is also well mixed in the upcast  with values ~2.5ug/L from 50m to the surface while the downcast quickly drops to about 0.2ug/L by 20m. So the upcast has a thick layer of warm, fresh, high DO, high Fluorescence and lower transmissivity than are seen in the downcast. The descent rate and ascent rate are both extremely noisy with many complete reversals of the CTD. The nutrients and titrated DO samples also show a well-mixed layer to about 50m. This looks like a real event such as movement through the edge of an eddy, not an instrumental problem.

##### Test Plots and Channel Check

The three deepest casts were plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2013-57-0010 | 1250 | -0.0012 | -0.00024 | -0.0016 | XNoisy, high |
| 2013-57-0011 | 1250 | -0.0015 | -0.00027 | -0.0016 | XNoisy, high |
| 2013-57-0012 | 1250 | -0.0015 | -0.00024 | -0.0012 | XNoisy, high |
| “ | 2200 | -0.0017 | -0.00026 | -0.0014 | XNoisy, high |

The temperature differences are much larger than we usually see while the salinity differences are not. The conductivity differences are only slightly larger than usual, but presumably there is one pair of sensors with compensating errors in the T and C sensors. Unfortunately, this means that the comparison with salinity bottles could not give conclusive evidence about which pair is more accurate.

A test using typical values at 1250m shows that if temperature is low by 0.0015 and conductivity low by 0.00025, salinity is low by 0.0014. The comparison with deep bottles showed the primary to be low by 0.0085 and the secondary by 0.0099, so differing by 0.0014.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  Station names were missing from events #2 to #7. Those were added.

Surface Check was run and shows an average surface pressure for the cruise was 2.7db but there was a wide range of values. A few casts were examined that included pressures <1 and the salinity recorded looks like near-surface values. But there are insufficient data to conclude that the pressure is accurate.

The cross-reference check was compared with the log book and no errors were found.

The cruise track was plotted and added to the end of this report.

The altimeter readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Despite very noisy altimetry at the bottom of casts, the header entries all looked appropriate, though for some casts there was obviously drift into waters of different depths while the CTD was at the bottom.

Water depths were missing but were available from the log book for all but 2 casts, so those were added to the headers of the CLN files.

The method used for deployment of the CTD varied. One technician started acquisition before or during the soak at 10m; the other waited until the CTD was back at the surface before recording data. Plots were made of pressure versus scan number to see which casts included soak period records. A text editor was used to remove such records from events 8, 9, 12, 16, 17, 22-25, 33, 36 and 37.

##### Shift

Fluorescence

The fluorometer was not pumped, so alignment was not required.

Conductivity

Tests were run on 3 casts to determine how to align the conductivity channels relative to temperature, so as to minimize noise in the salinity channels. A shift to the primary conductivity of -0.4 records looked best overall. For the secondary values all shifts tried made the noise in the salinity worse.

SHIFT was run on all casts using -0.4 records for the primary conductivity. The secondary conductivity was not shifted.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but as far as can be judged by such noisy data further alignment does not appear necessary.

Cast #2 was put through REVERSE and then DELETE to produce a REVDEL file.

##### DELETE

DELETE was run on all files.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: All warnings pertained to the upcasts. (events #11, 29, 35)

The DEL files were copied to EDT files in case some do not need editing.

The upcast file 2013-57-0002.revdel was renamed as 2013-57-9002.del.

##### DETAILED EDITING

The primary temperature and salinity were selected for archiving since the data appear to be slightly less spiky.

The descent rate of the CTD was extremely noisy for almost all casts.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems (chiefly miss-alignment of T and C) and likely to affect the bin-averaged values and many records corrupted by shed wakes including some surface records.

For most casts the descent rate of the CTD was extremely noisy with many complete reversals of direction.

There were some bad salinity points that were removed – these are likely due to alignment variations.

Editing was required for all casts.

The upcast file for event #2 was not edited because it was impossible to distinguish good records from bad. When a CTD is mounted below a rosette the upcast data are always of lower quality than downcast. In this case the very unsteady ascent rate (including many complete reversals) made it even worse than usual.

##### Other Comparisons

Previous experience with these sensors –We have no history for these sensors.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed and most data fell within those ranges. The only excursions in temperature were some high values between 20db and 60db during Event #40 and between 40db and 60db during the upcast of Event #2. For salinity there were some low values near the bottom of a few near-shore events and between 40 and 80db of the upcast of Event #2. The climatology is too severe for near-shore waters. These excursions are not considered evidence of calibration problems or instrumental malfunction, though the observations of the upcast data from Event #2 may be of research interest.

Repeat Casts –

There were no cases where casts were close enough in time and space to expect good repeatability.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### Initial Recalibration

No recalibration was considered necessary for pressure.

The primary salinity read low by an average of 0.007 but the analysis was run almost 3 months after collection so the result is not trusted. Based on some tests run at IOS, delayed analysis does result in bottle salinity increasing by from ~0.005 to 0.01 after 3 months. Most of the change seen is likely due to evaporation which appears to have increased in recent years with the use of silicon liners.

Given the uncertainties about the accuracy of the analysis, no recalibration will be applied was applied to the salinity.

Dissolved oxygen was recalibrated using the results of the comparison described in section 4:

DO_Bottle = DOX_CTD *1.1112 - 0.0334

CALIBRATE was run using file 2013-57-recal1.ccf to apply that correction to the dissolved oxygen channel in the SAM and MRGCLN2 files. COMPARE was rerun to check that the salinity was recalibrated appropriately and it was. The calibration was then applied to the EDT files.

(See file 2013-57-dox-comp2.xls.)

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to correct for errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The fit looks excellent below 25db with the CTD DO being low by an average of 0.004. Nearer the surface there is a lot of scatter and the CTD DO tends to be higher than bottles. This may partly be due to incomplete flushing of Niskin bottles in regions of high gradients and variations in response time in the SBE sensor and some temperature dependence in the response. No further recalibration will be applied.

Based on the final comparison a very rough estimate was made of the accuracy of the downcast Oxygen:Dissolved:SBE data:

±1  mL/L  from 0 to 30db

±0.6  mL/L  from 30 to 100db

±0.2 mL/L from 100 to 400db

±0.04 mL/L below 400db

(See 2013-57-dox-comp3.xls.)

##### Special Fluorometer Processing

There were no off-scale fluorescence data.

The fluorometer was a Wetlabs ECO and the traces were not very spiky; filtering is not useful.

There were some negative fluorescence values even after DELETE, so CLEAN was run to set those to 0.

The dissolved oxygen trace showed a spiky section in the DO channel for cast #11 and it included negative values. Values from 1431.1db to 1432.7db were replaced with pad values.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

##### Final CTD File Steps (REMOVE and HEADEDIT)

Plots were made to look for any problems with the data.

The downcast and upcast for event #2 were plotted together and there are only slight differences in T-S space while the profiles look very different. Fluorescence was much higher during the upcast between 10m and 60m while at the same depths transmissivity was lower, turbidity slightly higher and oxygen mostly much higher.

The Uploy0 and Upoly1 channels were examined and there are signals in Events #2 to 14 and 40 only. One set of traces looks like temperature and the other like the inverse of dissolved oxygen. These are for the Optode oxygen sensor mentioned in the log. The data are not to be archived at IOS. Files were kept to be delivered to the chief scientist with all channels.

REMOVE was run on all casts to remove the following channels:

Scan_Number,  Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Uploy0, Upoly1, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Turbidity and Fluorescence data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. SBE DO calibration was done using the method

described in the SeaBird Application Note #64-2 with a modification

allowing a small offset.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1  mL/L  from 0 to 30db

±0.6  mL/L  from 30 to 100db

±0.2 mL/L from 100 to 400db

±0.04 mL/L below 400db

Salinity data were not recalibrated because the delay in analysis was likely

to have affected the bottle data significantly.

For details on the processing see processing report: 2013-57-proc.doc.

The Standards Check routine was run and some problems were found and corrected; the previous step was rerun until no warnings were issued.

The cross-reference list was produced no errors were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values near the surface during offshore casts ranged from 105% to 110% which is reasonable. Inshore values ranged from 85% to 145% with the lowest values from casts with DO well mixed to 10 or 15m while the higher values were associated with higher near-surface DO gradients. High % saturation for DO was associated with high fluorescence at about 15m. These values look reasonable for this area and season.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Uploy0, Upoly1, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

A header check was run on the final files and a negative nitrate value was found – on the advice of the analyst this value was changed to 0.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. An error was found in the DO flags and some bottles without sample numbers were found in event #40, so those were removed.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up no further errors.

The track plot was produced on screen and no further errors were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

# Particulars

ALL – There is confusion over file names. They are not in sync with the log entries. This was due to a test cast that was not given an event number in the log book. Changing all file names to match the log would mean that the test cast would have event #0, but since it is not meant to be processed. Leaving the file names as they are mean they don’t match the log book, but there are frequent notes in the log book mentioning the difference. Since some preliminary processing was done using the file names as they are, it is best not to change them. It might lead to confusion for researchers who have already done some work on the data.

2. Up and downcasts very different.

9. Time in log approximate

19. Niskin top caps were not sealed.

# CRUISE SUMMARY

CTDs

|  | CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 | SEABIRD | 911+ |  | 1087 |  | Yes |  | Yes |  |  |  |
| Calibration Information CTD #1087 |  |  |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  |  |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  |  |  | 5483 |  | 23Feb2012 |  | Factory |  |  |  |  |
| Conductivity |  |  |  | 3916 |  | 15Mar2012 |  | Factory |  |  |  |  |
| Secondary Temp. |  |  |  | 5549 |  | 13Mar2012 |  | Factory |  |  |  |  |
| Secondary Cond. |  |  |  | 4003 |  | 07Mar2012 |  | Factory |  |  |  |  |
| Transmissometer |  |  |  | 1476DR |  | 13Oct2011 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  |  |  | 2331 |  | 26Feb2013 |  | Factory |  |  |  |  |
| WetLabs Fluorometer |  |  |  | 2578 |  | 6Mar2012 |  | Factory |  |  |  |  |
| WetLabs Turbidity |  |  |  | 2578 |  | 6Mar2012 |  | Factory |  |  |  |  |
| ARO-CAV DO |  |  |  | 0124 |  | 19Mar2013 |  | ? |  |  |  |  |
| Rinko III Aro-Cav |  |  |  | 0124 |  | 19Mar2013 |  | ? |  |  |  |  |
| Pressure Sensor |  |  |  | 1087 |  | 22Mar2011 |  | Factory |  |  |  |  |
| Altimeter |  |  |  | 59545 |  | 14Mar2013 |  | ? |  |  |  |  |