## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April 2025 | ### Updated channel names & formats in TOB files. GG |
| ### 22 Nov 2021 | ### Corrected Salinity:Bottle formats. S.H. |
| ### 20 Aug 2020 | ### HPLC Data added. S.H. |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |
| 11 August 2014 | Correction to CHL data in file 2013-58-0011.CHE and CHL spreadsheet; updated statistics in NH4 spreadsheet. |
| 13 June 2013 | Fluorescence data were removed from CTD and Bottle files due to further study showing the fluorometer malfunctioned. See section 5 for details. G.Gatien |

## PROCESSING NOTES

Cruise: 2013-58

Agency: OSD

Location: WCVI

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: 5 September 2013 –16 September 2013

Processed by: Germaine Gatien

Date of Processing: 24 March 2014 – 4 April 2014

Number of original HEX files: 93	Number of CTD files: 92 (1 file upcast only)

Number of bottle files: 	54		Number of bottle casts processed: 54

Number of original TSG files: 4 	Number of processed TSG files: 2 (2 had only few records each)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#2228) with a 10X cable on the primary pump, a Biospherical QSP-200 PAR sensor (#4615) and an altimeter (#1204).

A thermosalinograph (SeaBird 21 S/N 3363) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor and a flow meter.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette sampling log sheets were generally in good order. Sampling notes were provided which were helpful.

There was one split cast that was only noted in the log by having an odd event number (65/66). It is helpful to have a clear statement when a new file has to be started during a cast. When this happens during the upcast section of a rosette cast after some Niskin bottles have been closed, it is very helpful to fire again the previously closed bottles. This makes the addition of sample numbers to the appropriate Niskin bottles much easier and has no effect on the closed bottles. There were no CTD available  right at the time Niskin #9 was fired, but data acquired shortly before firing were entered in the CHE file.

An error was made in sample numbers on labels for event #32. Corrections were made that look appropriate, with salinity and dissolved oxygen showing good correspondence to CTD data. Flag 2 was added to all the samples from this cast to indicate that there was a problem, but the quality of the data are considered good.

The comparison of CTD salinity with bottle samples indicated that the primary salinity was low by ~0.0005 and the secondary salinity high by ~0.0048. Since the analysis was run about 2 months after collection there was likely some evaporation of samples. If the bottles are reading too high, this would imply that the secondary sensors are reading even higher than the comparison suggests. The difference between the two salinity channels is similar to what was observed to that seen during its previous use in 2013-18, but is markedly different from the cruise before that, 2013-50, during which they were both within 0.001 of bottles. There is some evidence that one sensor might have been damaged during a touch-down late in 2013-50. It seems more likely that one conductivity sensor had changed calibration suddenly than that two sensors had drifted fairly quickly. The primary salinity was less noisy than the secondary and likely more accurate, so it was selected for archiving. No recalibration was applied to the salinity.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.7mL/L from 0 to 30db

±0.5mL/L from 30db to 200db

±0.3mL/L from 200db to 300db

±0.1mL/L from 300db to 1500db

±0.03mL/L below 1500db

The SBE Fluorometer in use for this cruise produced odd results in 2013 with values relative to extracted chlorophyll being lower than in the past. The results for this cruise confirm that there is a serious problem with the sensor and it is getting worse. A comparison with extracted chlorophyll shows that while fluorescence is a little high for very low CHL values, it falls rapidly so that it is only about 25% of CHL for CHL=5ug/L and about 15% of CHL for the highest CHL values. The channel was removed.

There continue to be serious problems with the Thermosalinograph. The chief scientist noted in the cruise report that there was air in the system due to a switch to an old (inappropriate) pump so the bubble trap was replumbed; this had only a limited effect. Bubbles continued until the system was switched to the secondary filter system, indicating that there is an air leak in the 'outer' filter arrangement. The flow rate was often low and for a few hours flow was turned off.

There were no loop samples available so checks on TSG data can only be made during stops for CTDs. The TSG and CTD depths are not matched exactly, and the TSG system may operate differently in motion. A post-cruise calibration of conductivity and temperature was available. TSG salinity was very noisy with spikes to lower values, likely due to bubbles. The TSG fluorometer values are a little higher than those from the CTD fluorometer but lower than extracted chlorophyll samples taken from 5m rosette bottles when CHL >1ug/L; the low values may be related to the bubble problem and/or low flow rates. The TSG intake temperature compares reasonably well with the CTD temperature.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets as well as various analysis logs were in good order. There were notes about a few header errors and those were fixed in the raw files. There were notes about when the PAR sensor was moved on or off the CTD. However, the PAR must have come off sometime between casts #97 and #122 – a check of data files shows it was removed after cast #103.

There is a note “syringes on” beside events #86 and #139. For the first this was clearly a purposeful step as the CTD was not to be used for some time, but the latter instance is unclear since the note appears just before a CTD cast. The data look ok, so it will be presumed that the syringes came off before deployment.

There is a configuration file and BL file for event #162 but no hex file. There is an entry in the log for the cast but sampling notes provided indicate the file was lost for this and for event #164 which is not entered in the log.

Extracted chlorophyll, NH4, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The same calibration control file was used throughout the cruise. The calibration constants were checked for all instruments. No errors were found except that the

Based on log notes cast lists were prepared with and without PAR so that it will be easy to remove PAR as appropriate.

##### Hysteresis Study

Hysteresis tests were run for cruise 2013-17 and no adjustments of DO parameters was found necessary. A quick check was also made for 2013-18 and there was no evidence of hysteresis.

##### BOTTLE FILE PREPARATION

The ROS files were converted using file 2013-58-ctd.xmlcon with Tau and Hysteresis corrections selected and choosing the particular version by following the scheme provided by the chief scientist.

Those files were put through CLEAN to add event numbers (*.BOT).

Header Check was run on the BOT files and no problems were found.

Temperature and salinity were plotted for all BOT files. A few noisy primary salinity points were cleaned in file #84. The data from cast #153 indicate that the CTD was in motion when most bottles were closed. Notes on the rosette log indicate that only the surface bottles were intended to be sampled, so this is not significant.

The BOT files were averaged to enable an ADDSAMP file to be prepared. First the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. The sample numbers generally are in the order of bottle position though in a few cases sample #s were added or dropped, so checks were made against every sheet.

One cast is identified as 65/66 and this proved to be a split cast. Unfortunately when the second file was started the previously closed bottles were not fired. It is very helpful to refire such bottles as it makes the addition of sample numbers to the appropriate Niskin bottles much easier.

For events #65/66 there are no data for Niskin #9 in the CTD file, but there is a sample with that number.

The BOT files for casts 65 and 66 were merged. The merged file was opened in Ultraedit and bottle numbers were adjusted so that bottles 1-8 of file #66 became bottles #9-16 of the combined file #65. The bottle at 150m does not appear in this list but was obviously fired. There may be some data from the full data files that may be patched into the CHE file later.

SORT was used to ensure that the ADDSAMP file was in sample number order.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was ordered on sample number and converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2013-58-bot-hdr.txt.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2013-58.xls. Event numbers were added. A simplified version of the spreadsheet was saved as 2013-58chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2013-58oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2013-58oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2013-58SAL.xls. The file was simplified and the file was saved as 2013-58sal.csv. That file was converted to individual SAL files. The salinity data were analyzed 58 to 68 days after collection.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2013-58nuts.xls which included a report on precision.  The file was simplified, reordered on sample numbers and the file was saved as 2013-58nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2013-58_NH4.xls which included a report on precision. The file was simplified and saved as 2013-58-NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

At this point the full file was used to fabricate a line for the missing bottle of cast #65. There were data from the right level and the CTD was clearly stopped during that period. 100 records were averaged and added to the SAMAVG file for event #65. The MRGCLN1s file was edited, adding a line for Niskin #9 and adding the results of the analyses done for sample #295. The MRGCLN1s and SAMAVG files were merged on Bottle:Position and the results look fine. Flag “2” were added to the analyses for sample #295 just to alert users that while the sample values are fine, the associated CTD data are estimates.

A few problems were found in the course of running COMPARE:

- The ADDSAMP file needed correction

- For the Dissolved Oxygen data, samples #131-134 were in the analysis spreadsheet as coming from cast #32 and but those sample numbers are supposed to be for event #31. Event #31 did not have any DO sampling and since there was only surface sampling, it is clear that the values given for those sample numbers do not make sense. They do make sense if we assume that the data given for event #32 are from samples 135-152 rather than 131-148. So they were renumbered. Further investigation shows that this mislabelling occurred for other samples too, though in some cases the analysts discovered the correct attribution. All samples numbers were fixed for all variables and flag 2 added and a comment.

- File #39 was saved as file #38 for CTD data but all analyses had #39, so were missed in the merge.

- File #86 has many bottles that were fired at the surface but not sampled and no sample numbers were assigned. Those lines were removed from the SAMAVG file.

- File #153 has many bottles that were fired at depth but not sampled and no sample numbers were assigned. Those lines were removed from the SAMAVG file.

After correcting these errors the MERGE process was rerun as needed.

An initial header check did not turn up any problems.

##### Compare

Fluorescence  (Note: this section was amended in June 2014)

The comparison of SBE fluorescence versus extracted chlorophyll is markedly different from those found for this sensor in 2012. While the fluorescence is higher than extracted chlorophyll for very low CHL, it is not as high as was routinely seen in the past. The ratio of fluorescence to CHL drops rapidly so that for CHL>5 the fluorescence is only 25% of CHL. It eventually drops to about 15% for the highest CHL values. The sensor produced some very bad results during 2013-18 and rather large dark values on a number of 2013 cruises. The sensor has a serious problem and the channel should be removed.

For more detail see file 2013-58-fl-chl-comp.xlsx.

Salinity

Compare was run with pressure as reference channel. Salinity samples were analyzed about 2 months after collection. The first run turned up a reversal of the 2 samples from cast #65. That was fixed and the program was rerun.

There is a lot of scatter in the fit. Outliers were generally near the surface where flushing of Niskin bottles could be an issue and the CTD data were noisy. Both poor flushing and delayed analysis tend to lead to the CTD looking lower than it really is. There are some deep outliers with CTD reading higher than bottles. When outliers were removed, the primary salinity was found to be low by an average of 0.0005 and the secondary salinity was high by 0.0048. The standard deviations in those averages were ~0.002. There was pressure dependence in the result but it is identical in both sensor pairs so is unlikely to be due to calibration drift. All bottles above 500db are from a single cast. The precision study found a mean difference of 0.0018. The display of differences against time shows no trend but does show tremendous variability for the cast with the most samples; however, that is the only cast with samples above 500m.

This shows that the primary salinity is closest to bottles. If we reject all the values showing the salinity to be lower than bottles on the basis evaporation of samples and inefficient flushing, then the primary is still close to the CTD, reading high by about 0.001. The secondary salinity reads lower than bottles only in the upper 75m.

There is slightly more time dependence in the secondary than the primary salinity, but the difference is small and the scatter in the fit large.

For full details for the COMPARE run see file 2013-58-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. The first run turned up a problem with sample numbers for event #32. When those were fixed the results were good. There was only one major outliers, sample #371, event #86. This was one of duplicates, but the other value was rejected by the analyst and is even worse. The value is also out of line in profile.

The fit of differences against DO concentration looked good. There were a few other minor outliers, but they are all associated with noisy CTD DO data.

For this cruise the fits found were

DOX_BOT = 1.056 * DOX_CTD +0.0214 with R2 =0.9049

or

DOX_BOT = 1.0608 * DOX_CTD when a zero offset is forced and R2 =0.8950

The fit with a free offset looks better as well as having a higher R2 value.

There were many problems with the fit for 2013-18. When many outliers were excluded the fits were:

DOX_BOT = 1.052 * DOX_CTD +0.0163 with R2 =0.9501

or

DOX_BOT = 1.0555 * DOX_CTD when a zero offset is forced and R2 =0.9447

A fit was run with data from above 1800m excluded to see if there was any sign of hysteresis and there was not, though there are not really enough deep data to make this result significant.

For more details see 2013-58-dox-comp1.xls.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were detected.

Data from the MRG files were exported to a spreadsheet and casts were checked against rosette sheets to ensure all expected bottle data were present and to find any errors in the assignment of sample #s to bottle numbers by assessing if profiles look reasonable. No errors were found.

##### Conversion of Full Files from Raw Data

All files were converted using configuration file 2013-58-ctd.xmlcon with the hysteresis correction used.

After conversion file 2013-58-0038.cnv was changed to 2013-58-0039.cnv.

A few casts were examined and all expected channels are present.

The two temperature and conductivity channels are fairly smooth during the downcasts but noisy in the upcasts with some spikes and odd excursions in both channels. Downcast temperature and conductivity channel pairs are usually close but in high gradients show some vertical offset.

Fluorescence looks normal with a dark value of 0.13ug/L.

The altimetry, transmissivity, dissolved oxygen and PAR data look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

- When used on other recent cruises the DO sensor data were found to be improved by using advances of +4s, +4.3s and +4.5s. Tests were run on a few casts with few stops during the upcast but the results are not easy to interpret because of noisy upcast temperature data. The +4.5s was used for the previous cruise and is probably ok.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### CELLTM

The upcast data are noisy so the usual tests for CELLTM settings were also hard to interpret. Tests were run on a few casts using a variety of settings. The setting recommended by SeaBird does improve the data, so that was selected and was the choice made for all recent uses of this equipment.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity, dissolved oxygen concentration and descent rate.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are always very noisy, so these estimates are extremely rough. Data are also included from 3 other 2013 cruises that used this equipment and sampled to at least 1800m.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2013-38-0111 | 1900 | -0.0002 | 0.00001 | 0.0018 | High, Moderate |
| 2013-17-0025 | 1800 | -0.0003 | 0.00013 | 0.0018 | High, X Noisy |
| 2013-17-0036 | 1800 | -0.0001 | 0.00013 | 0.0018 | High, X Noisy |
| 2013-18-0042 | 1800 | -0.0001 | 0.0005 | 0.0064 | High, Noisy |
| 2013-18-0059 | 1800 | -0.0001 | 0.0005 | 0.0065 | High, Noisy |
| 2013-58-0037 | 1800 | +0.0005 | 0.0005 | 0.0058 | High, XNoisy |
| 2013-58-0075 | 1800 | +0.0003 | 0.0005 | 0.0060 | High, Noisy |
| 2013-58-0076 | 1800 | -0.0004 | 0.0005 | 0.0061 | High, Moderate |
| 2013-58-0091 | 1800 | -0.0002 | 0.0004 | 0.0055 | High, Noisy |
| 2013-58-0109 | 1800 | +0.0001 | 0.0005 | 0.0055 | High, Noisy |
| 2013-58-0111 | 1800 | -0.0003 | 0.0004 | 0.0058 | High, Moderate |

The temperature differences are not large, but they are much noisier and vary more from cast to cast than during earlier cruises. As was noticed during the processing of 2013-18 there was a shift in conductivity and salinity differences between 2013-58 and 2013-18. Cruise 2013-50 occurred between those 2 but included no deep calibration sampling, but there was some weak evidence that there may have been a change during that cruise. The salinity differences appear to be slightly lower for 2013-58 than 2013-18. The temperature differences vary not only from cast to cast but also within casts, looking noisier than usual even in areas where the descent rate was high and relatively steady. So the change in temperature differences might be due to one temperature sensor not performing well. While there are certainly differences between the two channels in plots, there is no way to determine which is better. The secondary salinity appears to be noisier than the primary, which may suggest that the secondary temperature is less reliable. But there could be other reasons for the noisy salinity such as alignment.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

Cast #26 could not be converted. The upcast contained many bad values, so most of the upcast data were removed and then conversion was successful.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  No problems were found.

The surface check indicated an average surface value of 2.8db which is reasonable. The CTD pressure was measured on deck before (0.65db) and after (0.48db) deployment during 1 cast. This may indicate that the pressure is reading too high by about 0.5db. During 2013-50 there were many pressure checks and the values varied over a range of about 0.4db with the upcast readings showing less variability. At that time the pressure appeared to be within 0.2db and no recalibration was applied. If the pressure really is high by 0.5db that means the average surface value would be 2.3db which seems rather shallow. Usually pressure drifts towards low values. No recalibration is justified yet, but an eye should be kept on this sensor.

The cross-reference check was compared with the log book and no problems were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings, maximum pressure and water depth were exported from the headers of the CLN and MRG files to spreadsheets. The water depths were checked against the log and a few corrections were made where a depth was wrong.  Where the CTD did not get within 15m of the bottom there should be no header altimetry entry and that was mostly the case. But occasionally there are spikes that are misinterpreted, so in cases where the CTD did not appear to get near the bottom, plots were made and the header removed when it was clear the data came from spikes. Plots were also made of some random casts. The same corrections were made to the SAMAVG files in case the MERGE process needs to be repeated in future.

##### Shift

Fluorescence

Normally a shift of +24 records is used to align fluorescence with temperature when the fluorometer is pumped. A few casts were tested to see if that setting improved alignment and the results showed that it did, though it is very hard to judge given the noise in the upcast temperature.

SHIFT was run on all casts to advance the SeaPoint fluorescence by +24 records.

Conductivity

For 2013-18 when these sensors were last used, primary and secondary conductivity were advanced by -0.7 and +0.4 records respectively. For some other 2013 cruises with this equipment different shifts were found appropriate. Tests were run on 3 deep casts using a variety of settings close to those used recently. The results looked best using -0.7 records for the primary conductivity though results vary from feature to feature. For the secondary the +0.4 setting looks better than smaller or larger shifts, but much noise remained no matter what choice was made.

Two runs of SHIFT were used to apply advance the primary by -0.7 records and the secondary by +0.4 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The alignment does not look very good, but that is due to extremely noisy upcast DO data. The alignment is likely the best we can achieve.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The conductivity sensors were both recalibrated in late January 2013 and were used on 4 previous cruises. The analysis of salinity calibration samples was done fairly quickly for 3 of them. The primary salinity was found to be low by 0.0005 for the 2 with the quickest analysis (2 to 4 weeks delay) and low by 0.0017 when there was a 3 to 6 week delay. For 2013-18 with a 2 to 3 month delay in analysis, the primary salinity was found to be low by an average of 0.0026..The secondary was high by about 0.0006 and 0.0010 when analysis was quick and low by 0.0005 when it was slightly delayed. For 2013-18 the secondary was high by an average of 0.0035. This gives some indication that the error associated with a delay of a few weeks is on the order of 0.001salinity units. But it is also clear that there was a change between June and August 2013 because the difference between the two salinity channels became much higher.

2. Dissolved Oxygen

The DO sensor was recalibrated in February 2013and has been used for 4 other cruises since then. The offset was forced to be 0 when there was little low DO data or data quality was lower than usual. The fits found were:

DOX_BOT = 1.0471 * DOX_CTD +0.0172 (2013-38)

DOX_BOT = 1.054 * DOX_CTD  (2013-17)

DOX_BOT = 1.0545 * DOX_CTD (2013-50)

DOX_BOT = 1.0528 * DOX_CTD +0.0123 (2013-18)

3. Pressure

The pressure sensor was recalibrated in March 2013 and the factory offset has been used for the 4 cruises on which it has been used prior to this one. Deck observations during 2013-50 showed it to be within 0.2db.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. A few salinity values were below the climatology minimum but these were only in the top 10m and near shore. The temperature data were all within the climatology except for high temperatures in the top 5 to 30m of casts along the LC and LB lines inshore of the 1000m contour. Theses look like real excursions, not evidence of calibration errors.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

The first issue is to decide which sensor pair to edit. During 2013-18 the primary sensors were selected because there was less noise in the salinity and the primary salinity was closer to calibration samples, though how close was not known. For this cruise the analysis delay was about 2 months, so slightly less than for 2013-18. The primary salinity was found to be low by about 0.0005 and the secondary high by 0.0048, but both comparisons had standard deviations of 0.002. So the difference between the two channels is slightly lower than for 2013-18 which agrees with the observations of section 10. If there is evaporation of samples then both are likely too high, but the primary would be closer.

The noise levels are high in both channels but slightly less in the primary, so it was selected for archiving.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. The descent rate was often noisy.

As noted on other recent Tully cruises, there are many odd excursions in salinity that do not appear to be related to shed wake corruption or stray spikes. These unstable features look like they are caused by misalignment of temperature and conductivity, possibly related to high deceleration rates of the CTD. Descent rates often vary from -1.5m/s to -0.7m/s within 3m, so within 2 or 3 seconds, and occasionally are much higher than that, most often during deeper casts.

In some cases these features are too large and complex to be corrected by cleaning, so the salinity values were removed. Temperature points were not removed for those features. Such editing was applied to most casts.

All ED1 files were copied to EDT.

##### Initial Recalibration

The pressure does not need recalibration.

Based on the analysis in section 16, no calibration will be applied to the primary salinity. If, in future, a factory report on the primary conductivity sensor shows that it had drifted significantly and the secondary had not, then this decision can be revisited.

Dissolved oxygen will be recalibrated using the fit:

DOX_BOT = 1.056 * DOX_CTD +0.0214

CALIBRATE was run on the MRGCLN2, SAM and EDT files using that file.

COMPARE was rerun to check that the DO correction was done correctly and it was. The average difference after removing a few outliers shows the CTD DO to be within 0.0023 of the bottles on average, though the standard deviation is 0.05mL/L. (See file 2013-58-sal-comp2.xls.)

##### Final Calibration of DO

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. There is a lot of scatter in the plot of differences versus CTD DO with the trend line varying greatly depending on how outliers are identified. A fit with a free offset does not look very good for DO<1mL/L. A fit with the offset = 0 and outliers identified based on residuals leads to the correction:

DOX_BOT corrected = 0.989 * DOX_CTD

It is questionable to apply a second correction given the scatter in the fit except for the fact that it is very similar to the result found for 2013-18 when the correction factor was 0.0991. To test the correction it was applied to the thinned files using file 2013-58-recal2.ccf and then COMPARE was rerun. The CTD DO values are still slightly high at the high end but the average difference is reduced from 0.04mL/L to 0.015mL/L. Unfortunately, the low end values look worse and there is considerable interest in resolving the oxygen minimum values. This 2nd correction does not appear justified. For more detail see 2013-58-dox-comp3.xls and 2013-58-dox-comp4.xls.

A plot of differences against pressure was used to make a final study of outliers and to make a rough estimate of the accuracy of downcast CTD dissolved oxygen.

##### Special Fluorometer Processing

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The files were bin-averaged (0.25db bins) and put through HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved.

A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channels in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and while there were a few small unstable features, these were in areas where that is expected; no simple editing would make them stable.

##### Final CTD File Steps (REMOVE and HEADEDIT)

A check was made of the cast lists with and without PAR sensor by plotting PAR. A few errors were found and corrected.

REMOVE was run on casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

Profile plots were examined on-screen. There was a large reversal in dissolved oxygen around 30m for cast #80 (LG5), but examination of the full file shows that it is seen in both the downcast and upcast and there are associated small reversals in fluorescence and transmissivity. A similar but smaller feature is seen in cast #99 (LBP2). There are a few other casts along the LD line in slightly deeper water with reversals found a little deeper and smaller and once again transmissivity also reverses. There are many small reversals in other casts, but #80 is the most dramatic with a very sharp gradient. These look like real features, so no editing was applied.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section, to fix the project name and to add the following comments:

---------------------------------------------------------------------------------

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and

unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Based on the recommendation from SeaBird, the method for calibration of

Dissolved Oxygen concentration was changed from that used for 2011

and some 2012 cruises. For more information see the SeaBird Application

Note #64-2, June 2012 revision.

The two CTD salinity channels differed by about 0.005 with one reading

slightly lower than bottles and the other higher. Delay in analysis likely

caused some evaporation of samples, so the CTD salinity channels are

likely reading higher than the comparison suggests. The little information

available suggests that there has been minimal calibration drift in

the primary salinity.

The Oxygen:Dissolved:SBE data are considered, roughly, to be:

±0.7mL/L from 0 to 30db

±0.5mL/L from 30db to 200db

±0.3mL/L from 200db to 300db

±0.1mL/L from 300db to 1500db

±0.03mL/L below 1500db

For details on the processing see the report: 2013-58-proc.doc.

---------------------------------------------------------------------------------

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged widely from 75% to almost 145%. The lowest and highest values were in inlets and close to shore. Most offshore values were about 105%.  These results look reasonable as fluorescence shows biological activity at the sites with high dissolved oxygen saturation. For example at cast #56 where the DO saturation was ~130% the SBE dissolved oxygen was ~7.8mL/L and a bottle taken near the surface during the upcast had a value of 8.2mL/L.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets and no errors were found other than some cases where changes were made in sampling but not noted on the log sheets.

A header check was run and no errors were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found.

##### Thermosalinograph Data

Data were provided in 4 hex files, but the first two contained only a few records each.

The external thermistor was connected.

There were no loop samples.

The following note is taken from the cruise report:

The only (ongoing) issue was the Thermosalinograph, it still has problems. In the beginning there were huge problems with air in the system as they had to switch to an old (inappropriate) pump so the bubble trap was replumbed. That had only limited effect... Bubble continued to plague the system until we switched to the secondary filter system, indicating that there is an air leak in the 'outer' filter arrangement. Then at some point it was discovered that the outflow valve of the TSG was closed for a period of time.

Indications are that the flow meter should be moved to measure the flow after the filters, before going into the instrument, and also to capture the fluorometer.

a.) Checking calibrations

The configuration files for the 4 casts are identical and the parameters are correct.

One file was saved as 2013-58-tsg.xmlcon.

b.) Conversion of Files

The files were converted to CNV files using configuration file 2013-58-tsg.con. They were then converted to IOS HEADER format.

The first 2 files contained only a few records each and will not be processed further.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Time-series plots were produced.

File #3 – The flow rate is very noisy. At first it was very low, ~0.5, and after about 12 hours it increased to ~0.7. Temperature and salinity both increased notably when the flow rate increased, but this also corresponds to the time the ship reached the west coast, so is likely not be a result of flow rate change. The flow stayed at 0.7 and sometimes closer to 1 for about 1.5 days. The flow was turned off at 22:00 on Sept. 8. When it was turned on again the flow was smoother and very high for a few hours and then it dropped to about 0.7 for the rest of the file, but remained smooth.

File #4 – The flow rate was ~0.7 and smooth throughout.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2013-58-ctd-tsg-comp.xls.

89 CTD casts with data from 4m overlapped with TSG data. The 2 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, temperature difference, salinity and fluorescence and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all <0.0004° and the median differences in both are 0.00000°.  So the TSG time and positions appear to be good.

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Rosette samples and TSG and CTD data

- T1 vs T2 Examination of the full TSG records shows great variability.

3. File #1 is the more complex one. During the first section the median difference/standard deviation is -0.33C°/ 0.51C°. Then there was a quiet section -0.27 C°/0.07 C°. That much heating in the loop is slightly higher than we would expect, but given the flow rate it may be reasonable. After that there is a short section when the flow was off and the intake temperature is higher than the lab temperature. When the flow started again, the differences were noisy but reasonable for a short time then became increasingly negative reaching differences ~ -8 C°. There was some improvement towards the end of the cruise, but the differences were still extremely high.

4. File #4 has variable noisiness but the median for the whole file is reasonable at -0.25 C°; the standard deviation of 0.66 C°. In one very quiet section the median was -0.20 C° with a standard deviation of 0.11 C°.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

1. Intake Temperature The median difference between the TSG intake and lab temperatures during stops for CTDs was -0.90C° but the standard deviation was 1.6 with differences ranging from -6.5 to +1.3 C°. When the differences were plotted against flow rate, most are close to -0.9 when flow >0.8. Below that the variability is very high. A group of casts was identified where the differences looked steady on the plot against event # and the median was -0.26C° with standard deviation of 0.04C°. That section was one where the flow rate was >0.8 and was also noted above as looking quiet when the underway data are included. That difference is a little higher than we might expect at this time of year when intake temperatures are closer to those on the ship, but the lower flow rate probably explains that. The comparison was simplified by removing cases of flow rate <0.7 and a few outliers. The TSG intake temperature was then found to be higher than the CTD temperature by a median of 0.013C°, but the standard deviation was 0.112 C°. When only 9 casts with flow rates >0.7 and well-mixed surface waters are selected, the intake temperature is high by 0.0099C°.

2. LAB TEMP Using the same events as used to compare the intake temperature with the CTD, the median heating in the loop was found to be 0.238C° and the lab temperature higher than the CTD by 0.254 C°.

3. SALINITY  TSG salinity data are extremely noisy. The spikes in the data are unidirectional so identifying outliers was based on removing cases with large differences and there is no obvious method to determine how far one should go. If we remove most of the large differences, then the salinity is low by ~0.2. If those are reduced to only 9 casts with flow rates >0.7 and well-mixed surface waters, the TSG salinity is lower than the CTD by a median value of 0.14. There was a problem with bubbles during this cruise, so this is most likely the reason for so many low TSG salinity values.

4. FLUORESCENCE Low flow rates are expected to be a problem for TSG fluorescence with bubbles as another possible source of error. The TSG fluorescence is higher than the CTD fluorescence by from 10% to 45% if only the higher flow rate data are included. The TSG is found to be significantly lower than the CTD fluorescence if all data are included. The lowest flow rates are associated with low TSG fluorescence. When 6 casts with well-mixed surface waters and high flow rate are used the TSG fluorescence is higher than the CTD by a median of 45%. Keep in mind that the CTD sensor did not perform well, reading low when CHL was >1ug/L but reading high for low CHL.

(See 2013-58-ctd-tsg-comp.xls.)

Loop Bottle - TSG Comparisons There were no loop samples.

- Surface rosette samples vs TSG observations

There was only 1near-surface rosette salinity sample that coincided with TSG sampling and for that case (cast #109) the TSG salinity was lower than the rosette sample by 0.56.

There were many near-surface CHL samples but they are either higher or lower than the TSG intake level. Values around 5m were compared to the TSG values.

The median value of the ratio of the TSG fluorescence to the extracted CHL was 0.49 and the ratio varied from 0.05 to 1.6. When the ratio is plotted against CHL values, we get the usual shape observed for the fluorometers with the ratio lower for high CHL, but the fluorescence is lower than usual. During the spring La Perouse cruise fluorescence was 50% of the CHL when CHL=20ug/L whereas it is 5% for this cruise. This may partly be due to a mismatch in depth and local variability. The CTD fluorescence during cast #86 varied from 1.5 to 5ug/L during the 5m stop, but the CHL sample was ~20ug/L, so the mismatch is unlikely to explain that. There cruise report notes problems with air in the loop and the flow rate was low, both of which may well affect the TSG fluorescence.

(See 2013-58-rosette-tsg-comp.xls.)

- Calibration History

The TSG pressure, temperature and conductivity were recalibrated in March 2012 and it was used during 2013-01, 2013-38, 2013-17 and 2013-18.

There was a post-cruise calibration in December 2013 and the salinity was found to have drifted low by about 0.0088 due to conductivity drift and temperature was low by ~0.0006 for a net drift in salinity of about -0.008.

Conclusions

1. The TSG clock appears to have worked well.

2. The flow rate varied greatly and was too low for much of the cruise. There was a section of zero flow in file #3 for which T0, T1, Salinity and Fluorescence data need to be removed.

3. The comparisons between the TSG and CTD are extremely noisy, with the exception of a group of casts during which the flow rate was reasonably high and the near-surface gradients low so that slight mismatches in depth are not critical. From that small group of casts we find that the TSG temperature is high by ~0.01 C°, the salinity low by 0.14 and the fluorescence ~45% higher than that from the CTD sensor. Since the CTD sensor was found to be low relative to CHL sampling, the TSG fluorometer is likely better than this comparison suggests.

4. The comparison with samples from the 5m rosette is only reliable where surface waters are well mixed since the TSG and CTD observations don’t match in time. The TSG salinity was lower by 0.56 than the only salinity sample. There were many CHL samples and they show that the TSG fluorometer to be reading a little high for very low CHL and low for CHL>1ug/L and much lower for CHL>2ug/L. The lack of loop samples hampers this study.

5. The temperature in the loop increases by about 0.25Cº when the flow rate is fairly high.

6. The TSG intake temperature appears to be higher than the CTD temperature by ~0.01Cº. A post-cruise calibration shows the TSG lab temperature to be too low by 0.0006.

7. The TSG Salinity is lower than that of the CTD by from about 0.14 when only a few casts with high flow rate and well-mixed surface waters are used. The presence of bubbles likely accounts for so many low TSG salinity values. The post cruise calibration shows that the salinity was low by 0.0088 by December 2013 based on the conductivity calibration drift. This is partly offset by a drift downwards of lab temperature. The net drift is -0.008. A reasonable choice for recalibration is +0.0075 as some drift may have occurred between September and December.

6. The TSG fluorescence is much lower than the CHL samples from the rosette and higher than the fluorescence from the CTD sensor.

f.) Editing

File 3 was opened in Ultraedit and the intake temperature, lab temperature, temperature difference, salinity and fluorescence data were replaced with pad values where fluorescence was near-zero (<0.001 and for a few records after the flow was turned back on.)

The ATC files were copied to *.EDT.

Salinity was too noisy to edit. A few points in file #3 were removed close to the time of flow being turned off or on. The output file was copied to *.EDT.

File #4 was not edited.

Plots were examined and no further editing was deemed necessary.

g.) Recalibration

CALIBRATE was run using file 2013-58-tsg-recal1.ccf to add 0.0075 to the salinity channel.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

The flow channel was not removed since it explains the gaps and some users may want to remove even more data based on this variable.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it still looks fine.

June 13, 2014: Based on analysis described in section 5, the fluorescence channel was removed from the CTD and Bottle files. G.Gatien

# Particulars (mostly notes from log book)

PAR on: 1-29, 42- 68, 83-86, 97-103, 124-162

PAR off: 30-41, 69-82, 90-96, 106 -122

1. Stop at 100m caused spike in all channels.

39. File saved as event 38.

# 44. TSG turned back on  at beginning of cast

76. Station name wrong in file – LD08 should be LG08

77. Station name wrong in file – LD07 should be LG07

79. Station name wrong in file – LC06 should be LG06

112. TSG restarted – someone had closed the outflow valve on the TSG. It was reopened and file #4 was started.

151. Test of pressure – acquisition started on deck and ended on deck.

153. Error in depth in header.

162. Event in log but no HEX file

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 2280 | 29Jan2013 | Factory |  |  |
| Secondary Temp. | 2668 | 1Feb2013 | Factory |  |  |
| Secondary Cond. | 2754 | 30Jan2013 | Factory |  |  |
| Transmissometer | 1396DR | 31Jan2013 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 16Feb2013 | Factory |  |  |
| SeaBird Fluorometer | 2228 |  |  |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Pressure Sensor | 63507 | 15Apr2011 | Factory |  |  |
| Altimeter | 1204 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2013-58

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 7Mar12 | Factory |  |  |
| Conductivity | 3363 | 7Mar12 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | Aug12 | IOS |  |  |
| Temperature:Secondary | 0603 | 7Mar12 | “ |  |  |
| Flow meter | ? | n/a |  |  |  |