## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: CFMETR Nanoose

Agency: Canadian Forces

Project: Canadian Forces Maritime Experimental Test Range

Area: Nanoose

Date: 10 January 2013 – 13 December 2013

Platform: CFAV Stikine (and others)

Processed by: Germaine Gatien

Date of Processing: 27 July 2015 – 30 July 2015

Number of original hex files casts: 73 (including some repeats and tests)

Number of casts processed: 62

INSTRUMENT SUMMARY

Five Sea Bird Model SBE 19 SEACAT CTDs were used – serial numbers #1451, 1511 and 1513. Each CTD had an altimeter listed in the configuration file, but the serial numbers were not always recorded.

SUMMARY OF QUALITY AND CONCERNS

Three different CTDs were used. The calibration information could not be confirmed.

To assist in analysis event numbers are helpful and the software in use will only allow 4 digits which suits most cruises. However a 7-character system is used for this project. A 4-character format was derived from those IDs based on the month and day of each cast. This means the same event # will appear if there is more than one cast on one day as did sometimes occur.

The soak period for the CTD was typically 1 -1.5 minutes long, but occasionally less than 1 minute. It is recommended that the SBE19 be soaked for at least 2 minutes to improve salinity quality, though in the absence of large differences between air and water temperatures it is likely that 1 minute is sufficient.

There was no salinity calibration sampling. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. Away from high gradients errors in salinity are expected to be on the order of 0.005.

The times given in the SeaBird headers do not include a time zone, so it was assumed that they are in PST. All times were converted to UTC by adding 8 hours.

There were errors in a few file names. There were often multiple casts on a single day. Data were converted by using the configuration file with the same name, if one was available. When there is no con file associated with a hex file, it required investigation to find out if there was good data in the file and if another con file would enable conversion.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

73 hex files and 68 configuration files were delivered. Three different CTDs were used.

A spreadsheet with positions was provided by Peter Chandler.

- Conversion of Raw Data

- The raw data were converted by matching the configuration file to the input file name. An initial conversion produced 68 CNV files. There were problems with the following files:

- N13020501_test.hex produced only surface data and will not be processed further.

- N13050702.hex did not have a matching con file; using con file N13050703 worked.

- Files n 13061201.hex and N13061201_as_received.hex were found. The former has a con file with the same name while the latter does not. There is a file N13061201b.con, so the “as received” hex file was converted successfully using that con file. The result looks identical to file N131061201.cnv except that it goes deeper, so file N13061201.cnv was deleted and N13061201_as_received.cnv was renamed as N13061201.cnv.

- N13071003 – no con file. This file looks like a copy of N13071002.hex except that 1002 is missing 2 characters in one line. N13071002 converted properly and the data look ok. There are only 2 position entries in spreadsheet. Couldn’t convert 1003 using N13071002.con, so there will be no file N13071003.cnv.

- N13082101b.hex  – looks like copy of N13082101.hex. There is no associated con file. Conversion using N13082101.con failed. Only 1 position entry in the spreadsheet for that day.

- N13112101b.hex – looks like copy of N13112101.hex but has different CTD serial #.  No con file with that name. Could not convert with N13112101.con.

- N13112101c.hex -  looks like copy of N13112101.hex but has different CTD serial #.  No con file with that name. Could not convert with N13112101.con.

So 4 files could not be converted, 1 more was converted but used to replace another previously converted, 1 was added and 1 that had been previously converted was found to contain no useful data. This leaves 68 CNV files to be considered.

Next casts were checked that had several hex files on a single day to ensure they are not copies. Problems found were:

- N13050701, N13050702, N13050703 occurred at different times. The first 2 were close together and the conductivity profile from the first looks bad, so the repeat a few minutes later was likely meant to replace it. So N13050701 should not be processed further.

- N13061201, N13061202, N13061203 occurred at different times, but the 2nd and 3rd were on August 21st, so these have been misnamed. Their profiles are similar and both are quite different from the first.  There is a file for that date but a little earlier in the day. Files N13061202 and N13061203 should be renamed N13082102.cnv & N13082103.cnv. N13061201.cnv is fine.

- N13071001.cnv, N13071002.cnv, N13071004.cnv – These all differ in time and the 2nd and 3rd are similar in profile. The first is shallower but similar in profile. All should be processed, but there are only 2 positions in the spreadsheet – use 1st position for 1st two since they are closer in time.

- N13082201.cnv, N13082202.cnv, N13082203.cnv, N13082204.cnv – all are at different times with the first 2 close together and the last 2 close together in time. The 1st and 4th used the same CTD and the 2nd and 3rd used another one. File N13082203.cnv has a profile that is very suspicious with a temperature excursion at depth. File N13082201.cnv looks odd. They did not use the same CTD. The descent rate does not show sudden stops or any explanation for these odd profiles. All 4 have positions shown in the spreadsheet. In T-S space all look stable, but the 3rd cast is notably different from the others. The 3rd cast is also quite far from the other 3, so that probably explains the differences. They will be processed.

- N13110702.cnv contains data from an upcast (but not a whole upcast) and N13110703 looks like N13110701 except that there is a stop at about 25m during the downcast. The profile plot looks ok. The times differ from the 1st file, so this one should be processed. But N13110702.cnv should not be processed.

Next, the spreadsheet with positions was checked against the converted files to see if any are missing:

- File n13061401- missing from s/s – name correct. Only a partial cast starting at 135m.  Do not process since don’t have position or know why it is missing data

- File n13070301- missing from s/s – shallow cast – somewhat out of line with casts before and after, and has unstable features in T-S. Given no position and no other positions that day, do not process further.

- File n13071004 -3 hex files at this site, 2 positions in s/s – likely correspond to n13071002 and n13071003 since n13071001 was shallow. Use 1st position for the 1st and 2nd casts.

- File n13100201 – very shallow, no entry in s/s and none for other casts that day. Do not process.

- File n13110702 – 3 casts very close together; 2nd only partial cast missing top and bottom. Do not process further.

- File n13110703 – 3rd cast bit noisy but maybe ok. Use same position as for 1st.

- File n1310417 – name is wrong – collected on April 17, should be 13041701. Fixed after conversion.

After corrections there were 64 entries in the spreadsheet including 2 casts that were not converted. There are 62 converted files.

The soak period was less than a minute for some casts. For most casts there was a soak of between 1 and 1.5 minutes and that was likely enough for most of them. A 2-minute soak is generally advised for the best salinity data if there is a large difference in temperature between air and water.

- The pressure sensors are only considered ±1db. The only information available suggests that the surface is well within 1db of 0.

- Temperature and conductivity generally look ok.

- The altimetry looks useful near the bottom of some casts where the CTD got within 20m of the bottom.

- WILDEDIT

No single-point spikes were noted, so this step was skipped.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 1s to increase the pressure resolution.

Plots of a few casts before and after this step show that reversals in pressure were removed without oversmoothing. The setting of 2s which is often used for SBE19 CTDs, looks too high for these data.

- ALIGNCTD

An initial run with no alignment applied to the temperature channel produced poor results with instabilities in T-S plots. Different settings were tried for the filter with no improvement made. Alignctd was applied using a variety of settings and the results were much better when temperature was advanced by 0.7s relative to pressure. The alignment can be fine-tuned later using SHIFT if further adjustment appears necessary.

STRIP was run to remove the salinity channel so that DERIVE will recalculate it without creating a second salinity channel.

- DERIVE

Program DERIVE was run to calculate salinity. A single CON file was used for the derivation since it has no effect on the operation.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

A spreadsheet of positions calculated from the *.act files was received from Peter Chandler in file NanooseCtdStations2013.xls. The spreadsheet did not contain the initial “n” found in the file names and was missing the final digit which was either a 1 or 2 since on some days there were 2 casts. Those items were added to the file names in the spreadsheet and the file was saved as a csv file.

The format of latitude and longitude needed to be changed to DD MM.MM N and DDD MM.MM W. Concatenate was used to do this. For example for column LOC:Latitude the entries are found using:  =CONCATENATE(INT(C2),"  ",TRUNC(((C2)-INT(C2))*60,2)," N ")

where C2 is latitude. There needs to be a space before the latitude and around the N and W. Do not add “! (deg min)” – that is done by the program.

This is a finicky job and usually takes many attempts before everything is right!

A column LOC:Station was added with every row having “Nanoose” entered.

A column LOC:Event Number was added with a 4-digit id based on month and day of cast. If there are more than 2 casts on a day they will have the same event number; the event number will not be used to rename the files in standard format because of that. (Use the RIGHT and LEFT functions to do this.)

The program MERGE CSV FILES TO HEADERS was then used to add that information to the headers.  A few errors were found and corrected in the CSV file.

8. Checking Headers

Header Check was run and a few errors were found in headers. The conversion to IOS Header format plus MERGE CSV Files to Headers routines were repeated until all problems were resolved.

The altimetry headers were exported to a spreadsheet file. This shows that only 4 casts got within 15m of the bottom. Plots of altimetry were checked to see if this result looks reasonable and it does. Most casts got to within 20 to 25m of the bottom.

Surface check was run and the average surface pressure was 0.6db and all readings are associated with low but definitely ‘in-water” salinity values.

9. SHIFT

These data were aligned earlier in the process, but tests were run to see if a further shift might help. T-S plots show a few unstable features though most plots look excellent. Some shifting of conductivity (followed by salinity recalculation) improved a few features but made others worse. So this step was skipped.

11. CLIP

CLIP was used to remove the first 60 records since they will all be within the soak period and this will eliminate some spikes and simplify the editing step.

12. DELETE

Then DELETE was run on the SHF files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings

12. CTDEDIT

The DEL files were copied to EDT.

All files were opened in a graphical editor, CTDEDIT. Most casts required light editing. Surface records were removed from many files and bottom records were removed from a few casts where there was clearly some corruption. Salinity needed editing in many files to smooth unstable features that appear to be due to minor misalignments of T and C in the presence of temperature gradients. A few records were removed because they were corrupted by shed wakes, but the descent rate was usually steady enough that this was not a big problem.

The edited files were copied to EDT.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure      	Averaging interval = 1.000

Minimum bin value = .000	Average value will be used.

Interpolated values are NOT used for empty bins

Profile plots were examined on-screen and no problems were found.

Page plots were examined and look ok with only a few very slightly unstable features.

14. Intercomparisons

Historic ranges – The only excursion from the local climatology was near-surface salinity that was a little low for one cast in May. This is not unusual so close to shore.

Repeat casts – There were some casts within a few hours of each other, but not at exactly the same place. The temperature and salinity are reasonably close, but the data are not suitable for assessment of CTD performance. Nonetheless, at depth there is often good correspondence. For example there were 3 casts on July 10 that were close together, with the 2nd about 2 hours after the 1st and the 3rd about 9 hours after the second. On a T-S surface at about 175m the temperatures along σt-lines varied by ~0.01C° and the salinity by ~0.003 over 11 hours.

15. CALIBRATION

No recalibration was applied.

16. REMOVE and HEADEDIT

The time zone for the times in the headers is unknown. The times suggest either PST or PDT, since they are mostly between 8 and 10am, so it will be assumed it is PST for all. We often find that computer times are not updated to mark the change from PST to PDT and vice versa. Program ADD TIME CHANNEL was used to add 8 hours to get UTC times. These times may well have an error of 1 hour during the time when PDT is in use locally.

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate Altimetry and Flag.

The HEADEDIT routine was used to fix formats and to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients,

but on the order of 0.005 where gradients are low and descent rate steady.

There were no salinity calibration data available for this cruise.

The standards check routine was run and no problems were found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced. No errors were found.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	Nanoose Test Range

Dates: 	Start:	10 January 2013	    End:	13 December 2013

Location: Nanoose Test Range 	Vessel:    Stikine, Sikani and possibly others

NOTE: Several SEACATs were used for this cruise and not all were identified. The calibration information used in conversion can be found in the headers of the CTD files.