## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 25 Mar 2025 | ### Updated channel names & formats in TOB & Loop files. GG&SH |
| ### 23 July 2024 | ### Added Cesium data to 7 casts. SH |
| ### 1 Dec 2021 | ### Corrected Salinity:Bottle and NUTS precision lost during the addition of HPLC, accounted for corrections below at the same time. S.H. |
| ### 17 Aug 2021 | ### Corrected HPLC flag comments for events 7, 22, 70, 79. Removed HPLC data from events 14 and 41. S.H. |
| ### 6 April 2021 | ### Corrected QF DMSP sample 400 for error in flags & comments. S.H. |
| ### 10 Dec 2020 | ### Corrected events for HPLC samples. S.H. |
| ### 4 August 2020 | ### Added HPLC Data. S.H. |
| 2 May 2016 | Added comments re DIC/ALK/pH; replaced blank flags with zeros. Removed DIC/ALK/pH from event 74. G.G. |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2014-01

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 10 February 2014 – 24 February 2014

Processed by: Germaine Gatien

Date of Processing: 27 May 2014 – 22 July 2014

Number of original HEX files:	42	Number of CTD files: 42

Number of bottle files: 		41	Number of bottle casts processed: 41

Number of original TSG files:  	2	Number of processed TSG files:  2

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1176) on the primary pump, a SeaPoint Fluorometer (#2356) with a 30X cable on the secondary pump, a WetLabs ECO-AFL/FL fluorometer (#2215), a Biospherical QSP-200 PAR sensor (#4615) and an altimeter (#1204).

A thermosalinograph (SeaBird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor and a flow meter.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette sampling log sheets were generally in good order, especially given the terrible weather conditions.

The descent rate of the CTD was extraordinarily noisy. During one cast the CTD rose by more than 20db in one minute with many smaller reversals within that excursion. A graphical editor was used to remove obviously bad data corrupted by shed wakes, but when there are complete reversals of the CTD it is harder to distinguish good from bad, so the quality of the data is somewhat lower than usual.  From event #67 onwards conditions, and quality of data, improved.

There were 2 fluorometers mounted on the CTD. The SeaPoint fluorometer malfunctioned, so only the WetLabs ECO fluorescence will be placed in the OSD Data Archive.

The comparison of CTD salinity with bottle samples indicated that the primary salinity was low by ~0.004 and the secondary salinity low by ~0.001. The analysis was done within 6 weeks of collection, so evaporation should not be a major problem, but there may be a slight effect making the CTD salinity look lower than it really is. There were a number of cases of missing or improperly inserted bottle liners and bottles with no head space. The secondary salinity was generally less spiky than the primary so the secondary channels were selected for archiving except for 2 casts (events 41 and 43) for which the secondary salinity data were very poor. So the primary salinity was recalibrated to match the secondary and that required a pressure-dependent correction since the primary had more depth-dependence than the secondary.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8mL/L from 0 to 100db

±0.6mL/L from 100db to 300db

±0.2mL/L from 300db to 800db

±0.05mL/L below 800db

The TSG salinity was extremely noisy presumably due to air bubbles with many values likely too low. Editing is not practical. No data were acquired from the intake thermometer for the loop. Based on a comparison with co-incident 5m CTD data, the lab temperature was recalibrated to create a proxy for intake temperature. The fluorometer used for the TSG was last recalibrated in 2003. Both TSG salinity and fluorescence were reported with fewer significant figures than usual. On a positive note, except for a short initial section, the flow rate was very steady throughout the cruise.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

An initial cast was sent from sea for initial processing. The configuration file looked fine. All sensors were producing sensible data. The secondary salinity seemed noisy at times, but a minor adjustment to the secondary conductivity alignment might fix that. The two fluorometers looked very similar on the upcast. The SeaPoint was higher on the downcast than the ECO, but both had the same general shape.

##### Preliminary Steps

The Log Book and rosette log sheets as well as various analysis logs were obtained.

There were few notes about problems. There are occasional notes about the PAR sensor being ON or OFF, but this is not entered consistently, so that will have to be figured out later by plotting profiles.

The post-cruise report indicates that there were problems with the TSG with no external temperature data and poor salinity data due to air in the system.

There was only a single version of the configuration file.

Extracted chlorophyll, DMS, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. The only error found was that the SeaPoint fluorometer serial number was missing and the calibration date of the ECO fluorometer was missing. The former was added to the file, but the latter was not known.

The PAR sensor was not always mounted. Cast lists of PAR and NO PAR will be prepared later when profiles are plotted.

##### BOTTLE FILE PREPARATION

The ROS files were converted using file 2014-01-ctd.xmlcon with Tau and Hysteresis corrections selected. Those files were put through CLEAN to add event numbers (*.BOT).

Temperature and salinity were plotted for all BOT files. The primary salinity had a few spikes not seen in other channels; CTDEDIT was used to remove a few values from that channel in events #21 (surface) and 38(25 and 300db).  The output files were copied to BOT.

Header Check was run on the BOT files and no problems were found.

The BOT files were then averaged to enable an ADDSAMP file to be prepared. First the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

Sort was used to ensure that the ADDSAMP file was in sample number order.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was ordered on sample number and converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2014-01-bot-hdr.txt.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-01.xls. A simplified version of the spreadsheet was saved as 2014-01chl.csv which was then converted to individual CHL files. Loop data were saved as file 2014-01chl-loop.csv.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-01oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2014-01oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2014-01SAL.xls. The file was simplified, loop samples were removed and the file was saved as 2014-01sal.csv. That file was converted to individual SAL files. The salinity data were analyzed 33 to 45 days after collection. Loop samples were saved to 2014-01sal-loop.csv. The event # 78 was changed to #79; the computer crashed on the first attempt at this cast and was restarted as #79.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-01nuts.xls which included a report on precision.  The file was simplified, reordered on sample numbers and loop samples were removed and the file was saved as 2014-01nuts.csv. The file was converted to individual NUT files. The loop data were saved as 2014-01nuts-loop.csv. Event #78 was changed to #79; the computer crashed on the first attempt at this cast and was restarted as #79. A CTD file will be prepared for #78 but no bottles were fired.

DMS

DMS data were obtained in file DMS summary (2014-01).xls. (There was a separate report on analysis techniques and problems.) The file was simplified and saved as 2014-01DMS.csv and converted to individual DMS files.

The SAL, CHL, OXY, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The MRG files were exported to a spreadsheet and compared to the rosette sheets. A few inconsistencies were noted:

- Cast #3 – there is no salinity sampling noted on the rosette sheet but there is a sample #42. The sample is higher than the CTD salinity but not unbelievable.

- Cast #4 – sample #54 is missing from salinity spreadsheet. Sample #42 does not match the salinity for this cast as well as it does cast #3, so this does not seem like a mislabelling problem.

- Cast #7 – sample #58 is missing. Not close to sample #42.

- Cast #9 – two DMS values had been entered wrong in processing files. The < entered by the analyst should have been changed to 0, but was entered as -99; that was corrected.

- Cast #22 - Sample given as #256 (event #22), should have been #258 (Event #25). There is no 256 on the salinity analysis sheets but there is a 258 and that was missing from the spreadsheet. The value 34.5151 makes no sense as #256 (surface) but does make sense as 258 (1500db). This just looks like a simple typo.

- Event #25 – Niskin 2 at 1500db – sample missing. This looks like the extra one from event #22. CTD salinity = 34.5124. Changed sample #256 to #258 in processing files and notified analyst about error in spreadsheet.

- Event #35 shows 9 salinity samples but there are none in the file because there were no liners they were not analyzed. There should be pad values and flags with comment to indicate this. This was changed in processing files and analyst notified that spreadsheet needs amendment.

- Event #43 – oxy sample 490 was lost but not flagged. This was done in processing files and analyst notified that spreadsheet needs amendment.

- Event #56, 59-70 and 79 – CHL appeared to be missing, but the data were found in the spreadsheet with the wrong event numbers. These were corrected in the working files, and the analysts were notified that corrections were needed to the QF spreadsheet.

- Sample #s 637 and 638 were used for both Event #s 74 and 79. There was no record of sampling at 74 on the sampling log, so it is assumed this will not be a problem as the cast was run for someone from another institution. But in case of confusion the numbers were changed to 9637 and 9638 and a note was entered in the header.

Cast #31 has many deep bottles that were taken just to provide water – those lines were removed from the IOS file and the process of adding sample numbers and merging with bottle data was repeated.

An initial header check did not turn up any problems.

##### COMPARE

Hysteresis Check

Before doing the full COMPARE runs, tests were run to check whether the hysteresis parameters for the dissolved oxygen sensor needed fine-tuning. COMPARE was run on dissolved oxygen with pressure as the reference channel. An initial examination suggested that there might be a problem with hysteresis. These checks had not been run since the sensor was last recalibrated at the factory. To fine-tune the hysteresis correction the manufacturer recommends varying parameters E, H1 and H3. Generally, varying E has been found to be most effective. Casts #25, 34, 35, 43 and 75 were selected for tests since they were deep casts. Those casts were converted using E values between 3.4 and 3.8 and COMPARE was run with the resulting files. E=3.75 proved to be the best choice.

The data were reconverted using an updated version of configuration file 2014-01-ctd.xmlcon, with E=7.5 for the DO calibration. Then the steps described in the previous section were repeated.

Dissolved Oxygen

COMPARE was run using the newly converted files.

There are 2 clear outliers and 2 others slightly out of line:

- Event #1 – Niskin #1 at 219db. The bottle value (2.25mL/L) is lower than the lowest SBE DO value (4.16mL/L) by almost 2mL/L, and is way out of line in COMPARE. This is a cast in Haro Strait where gradients are low. After recalibration of the SBE DO sensor data the bottle will be even more out of line.  This sample value was replaced with a pad value and flag 5.

- Event #18 – Niskin #17 at 50db – The SBE DO (4.44mL/L) is lower than the bottle (5.045mL/L) by 0.6mL/L. The standard deviation in the CTD data is high, and the local gradient is high so that the distance between the CTD and Niskin could account for the difference in light of the 3db range of pressures during the stop. This is a clear outlier in COMPARE but not clearly a bottle problem. No flag was added.

- Event #34 – Niskin #22 at 26.9db – This is a minor outlier and the standard deviation in the CTD data is slightly high, but it is probably not enough to explain the difference. Given this is a well-mixed surface area, the bottle value looks out of line in profile, being lower than those above and below it by ~0.07mL/L. A flag 3 was added since it is not seriously out of line and could be ok.

- Event #43 – Niskin #22 at 26.0db – The CTD data looks ok. This sample comes from well-mixed surface water but the bottle value is out of line with those above and below it by ~0.08mL/L. The value looks low but that is not certain. A flag 3 was added.

Checks were made of other variables for Niskin #22. DO looks ok for casts #35 and #75 for Niskin #22. The one available salinity sample from that Niskin looks a little closer to the CTD value than other bottles from that cast, but the difference is likely not significant for near-surface sampling. Nutrients do not look out of line for that bottle. There is no note of problems peculiar to DO such as a spigot leak. The draw temperatures look ok for those two bottles, though this is a very rough estimate. So if there was a problem with Niskin #22 it was minor.

When a few outliers were excluded based on residuals, fits were found with and without forcing the offset to be zero and with and without including cases where the SBE DO values were <1mL/L since SeaBird have raised some concern about the accuracy of analysis at low values.

Forcing a fit through the origin produced these fits:

DOX_BOT = 1.0541 * DOX_CTD  R2 =0.91 including bottles for which SBE DO<1mL/L

DOX_BOT = 1.0542 * DOX_CTD  R2 =0.81 excluding bottles for which SBE DO <1mL/L

Allowing a free offset produced these fits:

DOX_BOT = 1.0480 * DOX_CTD +0.0287 with R2 =0.93 including bottles for which SBE DO<1mL/L

DOX_BOT = 1.0443 * DOX_CTD +0.0474 with R2 =0.89 excluding bottles for which SBE DO <1mL/L

As expected there is very little difference between the two fits if the offset is forced to be 0.

The best fit judging by eye and by R2 value is with a free offset and including the low value DO bottles. There are a lot of bottles with low values which may help avoid any problems due to comparing small numbers.

None of these fits seems really convincing, nor can we really expect a great fit given all the influences on both titrations and SBE sensor performance. The free-offset, full-range fit was best and when a few more outliers were removed based on residuals, the fit was

DOX_BOT = 1.0497 * DOX_CTD +0.0234 with R2 =0.97

Another approach was taken by using just the bottles at the bottom and at the top where the gradients were mostly low. The offset found in that fit was then forced in a fit with all data included. A few major outliers were removed and then points determined as outliers based on residuals. The fit found using that approach was:

DOX_BOT = 1.0502* DOX_CTD +0.0142 with R2 =0.97 **

The fit excludes many points in the range 2mL/L to 4.5mL/L which is where high gradients would tend to make for a poor fit. Almost all the outliers in that range have the DO value being lower than bottles by more than points that are included in the fit, which would happen if they are in higher gradient zones. On the other hand many of the outliers in the 4.5-5.2mL/L zone are in the opposite direction. Those come mostly from the Haro Strait cast which was well-mixed throughout, so the effect of either poor response or distance between bottle and CTD is not significant.

The two approaches produce similar results with a slope near 5% and small offsets. Which fit is best is far from obvious and the difference between them is only 0.006mL/L near the top of the range.  Fit ** will be selected for recalibration. There will be another check later to see how downcast data compare with the bottles so there will be another opportunity to see if the fit looks reasonable.

For more details see 2014-01-dox-comp1.xls.

Salinity

Compare was run with pressure as reference channel. Salinity samples were analyzed within 33 to 45 days of collection.

There are comments about a group of samples having no headspace and compromised liners. They should be flagged. Flag 3 was added to samples 258-275 and the comment was amended to indicate they were not significant outliers in COMPARE.

There were a number of clear outliers, but most were from near the surface where high gradients and/or incomplete flushing of Niskin bottles may be significant.

Outliers requiring flag or correction:

- Cast #3 at 1.8db is the largest of the outliers but the standard deviation in the CTD salinity is high, and this is very close to the surface where slight inefficiency in flushing of the Niskin bottle near a significant gradient might explain the difference. The CTD is lower than the bottles by ~0.07.  No flag is justified.

- Cast #34 -Sample #354 ~1250db – CTD lower than bottle by ~0.01 which is fairly high for this depth. There were 3 readings taken. The standard deviation of the CTD salinity is very low at firing time and the descent rate during the stop would suggest flushing should be quite good. Flag 3 as an outlier in compare.

- Cast #65 at 5m – CTD lower than bottle by 0.033. The water is well mixed at this level and the bottle value is higher than CTD salinity at 5db and 50db. Flag 3.

- Cast #70 at 5m – CTD lower than bottle by ~0.03. CTD does not reach those values until below 80db. Flag 3

- Cast #75 -Sample #625 – 300db – CTD lower than bottle by 0.034. CTD data is very noisy before firing so likely bottle value is ok, but contents differ from water sensed by CTD. No flag added.

- Cast #75 -Sample #632 – 75db - CTD lower than bottle by 0.012. CTD data was very noisy before firing and Niskin likely contains water from shed wake. Likely no problem with analysis, jut a mismatch with CTD.  No flag added.

When outliers are excluded based on pressure <300db and differences >0.008, the two CTD channels are found to be lower than bottles by an average of 0.0037 for the primary and 0.0010 for the secondary. Both had a standard deviation of 0.0016. Given that either evaporation of samples or poor flushing could lead to CTD values looking lower, these are good results. There is not likely to be a major problem with flushing given that bottles bobbed around quite a lot in heavy seas. Evaporation should not be a big problem, given the quick analysis, but even a month might allow a little evaporation especially as there were problems with liners for some of the samples. There is some suggestion of poor Niskin flushing at the surface, with CTD being lower than bottles by about 0.002 more at low salinity than at high salinity.

There is no indication of salinity dependence. The primary has more pressure dependence than the secondary salinity. When the differences between the 2 salinity channels were plotted against pressure and bottles from above 100db were excluded, the fit was:

Sal1 – Sal0 = 6 * 10E-7 * Pressure + 0.0015

So at 100db the average difference was 0.0015 and at 4000db it was 0.0039. So while the average difference was 0.0027, that difference is only found around 2000db.

For full details for the COMPARE run see file 2014-01-sal-comp1.xls.

Fluorescence

Two fluorometers were mounted on the CTD.

COMPARE was run to see how each of the fluorescence channels compared with extracted chlorophyll and with each other. A table was prepared with data from both fluorometers and extracted chlorophyll samples and pressure and event number.

Extracted chlorophyll values were fairly low with a maximum of 2.2ug/L.

The SeaPoint fluorescence is higher than the ECO by a median value of 1.16 (standard deviation 0.23).

However, for high CHL, the SeaPoint is lower than ECO with the minimum ratio of SP/ECO being ~0.8.

Plots of the ratio of each fluorescence channel to CHL versus CHL look similar and typical of these fluorometers. The ratios are very high at very low CHL values but quickly fall to an average of ~2 when CHL reached 0.1ug/L. They continue to fall and are ~1 between 1 and 2ug/L., dropping to a minimum of ~0.66 when CHL=2.2ug/L.

There is a small group of points for which SP fluorescence is much lower than the ECO when CHL ~0.5. All of these points occur during casts 56, 65, 67 and 79. For those casts the two fluorescence channels are close at the surface, but the SP fluorescence is much lower than the ECO at 50m. The SP is closer to the extracted CHL at 50m. Both fluorometers are much higher than CHL at the surface. There is little difference between the CHL at the surface and at depth.

Full profile plots show some odd features in the SeaPoint fluorescence with sudden shifts in values especially in the deeper parts of downcasts early in the cruise and closer to the surface in upcasts late in the cruise, so this confirms that the SeaPoint is not as reliable as the ECO fluorescence for this cruise, at least in resolving the shape of the fluorescence profiles.

In plots of fluorescence versus extracted CHL it is clear that the data fall into 2 groups for both fluorometers. A plot of the ratio of FL/CHL versus Event # shows that the ratio is higher offshore. That pattern is seen in other cruises; possibly the type or health of phytoplankton offshore are different from those inshore.

For more detail see file 2014-01-fl-chl-comp.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were detected.

##### Conversion of Full Files from Raw Data

All files were converted using con file 2014-01-ctd.xmlcon with hysteresis and Tau corrections chosen.

A few casts were examined and all expected channels are present.

The two temperature and conductivity channels are fairly close to each other during the downcasts but noisy in the upcasts and there are many signs of shed wake corruption. In the downcasts the secondary temperature is smoother but in upcasts there are odd excursions in both channels.

The two fluorometer traces both look reasonable in the well-mixed surface layer (50 to 75m deep) in the early casts examined. But below the surface the SeaPoint fluorescence has odd shifts, particularly in the downcasts. Late in the cruise the downcast looks better while the upcast near-surface SeaPoint data are spiky and/or much lower than expected.

The altimetry, transmissivity, dissolved oxygen and PAR data look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on 4 casts with only a few stops during the upcast, but the results are not easy to interpret because of noisy upcast temperature data and much shed wake corruption of all channels. However, the +4s setting looks reasonable and is certainly an improvement over no adjustment.

ALIGNCTD was used to advance the DO Voltage by 4s relative to the pressure.

##### CELLTM

The upcast data are noisy so the usual tests for CELLTM settings were also hard to interpret. Tests were run on a few casts using a variety of settings. All settings tested improved the data with little difference among them. The choice of (α = 0.03, β=9) is probably best for the primary and (α = 0.03, β=7) best for the secondary.

CELLTM was run using (α = 0.03, β=9) for the primary conductivity and (α = 0.03, β=7) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are always very noisy, but these data are so noisy that no reasonable comparison could be made except to say there was little variation in the conductivity differences. Salinity differences were ~0.003 for event #35 and <0.001 for event #43.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are some bad values in pressure in cast, #31, so some values were replaced with pad value in the IOS file. CLEAN was rerun. This was repeated until the CLN file had no pressures <-1db. There may still be a problem with this cast but this can be addressed at the DELETE stage.

The surface check gave an average value of 2.45db with salinity values that make sense. There was one cast with a surface pressure of -0.31db and in-water salinity values, but there are just a small group of such values so this may be a case of the CTD being very briefly out of water. Or there may be a pressure spike. Pressure may possibly be slightly low but there is insufficient evidence to justify recalibration.

The cross-reference check was compared with the log book and two station names were wrong (events 79 and 82). They were corrected in the headers.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings, maximum pressure and water depth were exported from the headers of the CLN and MRG files to spreadsheets. There were some discrepancies in water depths between log entries and header entries in both sets. Where the differences were only a few metres they were left unchanged, but larger differences were investigated and most were changed. In the MRG headers there were some errors in the altimetry entry where spikes were misinterpreted by the algorithm and the deepest bottle was not close to the seabed. Such headers were removed.

##### 13. Shift

Fluorescence

The usual tests for alignment of the fluorescence channels are not likely to be useful for these data because there were many stops for bottles, the descent and ascent rates of the CTD were extraordinarily noisy and the temperature traces used as a guide for aligning fluorescence were extremely noisy.

Normally a shift of +24 records is used to align SeaPoint fluorescence with temperature. Tests are pointless for this sensor because it is known that there were odd shifts in values so that differences between upcast and downcast may be random. So the usual parameter was applied.

SHIFT was run on all casts to advance the SeaPoint fluorescence by +24 records.

For the ECO fluorometer a setting between 0 and 12 records is usually applied. As for the SeaPoint sensor, tests are not very useful though at least the upcast and downcast fluorescence values are similar. It is clear that some shift is required and +12 records looks reasonable.

SHIFT was run on all casts to advance the ECO fluorometer by +12 records.

Conductivity

Tests were run on a few casts using a variety of settings to see which shift produces reasonably stable T-S plots. Because the descent rates were extremely noisy it is difficult to judge what is noise that alignment SHOULD remove and what is not. A few casts with fewer shed wakes were examined (events #70, 75 and 82) and judging from quiet sections it appears that -0.7 records is best overall for the primary conductivity and +0.7records for the secondary conductivity. These settings do not look best for all features as variable descent rate and shed wake corruption complicate the picture.

Two runs of SHIFT were used to apply advance the primary by -0.7 records and the secondary by +0.7 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The alignment looks as good as we are likely to achieve.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds (taken from header)

COMMENTS ON WARNINGS: There were warnings for 3 casts. For event #22 there was noisy pressure around 250db but DELETE appears to have worked properly so that there are no large gaps in pressure. For event #34 all warnings were for surface values at the end of the cast, so of no significance for the downcast. Event #31 is the most problematic as the pressure was extremely spiky in some places and in others there were large excursions in the CTD that definitely look real. Some work was done earlier on this cast to remove some obviously bad values with pad values, but that just addresses the spiky areas. At about 160db the CTD rose by more than 20m with smaller reversals within that large excursion. The following illustration is a record of pressure and temperature versus scan number using data before any processing steps were applied. It shows the problem faced in selecting the best data for the profile. Not surprisingly the output from DELETE included various bits stitched together around 160db including data from areas where heavy corruption of data is anticipated.

To avoid this problem, the SHFc1 file was edited to remove scans 6157 to 9439. Doing that produces a file with data down to 150.9db before the reversal and starts again at 150.96db after the reversal. The match of temperature and salinity is far from perfect. It looks like the best that can be achieved at that depth. Other CTD reversals during this cast were smaller, none appear to be greater than 10db and DELETE will likely produce reasonable results for those features.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The conductivity and temperature sensors were both recalibrated in late January 2013 and this is the first known use since that time.

2. Dissolved Oxygen

The DO sensor was recalibrated in February 2013and this is the first known use since that time.

3. Pressure

The pressure sensor was recalibrated in March 2013 and this is the first known use since then.

Historic ranges – The DEL files were bin-averaged and profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The temperature data were all within the climatology except for some slightly low temperatures near the surface in Haro Strait and Juan de Fuca Strait. Salinity values were slightly lower than the climatology minimum in the halocline between P12 and P17 (near 150db) and at P4 the salinity is high around 50 to 75db, also in the halocline. None of these excursions suggest a systemic error, so are not evidence of calibration errors.

Repeat Casts –

There were some repeat casts. At P26 around 1000db along constant σt-lines temperatures are within 0.0015C° and salinity within 0.0005 units. The two casts occurred about 8 hours apart. This is good repeatability.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

The first issue is to decide which sensor pair to edit. There were severe problems with the secondary temperature during casts #41 and 43. Trial editing removed more data when the primary was selected. Overall, the secondary T and S seem like a better choice.

The secondary channels were selected for archiving except for casts #41 and 43 for which the secondary salinity looked bad and the primary was much better.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. The descent rate was extraordinarily noisy and much more data were removed than is usual.

As noted on other recent Tully cruises, there were some odd excursions in salinity that do not appear to be related to shed wake corruption or stray spikes. These unstable features look like they are caused by misalignment of temperature and conductivity, possibly related to high deceleration rates of the CTD. Where it is clear that the temperature points are ok, only salinity points were removed. Such salinity excursions were seen less often than during other recent offshore cruises, probably because they were masked by the heavy shed wake corruption and resulting complex traces.

All ED1 files were copied to EDT.

##### Initial Recalibration

The pressure does not need recalibration.

Since the secondary salinity will be used for most casts, but not all, the primary salinity needs to be recalibrated to match the secondary. It was established in section 4 that the difference between the salinity channels is pressure dependent. So the following recalibration will be applied.

Sal0 Corrected = Sal0 + 6 * 10E-7 * Pressure + 0.0015

Tests on a few casts show that the two salinity channels are much closer after this correction.

Based on the bottle comparison described in section 4 the dissolved oxygen channel was recalibrated using the fit:

DOX_CTD Corrected = 1.0502* DOX_CTD +0.0142

File 2014-01-recal1.ccf was prepared to apply the corrections to dissolved oxygen and primary salinity.

CALIBRATE was run on the MRGCLN2, SAM and EDT files using that file.

COMPARE was rerun to check that the DO correction was done correctly and it was. The average difference after removing the same outliers as in the fit upon which the recalibration was based, shows the CTD DO to be low by an average of 0.003, though the standard deviation is 0.02mL/L. The fit looks best at higher CTD DO values, with slightly low values at the low end of the DO range. However, the scatter suggests this is as good a fit as we can achieve. (See file 2014-01-dox-comp2.xls.)

##### Final Calibration of DO

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. When the differences were plotted against DO concentration, the fit had a lot of scatter but looked reasonable. When a few outliers were removed the SBE DO was high by an average of 0.0082mL/L and a standard deviation of 0.07mL/L.  SBE values were a little high at the high end of the DO range and a little low at the low end. Given the scatter no further recalibration is justified.

A plot of differences against pressure was used to make a final study of outliers and to make a rough estimate of the accuracy of downcast CTD dissolved oxygen.

##### Special Fluorometer Processing

It was noted during the bottle comparison that the SeaPoint fluorometer behaved oddly during upcasts starting at event #56 with one bad value at the previous cast. So at this point a study was made of the downcast data.

For the most part the two fluorometers have the same shape and values are similar. Generally the SeaPoint reads higher than the ECO except at the one cast where extracted CHL values are >2ug/L. However, the SeaPoint fluorescence channel has very odd variations at depth, below 300db. These don’t really matter since there is no meaningful signal at those depths, but it may indicate a problem with the sensor. For one cast only, there were odd excursions near the surface. The maximum extracted chlorophyll for this cruise was 2.1ug/L, so it is not surprising that the two fluorometers are reading much higher than the CHL most of the time.

Examining the full upcasts it is clear that there were problems from the beginning of the cruise with SeaPoint values flipping between reasonable ones and very low ones. The bottle stops appear to be unaffected until late in the cruise.

Given the problems seen late in the cruise with the SeaPoint fluorescence during stops, it is wisest to archive the ECO fluorescence and remove the SeaPoint fluorescence. Dr. Peña was consulted and agreed.

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The files were bin-averaged (0.25db bins) and put through HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved.

A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to reduce spikiness in both fluorescence channels. A few casts were examined before and after this step and showed that the filter was effective on the SeaPoint fluorescence but had little effect on the Wetlabs fluorescence.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was considered appropriate.

##### Final CTD File Steps (REMOVE and HEADEDIT)

Cast lists of events with and without PAR sensor were prepared.

REMOVE was run on all casts with a PAR sensor mounted except #41 and 43 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Fluorescence:URU:Seapoint, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted except for #41 and #43 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Fluorescence:URU:Seapoint, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

REMOVE was run on cast #41 and 43 (which had no PAR sensor mounted) to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

Profile and T-S plots were produced at this point to check for errors. No problems were found.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8mL/L from 0 to 100db

±0.6mL/L from 100db to 300db

±0.2mL/L from 300db to 800db

±0.05mL/L below 800db

The primary CTD salinity channel was used for events #41 and 43, so it was

recalibrated to match the secondary channel which was selected for

all other casts.

For details on the processing see the report: 2014-01_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were low in Haro and Juan de Fuca Straits at about 75-80%. For all other casts they were between 98% and 101%, as expected offshore at this time of year.

##### 24. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts with a PAR sensor mounted except #41 and 43 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Fluorescence:URU:Seapoint, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted except for #41 and #43 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Fluorescence:URU:Seapoint, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

REMOVE was run on cast #41 and 43 (which had no PAR sensor mounted) to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. A few problems were found in salinity and DMS sample data:

- Salinity sample 54 should have been assigned to event #4, not event #3.

- Salinity sample 35 is missing so was flagged 1 with a comment that it is lost or not taken.

- DMS samples 83, 407, 527 were the average of duplicates but were missing flag 6

- DMS sample 527 had the sum of the duplicates entered instead of the average.

The errors were corrected and the bottle merging process repeated.

A header check was run and no errors were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found.

##### 25. Thermosalinograph Data

Notes from the cruise report:

The thermosalinograph was still not working properly. The remote temperature was reading -9.999°C during the whole cruise, and the salinity data quality was very poor, maybe because of all the air in the filters, going into the main part of the TSG.

There were loop salinity nutrients, extracted chlorophyll and salinity samples taken, some while stopped and some while underway. There is some doubt about the source of one salinity sample (58) and one CHL sample (40), which should be investigated later.

a.) Checking calibrations

The configuration files for the 2 casts are identical. There is some confusion about what fluorometer was used as the log book lists 2847. That makes no sense, so the configuration file is assumed to be correct. The log entry was probably meant to be 2487 which is the serial number of the TSG not the fluorometer. One file was renamed as 2014-01-tsg.xmlcon. Note that the fluorometer calibration parameters are from 2003.

b.) Conversion of Files

The 2 files were converted to CNV files using configuration file 2014-01-tsg.con. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. As noted in the cruise report the intake temperature was off-scale throughout the TSG record. The first part of the first file is bad; while the flow rate is ~0.5 the salinity is ~0 and primary temperature is high, close to the expected lab temperature. Those records should be removed later. The flow rate after that section is unusually steady at about 1.05. The temperature data are quite smooth. The fluorescence is fairly high ranging from about 2 to 4ug/L. The salinity is very noisy with what look like one-sided spikes of size 0.5 to 1 salinity units; filtering will not be useful. The second file is much the same though fluorescence starts somewhat lower but increases towards the end.

The track plot looks a little odd but this is likely because the ship took some odd jogs to avoid storms. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2014-01-ctd-tsg-comp.xls. There were 39 casts which overlapped with TSG files but for the first few hours the flow was very low and salinity was very close to 0. Those casts were removed from the spreadsheet.

The 2 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all≤0.0003°. This shows both the times and positions are reliable for both systems.

The loop samples were combined in file 2014-01-loop-tsg-rosette-comp.xlsx. There were a few problems with the labelling and event numbers. The CHL sample said to be from P21 and event 40 is believed to be from Loop 60, Event 60. Samples said to be from events 57 & 58 are really from event 56 & 59. Times were added and then any overlapping TSG data was added. There were 8 such records but 1 was from the beginning of the file when there was a large section of near-zero salinity. Then 5m rosette bottle data were added from 3 casts though only 2 had salinity and CHL samples.

The two spreadsheets will be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature was not working.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

1. Intake Temperature Unavailable.

2. LAB TEMP The lab temperature was higher than the CTD temperature by from 0.15 C° to 0.30 C° and a median of 0.23C°.  As expected the difference was temperature dependent with the greatest warming for the offshore casts where the in situ temperature was lower. A correction will be applied to the lab temperature as a proxy for the missing intake temperature, as follows:

Temperature Primary = Lab Temp * 1.0581 – 0.7049

This produced corrections of from 0.23 C° to 0.27 C°.

3. SALINITY  The TSG salinity data are lower than the CTD salinity by from 0.03 to 0.93 and a median of 0.07. When obvious outliers are excluded, the median is 0.05. The differences appear random with no temperature or time dependence.  The salinity is reading low but that is likely due to causes other than calibration drift, air bubbles being the most likely problem.

4. FLUORESCENCE The TSG fluorescence was 2.1 times higher, on average, than the ECO fluorescence from the CTD with a range of 1.8 to 2.9.

(See 2014-01-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons

The TSG salinity was found to be lower than the loop samples by from 0.02 to 0.09.

The TSG fluorescence is higher than the loop extracted chlorophyll by a median factor of about 3 but the range is large, 1.2 to 6.7.  The ratio appears to be falling as CHL increases so that the two values are fairly close for CHL=2.7ug/L.

(See 2014-015-loop-tsg-rosette-comp.xlsx.)

- Surface rosette vs TSG (salinity and chlorophyll)

The rosette CHL sample and loop CHL sample from the only 2 overlapping case have very similar values, 0.37/0.36 and 0.55/0.50ug/L.

The rosette salinity was lower than the loop salinity by 0.0001 and 0.0025, which is good correspondence given the times and depth, are not well matched.

- Calibration History

The TSG pressure, TSG temperature and conductivity sensors were recalibrated in February 2013 and this was the first known use since then.

Conclusions

1. The TSG clock appears to have worked well.

2. The salinity was extremely noisy, likely due to bubbles.

3. The temperature in the loop increases by from ~0.15 to ~0.30Cº based on comparisons during stops for CTD casts, and was temperature-dependent. A T-dependent fit was found.

4. The TSG intake temperature was either missing or malfunctioning. There was no signal.

5. The TSG Salinity is lower than that of the CTD by a median of 0.07, and the TSG Salinity is lower than the loop salinity by 0.03 to 009. Recalibration is not appropriate since the differences are likely mostly due to bubbles in TSG salinity. Fewer significant figures than standard should be used to draw attention to the problem.

6. The TSG fluorometer was a little higher than the CTD Wetlabs fluorometer in one case and a little lower in another. It is 1 to 7 times higher than the Loop CHL. Fluorescence is always considered nominal, and we don’t have much experience with this sensor, so the number of significant figures will be reduced to draw attention to this.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT. The first 933 records were removed from the first file since the salinity was near-zero and the flow rate was very low. The salinity was too noisy to edit except that a few very large spikes were removed from both files.

g.) Recalibration

ADD CHANNEL was used to add channel Temperature:Lab which was set equal to Temperature:Primary.

CALIBRATE was run using file 2014-01-tsg-recal1.ccf to apply the following correction to channel Temperature:Primary:   Temperature Primary = Temperature:Lab * 1.0581 – 0.7049

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Secondary, Temperature:Difference, Conductivity:Primary. UPloy0, Position:New,  and Flag channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats (except for salinity and fluorescence which were given fewer decimal places than standard). Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

A header check looked fine.

The cruise plot was added to the end of this report.

##### 26. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### 27. Loop File

An initial surface & loop file was prepared with rosette samples and all loop samples.

The CHE files were put through DERIVE to calculate sigma-t. They were then exported to a spreadsheet and sorted on pressure. All lines were removed except for the near-surface bottles (above 8m). (2014-01-loop-che.csv) The loop sample data prepared for the TSG processing were added to the file and lined up appropriately with the CHE data. The data were then ordered on time.

The sampling method column was added entered ROS or UWS for rosette data and true loop data, respectively.

Comments were added as required; for each flag other than 6, the files were checked for the relevant comment and those were entered in the 2014-01-loop-che.csv file.

A 6-line header was added and the original header was removed.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was then saved as 2014-01-surface-6linehdr.csv.

CONVERT was run to get an IOS Header file, followed by CLEAN to get start and stop times and positions.

Header Edit was used to add general comments from the CHE files. Comments were added concerning flags on samples from Niskin bottles. The flags from the loop samples were entered automatically in the conversion process.

The final file was named 2014-01-surface.loop. A track plot looks reasonable and a plot of temperature and salinity versus date looks reasonable.

# Particulars

PAR on: 1-11, 28, 46-48

12. Bottle #24 did not trip.

31. 22 Niskins at 2000db, not to go in CHE file. Pressure spikes.

59. Depth wrong in header - 3854

# 78. Partial cast – computer crash so returned to surface and restarted as 79.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 31Jan2013 | Factory |  |  |
| Conductivity | 2128 | 29Jan2013 | Factory |  |  |
| Secondary Temp. | 2374 | 31Jan2013 | Factory |  |  |
| Secondary Cond. | 3396 | 30Jan2013 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 5Mar2013 | Factory |  |  |
| SeaBird Fluorometer | 2356 |  |  |  |  |
| WetLabs ECO Fluorometer | 2215 |  |  |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Pressure Sensor | 75636 | 25Mar2013 | Factory |  |  |
| Altimeter | 1204 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2014-01

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 7Feb13 | Factory |  |  |
| Conductivity | 2487 | 7Feb13 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-953P | 26Mar03 | IOS |  |  |
| Temperature:Secondary | 0603 | Feb13? | “ |  |  |
| Flow meter | n/a | n/a |  |  |  |