## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 19 November 2014 | Note added about factory post-cruise calibration. See end of report for details. GG |
| 6 August 2014 | Corrected cruise dates in processing report. GG |
| 21 July 2014 | Salinity in CTD and CHE files recalibrated. See note at end of report. GG |

## PROCESSING NOTES

Cruise: 2014-03

Agency: Ocean Sciences Division

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia - Venus

Party Chief: Macoun P.

Platform: John P. Tully

Date: March 4, 2014 – March 11, 2014

Processed by: Germaine Gatien

Date of Processing: 18 June 2014 – 3 July 2014

Number of original HEX files: 16

Number of CTD files: 15 (1 cast surface only)	Number of bottle casts:	12

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1119), a WetLabs ECO Fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#4615), a pH sensor (#0692) and an altimeter (#1204).

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log analysis logs were in good order. There were no rosette sheets. Photos of the setup were not available.

CTD #0585 was used for 4 successive Vector cruises in early 2014; this was the first.

There were no rosette sheets and this likely led to the non-standard assignment of sample numbers. The dissolved oxygen and extracted chlorophyll samples were each given their own set of consecutive numbers. So from Niskin 1, Event 3, the CHL and DO samples were labelled as 1 and 8, respectively. This made normal construction of rosette files impossible, so new sample numbers were assigned. A table in section 4 of this report shows the correspondence between the original sample numbers and those assigned in processing; the new numbers start from 20 to avoid any confusion with the original numbers. A unique sample number should be assigned to each Niskin bottle fired and every sample from that Niskin should have the same sample number. If there were no samples of any sort taken, you can skip assigning a sample number, though it does not hurt to have them. For duplicate samples use a and b with the sample number, as was done for these data.

The dissolved oxygen samples were not suitable for recalibration of the CTD DO. They were few in number and were all taken at depths that are not ideal for this purpose, namely within a few meters of the seabed or surface. It is recommended that calibration samples be taken at least 10m above the bottom. Titrated DO values from pairs of bottles fired at the same level differ significantly. The data were recalibrated by using the results of the cruise that followed, 2014-06. For events #3 and #9 there was no draw temperature recorded, so Oxygen:Dissolved could not be derived in mass units.

PAR values are all very low, but all casts were run at night, so this is reasonable.

There was no salinity calibration sampling. (July 2014: Recalibration was applied by subtracting 0.004 based on the results of calibration sampling during other cruises.)

The ECO fluorescence values are about 40% higher than extracted chlorophyll. The range of CHL was small, from 0.7ug/L to 1.1ug/L. There are no significant outliers.

A CHE file was prepared for every cast with Niskin bottles fired, but in some cases

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained as well as analysis sheets for dissolved oxygen and extracted chlorophyll.

The PAR sensor is included in the configuration file but is not listed in the log. All casts were at night so any signal is likely to be small.

Event #9 contains only surface data as acquisition was run only to fire surface bottles. A bottle file should be prepared, but no profile file.

Extracted chlorophyll and dissolved oxygen data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained, but none of the sensors have been used since they were last calibrated at the factory. There were 3 April cruises with the same equipment, so all 4 will be examined together.

The calibration constants were checked for all instruments and the only errors found were some missing dates. One corrected file was saved as 2014-03-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2014-03-ctd.xmlcon to create CNV files.

Acquisition for event #6 did not begin until the CTD was coming up, so it was raised to the surface and then a complete cast was acquired. The first 6623 records were removed using a text editor since these records are from the first upcast.

A few casts were examined and all expected channels are present. The temperature and conductivity channels are fairly close on downcasts, but differ significantly on upcasts. It is not obvious that one is better than the other. There are many spikes in both channels in both downcast and upcast traces.

The PAR signal has the right shape but values are very low, as expected from night casts.

The pH and DO signals look normal.

The altimetry looks noisy at the bottom, but a useful estimate is possible for the bottom reading.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2014-03-ctd.xmlcon.

The files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. No problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file

ADDSAMP.csv.

At this stage sample numbers are usually added to the ADDSAMP file based on the rosette log records. However, there are no rosette sheets and the CHL and DO samples were each numbered sequentially starting with 1, but samples with the same # usually come from different casts. There were also more bottles fired than were sampled. The DO analysis log sheet offers some help, and the CHL sample labels are available which state what Niskin bottle was sampled and usually includes the station name. Using this information, it is possible to establish the source for each sample, but it was a tedious job. New sample numbers (Rosette Sample #) were assigned starting with #20 and the following table shows how these correspond to the numbers assigned at sea:

| Event # | Niskin # | Station | Pressure | Rosette Sample # | CHL Sample # | DO Sample # | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1 | SI-VIP | 95.3 | 20 |  | 1 |  |
| 1 | 2 | SI-VIP | 95.3 | 21 |  | 2 |  |
| 1 | 3 | SI-VIP | 95.3 | 22 |  |  |  |
| 1 | 4 | SI-VIP | 95.3 | 23 |  |  |  |
| 2 | 1 | SI-BPS | 211 | 24 |  | 3 |  |
| 2 | 2 | SI-BPS | 211 | 25 |  | 4 |  |
| 2 | 3 | SI-BPS | 211 | 26 |  | 5 |  |
| 2 | 4 | SI-BPS | 211 | 27 |  | 6 |  |
| 2 | 5 | SI-BPS | 211 | 28 |  | 7 |  |
| 2 | 6 | SI-BPS | 211 | 29 |  |  |  |
| 3 | 1 | QA1 | 5.1 | 30 | 1 | 8 | DO analysis log not completed at sea-draw temperature not available |
| 3 | 2 | QA1 | 2.6 | 31 | 2 |  | DO analysis log not completed at sea-draw temperature not available |
| 4 | 1 | QA2 | 5.2 | 32 | 3 |  |  |
| 4 | 2 | QA2 | 1.8 | 33 | 4 |  |  |
| 5 | 1 | QA3 | 5.4 | 34 | 5 |  |  |
| 5 | 2 | QA3 | 2.2 | 35 | 6 |  |  |
| 6 | 1 | QA4 | 5.4 | 36 | 7 |  |  |
| 6 | 2 | QA4 | 5.4 | 37 |  | 9 |  |
| 6 | 3 | QA4 | 1.9 | 38 | 8 | 10 |  |
| 6 | 4 | QA4 | 1.9 | 39 |  |  |  |
| 7 | 1 | QA5 | 5.3 | 40 | 9 |  |  |
| 7 | 2 | QA5 | 1.9 | 41 | 10 |  |  |
| 8 | 1 | QA6 | 5.6 | 42 | 11 |  |  |
| 8 | 2 | QA6 | 2 | 43 | 12 |  |  |
| 9 | 1 | QA7 | 4.8 | 44 | 13 | 11 | DO analysis log not completed at sea-draw temperature not available |
| 9 | 2 | QA7 | 4.8 | 45 |  | 12 | DO analysis log not completed at sea-draw temperature not available |
| 9 | 3 | QA7 | 2.1 | 46 | 14 |  |  |
| 9 | 4 | QA7 | 2 | 47 |  |  |  |
| 10 | 1 | HB | 5.1 | 48 | 15 |  |  |
| 10 | 2 | HB | 5.1 | 49 |  | 13 |  |
| 10 | 3 | HB | 1.9 | 50 | 16 |  |  |
| 10 | 4 | HB | 1.8 | 51 |  | 14 |  |
| 11 | 1 | CNODE | 300 | 52 |  | 15 |  |
| 11 | 2 | CNODE | 300 | 53 |  | 16 |  |
| 11 | 3 | CNODE | 300 | 54 |  |  |  |
| 11 | 4 | CNODE | 300 | 55 |  |  |  |
| 12 | 1 | ENODE | 174 | 56 |  | 17 |  |
| 12 | 2 | ENODE | 174 | 57 |  | 18 |  |
| 12 | 3 | ENODE | 174 | 58 |  |  |  |
| 12 | 4 | ENODE | 174 | 59 |  |  |  |

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2014-03-bot-hdr.txt. Comments were added about the sample number problem.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-03chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2014-03chl.csv. The sample numbers were adjusted as explained in the table above. The file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-03oxy.xls which includes flags, comments and a precision study. Draw temperatures are available for most samples. The spreadsheet page with the final data was simplified and the file was then saved as 2014-03oxy.csv. The sample numbers were adjusted as explained in the table above. That file was converted into individual *.OXY files.

The CHL and OXY files were merged with CST files in 2 steps.

After the 2nd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet. There was an error in event numbers for one cast in the analysis spreadsheet. That was corrected and the merge process repeated and no further errors were found.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There are too few data for a believable fit, and all the data available come from either the bottom or the top 5m where comparisons are usually poor. At the bottom poor flushing is likely an issue and at the surface incomplete flushing and high gradients often occur. The following table illustrates that 2 bottles at the same level differed greatly except in the anoxic waters in Saanich Inlet.

| event # | differences between bottles from same level | pressure | Comment |
| --- | --- | --- | --- |
| 1 | 0.086 | 95 | bottom |
| 2 | 0.000 | 211 | bottom |
| 9 | 0.243 | 5 | surface |
| 11 | 0.053 | 300 | bottom |
| 12 | 0.055 | 174 | bottom |

A flag is suggested for sample #42 (#11 in the original numbering), event #11 as it is out of line with both the CTD and the other bottle at that level. A 3 flag is sufficient since it is a surface sample and may be ok given that there is often high variability near the surface.

To recalibrate these data, we could choose results from 2 other cruises that used these sensors in April 2014 and had good calibration sampling. The first was 2014-06 for which the following correction was found:

CTD DO Corrected = CTD DO * 1.0300 + 0.0744   (2014-06)

To check if that is a reasonable fit for this cruise, the offset was forced to 0.0772 and a few outliers removed including all points from the hypoxic layer. The resulting fit had a slightly higher slope:

CTD DO Corrected = CTD DO * 1.0322 + 0.0744   (2014-03 with offset set to 0.0744)

The result of cruise 2014-15 was also considered:

CTD DO Corrected = CTD DO * 1.0325 + 0.1019   (2014-15)

Fixing the offset to 0.1019 the slope is lower:

CTD DO Corrected = CTD DO * 1.0271 + 0.1019   (2014-03 with offset set to 0.1019)

It seems wisest to use the results of cruise 2014-06 since it is closest in time to 2014-03. The correction found for 2014-15 is larger than that for 2014-06 and that is likely to be due to calibration drift. Since such drift generally increases with time, it is less likely that 2014-15 suits this cruise as well as that from 2014-06.

Plots of titrated DO and CTD DO against CTD salinity were examined and no problems were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The CTD fluorometer was a WetLabs ECO sensor

The range of CHL is small. The ratio of ECO/Extracted CHL is ~1.5 for CHL~0.7ug/L and it drops to ~1.3 for CHL ~1.1ug/L.

The ECO readings are high by an average of 40%.

For full details of the comparison see file 2014-03-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on data from 2014-06 with results considered difficult to judge. A setting of +4s was applied, but later in processing it was decided that value was too high and a second shift of -1.5s was applied. So for these data a setting of +2.5s was tried and the results looked good overall.

##### CELLTM

The results of the tests run on 2014-06 were applied.

CELLTM was run using (α = 0.02, β=9) and (α = 0.0245, β=9.5) for the primary and secondary conductivity.

##### DERIVE  and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 2 of the deeper casts to examine differences between sensor pairs.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-03-0006 | 320 | ~0 XN | +0.0004 | +0.004 | Very steady |
| 2014-03-0014 | 320 | +0.0004 XN | +0.00035 | +0.0042 | Very steady |

The differences between conductivity and salinity channels are fairly steady, but the temperature differences are extremely noisy.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. No errors were found, though times and positions were a little different in a fairly random way. The log time is later than the header time by from 1 to 8 minutes. The same thing was seen during the cruise that followed. The difference is likely due to the timing of when the log record was entered, though the computer time may be a little fast.

A header check was run. Negative pressure and fluorescence signals were found but these are likely just near-surface spikes.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.7db which is reasonable.

The altimeter and water depth readings from the headers of the CLN and SAM files were exported to a spreadsheet.

Corrections were made to SAMAVG files as follows:

- The altimetry header was removed from casts 9 and 10 as a spike had been interpreted as a signal.

- The water depth was changed for casts 3 and 11 to match the log entries

- The water depth was added to cast #4 based on the log entry.

- The water depth was removed from cast #7; it is missing from the log and clearly wrong in the header.

After those corrections the SAM files were bin-averaged again and re-merged with the bottle data.

Corrections were made to the CLN files as follows:

- Water depths were changed to match the log for cast #3 and added for casts 4 and 14.

- Water depth was removed from casts #7 and 11 as it is clearly wrong both in the file and it is missing from the log book.

- Event #9 was removed since there is only surface data. A CHE file will be prepared, but no CTD file.

##### Shift

Fluorescence

SHIFT was run on the WetLabs ECO fluorescence channel in all casts using the usual advance of +12 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. During 2014-06 tests were run using values between +60 and +80 and the best results were found with +70 records, though the best setting varied with depth. For some features that reading does not seem high enough, but that is likely because there are some reversals and the pH sensor does not reach the peak which complicates the comparison.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration. The alignment looked good.

Conductivity

Tests were run on 2 casts using a variety of shifts and the best results were found with -0.7 records for the primary and +1.2 records for the secondary. For cruise 2013-06 which followed the same setting was found for the primary, but the secondary was aligned using -0.6 records. A test was made on one cast of 2013-06 using +1.2 records just to be sure it was not a better choice and the result was very poor. So clearly something changed between these two cruises. The CTD was moved from one ship to another so the setup might have changed a little.

SHIFT was run using -0.7 records for the primary conductivity and +1.2 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity sensors were both recalibrated in late December 2012 or early January 2014 and they were used for 4 cruises in spring 2014. For 2014-03 there was no salinity calibration sampling, but for 2014-06, 2014-27 and 2014-15 there was. The primary salinity was found to be low by 0.0030, 0.0040 and 0.0032, respectively. The secondary salinity was found to be high by 0.0011, 0.0000 and 0.0007, respectively. The best sampling was for 2014-06 but even for that cruise there was a lot of scatter in the comparison; when only a few deep bottles were included the primary was low by 0.0035 and the secondary was high by 0.0002. The difference between channel pairs was consistently ~0.004. During a cruise in June 2014 it was found that the differences between channels had increased somewhat, but salinity calibration data for that cruise are not yet available.

Dissolved Oxygen

The sensor was calibrated at the factory in January 2014. This was the first use and it was also used for 2014-06, 2014-27 and 2014-15. From 2014-06 there was a good comparison and the fit had slope/offset = 1.03/+0.0744, though there were 2 casts with severe outliers between the surface and 100m. The calibration sampling of 2014-03 and 2014-27 were each insufficient for good fits, but both looked as though the 2014-06 correction is likely appropriate. 2014-15 had more sampling and was in a different area; the fit had a slope/offset = 1.0325/0.1019. While the fits for 2014-15 varied depending on how outliers were identified, no choice would lead to a fit similar to 2014-06. So there was likely some calibration drift, though there could be some effect on the fit due to differing water properties or sampling or analysis details.

Pressure

The sensor was recalibrated in December 2013 and this was the first known use. In 3 cruises that followed there was no evidence of a need to adjust the offset.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The temperature values were low at many stations. The depths of the low values were between 20 and 150m in the south-east part of the survey and those casts were well-mixed at those levels.  The low temperatures were near the bottom for the in the north-west section of the survey.

Salinity values were all within the climatology except for a few low salinity values in the top 5m of one cast.

None of the outliers look like they are due to calibration problems. The 3-standard deviation climatology is too severe for this region.

Repeat Casts – There were no repeat casts; there were some nearby casts, but given the local variability there is no expectation of good repeatability, so no comparisons were done.

Post-Cruise Calibration – There were no post-cruise calibrations available at the time of processing.

##### CHANNEL CHOICE STUDY

From 2014-06, 2014-27 and 2014-15 the secondary salinity was closer to the bottles than the primary. There may be a little evaporation of salinity samples, but analyses were done quickly so this should not be a major problem. There could be some effect from incomplete flushing of Niskin bottles. However, when only data below 175db were included for 2014-06, the primary salinity was found to be low by an average of 0.0035 and the secondary was high by 0.0002. These samples are from areas with low salinity gradients where complete flushing is not critical. Similar results were found for 2014-27 and 2014-15.

There is not much difference between the T-S plots though the primary seems a little noisier.

The secondary temperature and salinity channels were selected for archiving.

##### DETAILED EDITING

The secondary temperature and salinity channels were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to minor misalignments or other instrumental problems and likely to affect the bin-averaged values, and records corrupted by shed wakes, including some records from near the top and bottom of the casts.

All files required some editing.

All EDU files were copied to *.EDT.

##### Initial Recalibration

The pressure does not appear to require recalibration.

The secondary salinity was found to be within 0.001 from 3 other cruises that used these sensors after 2014-03. The best comparison was from 2014-06, but even for that cruise there was a lot of scatter. There is insufficient evidence to justify recalibration of secondary salinity which is likely good to within 0.001.

For dissolved oxygen there were too few bottles for a good fit, but the results look compatible with the results of 2014-06. For that cruise there were some odd severe outliers from 2 casts, but when those were excluded the comparison produced a fairly tight fit. That result will be used for 2014-03.

CALIBRATE was run using file 2014-03-recal1.ccf to apply the following correction to the SAM and MRGCLN2 files:

CTD DO Corrected = CTD DO * 1.03 + 0.0744

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibrations were done properly. (See 2014-03-dox-comp2.xlsx.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

A comparison is usually made between downcast CTD DO and upcast bottles, but there are too few bottles and too much scatter to expect reasonable results for this cruise, so this step was skipped.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots look fine except for small instabilities near the surface that may well be real. There is a PAR signal but the maximum values are all extremely low as is expected because all CTD casts were run at night. Profiles mostly look fine though there are a few deep spikes in transmissivity and fluorescence. No editing was applied as these features are at sites with fairly strong mixing.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH and PAR  data are nominal and

unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the results of cruise 2014-06.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2014-03-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values at 2m or 3m were all between 85% and 100%, with the highest values associated with higher near-surface DO gradients. The lower values are in better mixed surface waters.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

A note was added to the headers of events #3 and #9 to explain that Oxygen:Dissolved values are not available in mass units because the draw temperature was not available.

Standards check was run on all files and no errors were found.

A header check was run on the final files.

No further errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

6. Acquisition started during upcast, so brought to surface and full cast run.

7. Repeated because archiving started during upcast.

9. No CTD file – only used to trip surface bottles – prepare CHE file only.

NOTE: RECALIBRATION 21 July 2014

When these data were originally processed, it was believed that the secondary salinity did not need recalibration based on comparisons of CTD and bottle salinity during cruises 2014-06. There were some doubts about that comparison because of incomplete flushing of Niskin bottles, but it was the best available. Later, comparisons from 2014-21 and 2014-18 showed that the primary salinity was close to bottles, while the secondary was high. So, the primary CTD salinity is presumed to be reliable. The secondary salinity data from this cruise were recalibrated by subtracting 0.004, which was roughly the difference between the two channels.

The two conductivity sensors will be returned to the factory for a post-cruise calibration. When the results are available, further recalibration will be applied, if needed.

G. Gatien

Note: 19 November 2014: Post-cruise factory calibrations show that the secondary salinity was high by about 0.005 by mid-August. Based on observations during a series of cruises, it appears that the calibration applied to these data in July was appropriate, though there remains some doubt about when drift occurred.

G. Gatien

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 31Dec2013 | Factory |  |  |
| Conductivity | 1766 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 4700 | 4Jan2014 | Factory |  |  |
| Secondary Cond. | 3321 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 21Jan2014 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | 14Jan2014 |  |  |  |
| WetLabs ECO Fluor. | 2216 | n/a |  |  |  |
| Pressure Sensor | 77511 | 30Dec2013 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |