## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |
| 19 Nov. 2014 | Note added to summary section giving results of post-cruise factory calibrations. G. Gatien |

## PROCESSING NOTES

Cruise: 2014-06

Agency: Ocean Sciences Division

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: April 2, 2014 – April 5, 2014

Processed by: Germaine Gatien

Date of Processing: 18 June 2014 – 22 July 2014

Number of original HEX files: 80 (1 split cast)

Number of CTD files: 79	Number of bottle casts:	22 (+1 with no sampling, not processed)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1119), an ECO Fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#4615), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (#1204).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order except that the personnel list was missing. Photos of the setup were available.

Dissolved oxygen samples from 2 casts were rejected due to unbelievably high values, including much higher surface saturations than previously seen at those sites. The data from the 2 casts fell into a linear group on a plot of bottle DO versus CTD dissolved oxygen. No problem was found in the CTD data and the linearity argues against contamination, mis-sampling, misfiring Niskin bottles or errors in entering flask numbers. This would appear to be an erratic problem with the analysis system.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1mL/L from 0 to 20db

±0.4mL/L from 20db to 100db

±0.2mL/L from 100db to 250db

±0.05mL/L below 250db

The ratio of ECO Fluorescence / Extracted CHL is >1 at the lowest end of the CHL range and then falls gradually it reaches about 0.6 for CHL~3ug/L and stays at about that level for higher CHL. A plot of ECO fluorescence versus extracted chlorophyll has a slope of 0.58 when the trendline is forced through the origin, but that fit is poor for low CHL.

CTD #0585 was used for 6 successive cruises in early 2014. The two salinity channels differed by ~0.004 in March even though all T and C sensors had recently been recalibrated at the factory. The difference between the sensor pairs increased slightly with time and was ~0.0054 in June.

The salinity analysis was completed within 2 weeks of the beginning of the cruise.

The secondary temperature and salinity channels were chosen for most casts, but the data were poor near the surface for event #33, so the primary temperature and salinity were chosen for archiving for that cast. It was later discovered that during 2 offshore cruises the primary salinity was closest to the bottles, while the difference between the 2 channels had not changed much. So for this cruise it was assumed that the primary salinity is correct. The secondary salinity was recalibrated to match the primary by subtracting 0.0041, which is an average of the difference between the channels as noted during bottle stops. The difference between the results of the comparison for this cruise and that from the later cruises is likely how well the Niskin bottles flush. The descent rate of the CTD was relatively quiet for this cruise; for the few casts where it was noisier, the comparisons looked a little closer to the offshore ones.

Note: 19 November 2014: Post-cruise calibration shows that the secondary salinity was high by about 0.005 by mid-August. Based on observations during a series of cruises, it appears that the calibration applied to these data was appropriate, though there remains some doubt about when drift occurred.

G. Gatien

This study shows the limitations of comparisons in inshore, waters. The best results will come from well-mixed waters and wherever the descent rate of the CTD is noisy. It is still worth doing the sampling, but recalibration should only be considered if the change seems too large to be explained by poor flushing, or if there is offshore sampling to confirm the result.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

One cast was interrupted on the way up. Acquisition was restarted and the files were named 2014-06-0001up.hex. The standard approach to is to assign a new event # for the 2nd half of such interrupted casts. However, since all the downcast data are in the 1st file and the bottles were all fired after the 2nd file was started, the rosette file can just be renamed as 2014-06-0001.ros after conversion.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen, extracted chlorophyll and salinity.

The deck pressure was recorded once; it was 0.56db.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments and the only errors found were some missing dates. The only change to the configuration files through the cruise was to add SPAR parameters that were omitted earlier in the cruise.

One corrected file was saved as 2014-06-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2014-06-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. The temperature and conductivity channels are sometimes close but there are many spikes in both channels in both downcast and upcast traces. Unusually, the upcast traces are often closer than the downcast ones, but that is not always the case and there is more detail in downcast traces.

The altimetry looks useful and fluorescence, transmissivity, PAR and SPAR look normal. The pH signal looks odd for some early casts with downcasts having odd shapes, but these were well-mixed casts with little variation, so this is likely insignificant. In later casts checked the signals look normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2014-06-ctd.xmlcon.

File 2014-06-0001up.ros was renamed as 2014-06-0001.ros.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. There was some noise in a few points of the primary salinity around 100db in cast #76; this was cleaned using CTDEDIT. The output was copied to BOT.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

All bottles were fired at the surface of event #1, so those lines with no sampling were removed from the ADDSAMP file.

For event #73 two bottles were fired at 300db, but they were given the same sample number. This is a problem since the merging process can’t cope with it. The contents of 2 bottles fired at the same depth do differ. The second case was renamed sample #9248. This will require some manipulations later in the merging process. (The 2nd bottle was fired just to get water, so this line will be removed later.)

There were bottles fired for event #85, but there was no sampling and no need for a bottle file.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2014-06-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-06chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2014-06chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-06oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2014-06oxy.csv.

An error was found in event #s – bottles were from event #7 not #6.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2014-06SAL.xls. The analysis was done within 9-14 days of collection. The files were simplified and saved as 2014-06sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-06nuts.xls. This includes a precision study.

Then the file was simplified, reordered on sample numbers and saved as 2014-06-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A few problems were found and corrected:

- Flag 1 was added to 3 CHL entries with the note “no sample found”.

- Event #49 had DO and Salinity samples marked as from event #47.

- Event #73 – because of sample # irregularities the order of samples had to be rearranged.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure shows no severe outliers. Most outliers are associated with large standard deviations in the CTD salinity channels and others are at depths where flushing in high gradients might explain the differences. When 6 out of 22 samples are excluded, the primary salinity is lower than bottles by an average of 0.0030 and the secondary is high by 0.0011. The standard deviation in the fits was 0.0016 for both channels. When only data below 175db are included, the primary salinity was found to be low by an average of 0.0035 and the secondary was high by 0.0002.

When the only 5 bottle casts with a noisy descent rate and low standard deviation in the CTD salinity are selected, the primary salinity is low by an average of ~0.002. This may be a sign that the results of the comparison are weak due to incomplete flushing of Niskin bottles.

There are no obvious bad bottles.

A plot of differences against time shows no obvious pattern

For full details for the COMPARE run see file 2014-06-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter, but there are many points that fit in a fairly tight group. When outliers are removed from that group, the fit is:

CTD DO Corrected = CTD DO * 1.0300 + 0.0744

If the offset is set to 0 the fit is:

CTD DO Corrected = CTD DO * 1.0446

SeaBird recommend doing a fit with the offset set to 0. This does not always seem to produce the best fits. For these data the R2 value is higher when the free offset is allowed and it looks better to the eye.

A study of the outliers shows many from the Saanich Inlet cast which is not unusual, and some from the top and bottom of casts, again not unusual. But the most severe outliers come from 3 casts including all bottles from above 100m for events #7 and #19 and 1 bottle from cast #76 at 149m with differences between bottles and CTD of up to 4mg/L. Even the deepest bottle from cast #19 is an outlier, just not as notable as the others. And the deep bottles of event #7 while mostly not out of line, do fall close to the odd fit of the other bottles from that cast. Only one of the severe outliers was flagged by the analyst. When only the bottles from events #7 and 19 above 100m are included, the fit found was:

CTD DO Corrected = CTD DO * 1.881 - 2.0465

What could explain these odd outliers?

- The one bottle that fits the pattern from cast #76 stands out as an outlier in the profile, but the CTD dissolved oxygen and temperature traces are quite complex around that depth with several reversals. This is not a large outlier and is likely due to high local gradients.

- Examination of full profiles of CTD DO shows nothing remarkable about the other 2 events.

- Event #7 was the middle station of the line across the mouth of Juan de Fuca (station 102) and Event #19 was at station ADCP near Victoria; they were not consecutive or close to each other.

- The titrated DO values at the mouth of Juan de Fuca are ~10mL/L down to 30m, which seems remarkably high. The CTD DO is also well mixed to 30m but with maximum values ~7mL/L.

- The temperature shows that event #19 is very well mixed between 20db and 75db; the nearby stations are also quite well mixed but not as much as #19. However, both the CTD and bottle DO values increase through the well-mixed section.

- The draw temperatures mostly increase in step with the CTD temperature; there are some obvious outliers but those are near the surface.

- The transmissivity and fluorescence show nothing particular to those casts compared to neighbouring ones. In the central Strait of Georgia DO values of >8mL/L are found near the surface, but they are associated with high fluorescence, which is not the case for casts #7 and 19.

- There were many reports of particles (including zooplankton) in samples including at #7 and 19, but they were only near the surface, and similar remarks are found for other casts which have no such outliers.

- The maximum DO values are 12.7 and 8.6mL/L for casts #7 and #19, which seem remarkably high.

- One of the odd casts was sampled by the night watch, one by the day watch and the same experienced analyst ran the titrations for the whole cruise.

- Sampling from the wrong bottles doesn’t explain the very high values at the surface since there are no other bottles from which such values would look reasonable.

- Similarly, bottles misfiring consistently won’t explain the high DO values. The values seen at the surface would look even worse if they occurred deeper in the cast, and it would be unlikely that such misfires would lead to a linear fit of bottles versus CTD DO.

- Contamination of every bottle from 2 well-separated casts seems unlikely, and it is especially difficult to believe that it would lead to a linear fit against CTD DO.

- The flasks used for these casts were mostly the same, but those flasks were used for other casts as well and the volumes are consistent. Even if there were errors in flask identification or volumes, this is unlikely to lead to a linear fit against CTD DO for those casts.

- It looks most likely that these odd values are due to an erratic problem in the analysis system, leading to erroneously high thiosulphate readings.

All bottles from casts #7 and 19 will be replaced with pad values at the MRGREO stage, flagged 5 and have the comment “Severe outlier in comparison with CTD DO; likely analysis system error” added.

NOTE: The analyst found an explanation for these problems – see the NOTES page in spreadsheet QF2014-06-OXY.xlsx. July 25, 2014.

Plots of Titrated DO and CTD DO against CTD salinity were examined and casts #7 and 19 look odd. No other outliers were noted.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a WetLabs ECO sensor

The ratio of ECO/Extracted CHL is >1 at the lowest end of the CHL range and then then falls gradually until CHL~3ug/L at which point it is about 0.6. It stays at roughly that level for CHL>3ug/L.

A fit of ECO versus CHL has a slope of 0.58 if forced through the origin. That fit is not very good for CHL<2ug/L.

For full details of the comparison see file 2014-06-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a 4 casts to see what setting brought the vertical offset between the downcast and upcast DO traces into the best agreement with the temperature traces. The best setting was between 3.6s and 4.4s. The temperature traces are very noisy making a judgment difficult, but +4s looks like the best choice overall. (It was later found that this setting was too high and a further alignment of -36 records or -1.5s improved the alignment of dissolved oxygen with temperature. See section12.)

##### CELLTM

A variety of settings were tested on 4 casts. The upcast and downcasts differed so much that the results were difficult to interpret. The primary seemed slightly improved by using (α = 0.02, β=9), while the results for the secondary were too confusing to decide, so the default setting of (α = 0.0245, β=9.5) looks like the best choice.

CELLTM was run using (α = 0.02, β=9) and (α = 0.0245, β=9.5) for the primary and secondary conductivity.

##### DERIVE  and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-06-0043 | 300 | +0.0002 | +0.00042 | +0.0042 | Very steady |
| 2014-06-0053 | 300 | ~0 VN | +0.00041 | +0.0043 | Very steady |
| 2014-06-0053 | 350 | ~0 VN | +0.00044 | +0.0044 | Very steady |
| 2014-06-0073 | 300 | +0.0013 VN | +0.0005 | +0.0044 | Very steady |
| 2014-06-0073 | 340 | +0.0017 VN | +0.0006 | +0.0043 | Very steady |

The differences between channels are mostly small except that there were some large temperature differences during cast #73 near Texada Island.

Plots of the pair of temperature channels together with the differences between them suggest that more difference spikes are associated with spikes in the primary than the secondary. At other times such spikes are associated with features in both channels that are misaligned. A T-S plot of each channel pair shows stability being a little higher for the secondary for one cast and for the primary for another. This will be investigated more thoroughly after conductivity channels have been aligned.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book.

- Times and positions were a little different in a fairly random way; the log time is later than the header time by from 0 to 13 minutes. The larger time differences are associated with larger position differences. Most differences are <5 minutes, with only 9 cases larger. So rather than a faulty clock time, this is likely due to the log entry being recorded at slightly different stages of the operation from cast to cast and possibly some variation among clocks used to record log time.

- The station name for event #76 was entered as 12 in both the log and the header. There was another station 12 entry with a different position. It is clear from a track plot and previous cruises that the entry for #76 is incorrect and should be station 14.This was corrected in the IOS and CLN files and in the bottle files.

- One station name was missing – cast #35. That was fixed in IOS and CLN files. There was no bottle cast at that site.

A header check was run. The fluorescence signal from cast #35 had a minimum of -0.53 but this looks like a surface spike that will be removed by DELETE.  All other parameters had reasonable values.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.1db which is reasonable for the Vector. The single deck pressure recorded was 0.56db. This might suggest the pressure is reading too high, but a single deck reading is not significant. The lowest pressure recorded by the CTD, 0.485db, occurred at the end of a cast with the pumps turned on; this was not just a single point spike and all variables show the CTD was definitely in the water. No recalibration of pressure is justified.

Depth readings were missing for the first 5 casts, so the information was added based on the header entries except for event #5. There is a note in the log that station 61 (event #5) usually has a depth of 274m but the depth entered in the log was 214m. Checks of a few other cruises showed some variation form 205m to 285m. Given a minimum altimeter reading of ~33.4m and maximum pressure of 233db (~231.5dbm), we get a rough estimate of depth as 265m. There could have been significant drift between the depth reading and the CTD getting close to the bottom. Whatever the reason, the depth entry is inappropriate for the time of the cast. The depth entry was not in the headers and was not added.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for most casts and the headers look appropriate despite a very noisy signal near the bottom of some casts. However, in some cases there was no entry due to such noise and an estimate can be made by examining the plots of altimetry at the bottom. Those estimates were entered into the headers of the CLN files, together with a comment how the value was determined. (Events #47, 61, 72 & 75)

The bottle files were also examined and 4 of them lack an altimeter header due to heavy noise in the altimetry at the bottom of the casts. The maximum pressures were the same as those found for the full files, so the values from those files were used to add header entries to the bottle files, and the comment was adapted to explain that.

##### Shift

Fluorescence

SHIFT was run on the WetLabs ECO fluorescence channel in all casts using the usual advance of +12 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +60 and +80. The best setting seems to vary with depth, but +70 looks like the best choice overall.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel and it does look like the original alignment went too far. So SHIFT was run using a setting of -36 records. The alignment looks better after this step. The cumulative shift is +60records or 2.5s.

Conductivity

Tests were run on 5 casts using a variety of shifts and the best results were found with -0.7 records for the primary and -0.6 records for the secondary.

SHIFT was run using -0.7 records for the primary conductivity and -0.6 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

This was the second cruise using these temperature and conductivity sensors since they were calibrated at the factory. The previous cruise had no salinity sampling.

Dissolved Oxygen

The sensor was calibrated at the factory in January 2013 but not used until 2014. There was only one use before this cruise with limited DO sampling. A comparison with bottles is too limited to suggest a recalibration method, but the results are reasonably close to those of 2014-06.

Pressure

The sensor was recalibrated in January 2014 and it was used for only 1 other cruise since then. The factory offset was used for that cruise and the results look reasonable.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.

Temperature values were low below 220m in the Strait of Georgia. All temperatures were within the climatology in Juan de Fuca and Haro Straits.

Salinity was low near the surface at most casts in Juan de Fuca Strait, with low values as deep as 50m on the south side near the mouth and below 120m at station 70, cast 15. All salinity data was within the climatology in the Strait of Georgia.

None of these outliers look like evidence of calibration drift. The 3-standard deviation climatology is too severe for this region.

Repeat Casts – The only repeat casts include one very shallow cast and another deeper, so they are not suitable for testing repeatability of data, but on a T-S plot they do look reasonably close.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### CHANNEL CHOICE STUDY

COMPARE shows that the secondary salinity data are closer to the bottles than the primary. Evaporation of samples and/or incomplete flushing could mean the secondary is not as low as it appeared in the comparison. But the analysis was done fairly quickly so evaporation should be small. And even at depths where local gradients make flushing less critical, the secondary does look close to the bottles.

An examination of T-S plots using the two channel pairs suggests there is little difference in quality, though the primary had a few more spikes and unstable features.

##### DETAILED EDITING

Since the secondary temperature and salinity channels seem a little less spikey and the salinity closer to the bottles, those channels were selected for archiving, and hence, editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts.

Cast #4 required no editing.

All other files required some editing.

For cast #33 the primary channels looked much better near the surface, so those channels were edited.

All EDU files were copied to EDT.

##### Initial Recalibration

At this stage processing was suspended to see what could be learned about the salinity calibration from later cruises using the same sensors. The results of 2 offshore cruises showed that the primary salinity was very close to bottles rather than the secondary as was concluded in section 5. The difference between the results of the comparison for this cruise and the later cruises is likely how well the Niskin bottles flushed. The descent rate of the CTD was relatively quiet for this cruise; where it was noisier, the primary salinity was closer to bottles than at the quieter casts. And for 2014-27 which followed and had even steadier descent rates the CTD primary salinity was even further from the bottles.

During cruise 2014-03 which preceded this cruise there was no salinity calibration sampling. The difference between the two channels was found to be 0.004, which is larger than expected given that the sensors had just been recalibrated. The difference grew slightly through the series of cruises examined until it was about 0.0054. The primary salinity was very close to bottles during 2014-21 in May, a cruise that was in offshore waters with deep sampling and quick salinity analysis. For 2014-18 in June, the primary was slightly lower than bottles, but that is likely explained by the fact that salinity analysis was slightly delayed which may have led to a little evaporation of bottle samples. So, it looks best to assume that the primary CTD salinity is correct until post-cruise calibration information becomes available. Recalibration of the secondary salinity channel is thus best done by subtracting the difference between the two salinity channels.

The secondary salinity will be recalibrated to match the primary by subtracting 0.0041, which is an average of the differences between the channels as noted during bottle stops.

The pressure does not appear to require recalibration.

CALIBRATE was run using file 2014-06-recal1.ccf to subtract 0.0041 from the secondary salinity and to apply the following correction to DO in the SAM and MRGCLN2 files:

CTD DO Corrected = CTD DO * 1.0300 + 0.0744

COMPARE was rerun for salinity and dissolved oxygen using the recalibrated values. The results confirm that the recalibrations were done properly. The average of differences in the fit for dissolved oxygen is +0.004mL/L with a standard deviation of 0.028mL/L. For salinity the average of differences is -0.041 for both channels.

CALIBRATE was then run on the EDT files.

One cast was checked to see that step brought the two salinity channels closer, and it did.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. Bottles from casts #7 and 19 were excluded as were large differences.  In a few cases there were no CTD data available sufficiently close to surface bottles. The differences show that the CTD DO is higher than the bottles by an average of ~0.04mL/L, but the results vary with depth. The descent rate was very steady for many of these stations which may decrease the efficiency of flushing. However, the DO gradient is mostly low for this cruise, so poor flushing is not likely to lead to large errors. Bottom bottles seem a little out of line and flushing can be a problem there too, though the error is in the opposite direction. When those bottles are excluded, the fit of differences versus pressure is fairly flat with an average difference of +0.042mL/L. Two very well-mixed casts were examined in detail. While the vertical gradients are low, there is a lot of fine-scale noise in these profiles in both temperature and DO, and differences on the order of 0.04mL could still be due to incomplete flushing. The area of lowest gradient for cast #32 is between 150m and 200m, and the CTD DO values are lower than the bottles there. So it is not clear that the differences are due to errors in the CTD calibration.

No further calibration was applied.

See 2014-06-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2014-06-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots are fine – the well-mixed casts do show many slightly unstable features, but that is normal for this region. Profiles look fine.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except #33 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

REMOVE was run on cast #33 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH, PAR and PAR:Reference data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1mL/L from 0 to 20db

±0.4mL/L from 20db to 100db

±0.2mL/L from 100db to 250db

±0.05mL/L below 250db

The primary CTD salinity channel was used for events #33; it was

recalibrated to match the secondary channel which was

selected for all other casts.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2014-06-proc.doc.

One line was removed from event #73 and the headers adjusted. There were no associated analyses – the bottle was just fired to get water.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values <85% were all in Haro Strait or Discovery Passage, areas known for strong vertical mixing. The highest values, >125%, were found in the central Strait of Georgia and Saanich Inlet. Near the mouth of the Strait of Georgia, values were between 95 and 115%. In the eastern section of Juan de Fuca Strait and the southern part of the Strait of Georgia, values are mostly between 85% and 95%. The lowest DO saturation values are at casts with low CTD fluorescence maxima and vice versa. The values are a little higher than were seen in this region in April 2013, but the pattern is similar. During the 2013 cruise there were more variable weather conditions, likely leading to stronger mixing, and the fluorescence maximum was much lower than in 2014. The saturation values do not suggest any problem with calibration.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

Bottles were removed from file #1 where no sample numbers had been created.

To study the apparent bad bottle data from 2 casts, % DO saturation was calculated for the bottle values and CTD DO values in the bottle files. Then the surface saturations were compared for stations along Juan de Fuca Strait. The following table shows the results.

| Station Name | Event # | Surface % saturation based on titrated sample | Surface % saturation based on CTD DO at time of surface bottle firing |
| --- | --- | --- | --- |
| SI | 1 | 133% | 129% |
| 102 | 7 | 161% | 106% |
| 75 | 10 | 107% | 107% |
| 72 | 13 | 104% | 103% |
| 69 | 16 | 105% | 105% |
| ADCP | 19 | 128% | 87% |
| 62 | 26 | 82% | 82% |

The Saanich Inlet cast is the one that usually has the highest % saturation observed in the Strait of Georgia/Juan de Fuca Strait region, and it is for this cruise too based on CTD DO data. For the bottle data the highest value is at the mouth of Juan de Fuca Strait where we expect values in the 95-115% range. Even more striking is the observation at station ADCP where strong mixing usually keeps DO saturation low.  This confirms that for events #7 and #19 the CTD DO data are reasonable and the bottle data are not.

The bottle DO values from casts #7 and #19 were replaced with pad values and flag 5, and the following comment was added:

Severe outlier in comparison with CTD DO; likely analysis system error.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check was run on all files and no errors were found.

A header check was run on the final files.

No further errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars (Notes from Log Book)

1. Two files – one down, one up from 100db. The second was named 1UP.

5. Min Altimeter reading 33.4m – Depth reading 214 - Usual depth for this station 274m.

37. Big fluorescence signal >16mg/m3

40. Fluorescence max ~12

44. Fluor >20

46. Fluor spiking to 27

48. Altimeter dropping out close to bottom.

49. PAR cap was left on the rosette; brought back to surface to remove it.

53. SPAR coefficients entered in configuration file.

64. Chains on – SPAR seems better

67. Windy during sampling.

72. Altimeter drops out in last 5-7m from bottom.

82. Tidal current

83. Go pro test - 0.5m/sec - at station 20

84. “Real” cast at station 20

85. Go pro test #2 1m/sec at station 22

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 31Dec2013 | Factory |  |  |
| Conductivity | 1766 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 4700 | 4Jan2014 | Factory |  |  |
| Secondary Cond. | 3321 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 21Jan2014 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | 14Jan2014 |  |  |  |
| WetLabs ECO Fluor. | 2216 | n/a |  |  |  |
| Pressure Sensor | 77511 | 30Dec2013 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |