## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |
| 19 Nov. 2014 | Note added to summary section on post-cruise sensor calibrations. GG |

## PROCESSING NOTES

Cruise: 2014-15

Agency: IOS, Ocean Sciences Division, Sidney, B.C.

Location: Douglas Channel

Project: WC-Douglas channel

Party Chief: Vagle S.

Platform: Vector

Date: April 12, 2014 – April 19, 2014

Processed by: Germaine Gatien

Date of Processing: 27 June 2014 – 28 July 2014

Number of original HEX files: 37

Number of CTD files: 37 (2 from upcast data)

Number of CHE files: 15 (3 other casts with bottles fired but no samples or sample #s)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1119), an ECO Fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#4615), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (#1204).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order. No cruise report was available at the time of processing. There were no photos of the setup, but it is believed to be the same as that used for 2014-06.

There were a number of errors in sample numbers and/or event numbers. In one case the rosette sheet had obviously been changed, but the labels were not corrected.

The files for events #16 and #70 contain data from the upcast because no downcast data were available. These data were edited to remove records corrupted by shed wakes when the rosette was stopped suddenly for bottle sampling. The quality of the data for those 2 files is considered lower than usual because the rosette package mixes water ahead of the CTD and carries deeper water with it. Typically there is an offset of from 2 to 10m between the downcast and upcast temperature traces.

CTD #0585 was used for 6 successive cruises in early 2014; this was the 4th in the series. It was noted during the first of these cruises that the two salinity channels differed by about 0.004 even though pressure, temperature and conductivity sensors had all been serviced recently. That difference grew slightly through 4 months of use. Examination of an offshore cruise (2014-21) with a noisy CTD descent rate and quick analysis of salinity samples shows the primary salinity channel to be close to bottles. This result is trusted more than those of the inshore cruises where a combination of poor flushing of Niskin bottles (due to very steady descent rates) and shallow sampling (with associated higher salinity gradients) can lead to the bottle values being significantly higher than ambient conditions. An exception is for bottles fired near the bottom where flushing has the opposite effect and bottle values are lower than the ambient ones. For the other offshore cast (2014-18), the primary was closest to the bottles, but not as close as for 2014-21, likely due to evaporation of samples as the analysis was not done as quickly as for 2014-21.

NOTE: 19 November 2014: Post-cruise factory calibrations show that the secondary salinity was high by about 0.005 by mid-August. Based on observations during a series of cruises, it appears that the calibration applied to these data in July was appropriate, though there remains some doubt about when drift occurred. G. Gatien

The salinity analysis was completed within 13 days of the beginning of the cruise. Salinity samples were frequently taken from the bottom bottle. This is not a good practice if the intention is to use the samples for calibration purposes, as even in the best sampling conditions, flushing of Niskin bottles is poor near the bottom. For this cruise sampling conditions were far from ideal as the descent rate was generally very quiet and local gradients fairly high. The effects of poor flushing are obvious in the comparison of CTD with bottles as the CTD salinity looks much lower than bottles near the surface and gradually grows closer to the bottles with increasing depth. For the near-bottom bottles the CTD salinity is higher than the bottles.

The correction applied to the Oxygen Dissolved:SBE data amounts to an increase of ~5% near the top of the range and ~8% for the lowest values. The fit of the CTD DO against titrated samples had a much larger offset than usual, but the same was true for 2014-06 in the Strait of Georgia. An initial assessment of the sensor’s performance during the offshore cruise that followed this one found a slightly larger slope for the correction, but a small offset, so the larger offsets found in the earlier cruises are not associated with the sensor itself.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 50db

±0.15 mL/L from 50 to 200db

±0.05 mL/L below 200db

The ratio of fluorescence to extracted chlorophyll is generally >1 for CHL<2ug/L and then decreases to about 0.65 at 4ug/L, and stays close to that for higher values.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets and final spreadsheets for dissolved oxygen, nutrients, extracted chlorophyll and salinity.

There is mention in the log of several casts with no downcast data acquired, some cases of sampling for non-chemistry items so no sample #s were assigned and noisy altimetry.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments and the only errors found were in some dates. One corrected file was saved as 2014-15-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2014-15-ctd.xmlcon to create CNV files.

File 2014-15-0040.cnv was renamed as 2014-15-0038.cnv since the log book mentions it was saved with the wrong name. There was no bottle file for that cast.

A few casts were examined and all expected channels are present.

The temperature and conductivity channels are close on downcasts, but the upcast traces are noisy with many excursions in both channels.

The altimetry looks extremely noisy at the bottom, but the trend above may enable an estimate if the header reading algorithm fails. Similar noise was noted during 2014-27.

The fluorescence, transmissivity, DO, pH and PAR look normal.

The SPAR signal looks odd. Before running REMOVE this should be checked.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2014-15-ctd.xmlcon.

The files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files. There was some noise in a few points of the primary salinity for cast #69 and the secondary salinity for cast #70. Those points were cleaned and the output file was copied to BOT.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

Events #48, 51 and 73 had bottles fired but no sample #s were assigned and no chemistry sampling done.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2014-15-bot-hdr.txt; it may need further editing to reflect problems found during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-15chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2014-15chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-15oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2014-15oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2014-15SAL.xls. The analysis was done within 45-51 days of collection. The files were simplified and saved as 2014-15sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-15nuts.xls. This includes a precision study.

Then the file was simplified, reordered on sample numbers and saved as 2014-15-nuts.csv. The file was converted to individual NUT files.

The CHL, SAL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions and a number of errors were found. In some cases the samples were given a number that was different from the rosette log sheet entry and in one case the event numbers differed. These errors led to missing data in the merge process.

- Event 22 – Event number for all samples was given as 23; this was also entered on the rosette log sheet, but 23 was not a CTD cast. The entries were changed to 22.

- Event 29 – DO, Nuts and CHL samples given as 122 were changed to 123 and salinity sample given as 123 was changed to 124 for salinity. There are erasures on the rosette log sheet, so the problem likely arose from changes after labels were prepared.

- Event 69 – one nutrient sample was entered as event 70 – this is presumed to be a typo. The nutrient file was reconverted after this correction.

- Event 69 has no sampling from Niskin #3 so this was removed from the SAMAVG file.

After these corrections, the merge process was repeated and when exported to a spreadsheet the results looked fine.

##### Compare

Fluorescence

COMPARE was run with pressure as reference channel. The fit of fluorescence versus extracted CHL looks typical with fluorescence generally reading higher than extracted CHL when the latter is lower than 2ug/L. Then fluorescence drops relative to CHL until the ratio FL/CHL reaches about 0.65 when CHL is ~4ugL. It stays fairly steady for higher values.

For full details see document 2014-15-fl-chl-comp.xlsx.

Salinity

Compare was run with pressure as reference channel. There are many outliers, with the most severe being near-surface bottles with high standard deviations in the CTD salinity. There are also a few outliers that come from samples taken at the bottom of casts; those do not have noisy CTD salinity. Salinity samples from bottles fired within 10m of the bottom are not usually included in COMPARE because they are routinely out of line in such a way that the CTD salinity looks higher relative to bottles compared to other samples. This is presumably because of flushing problems so close to bottom. So the outliers are not flagged since they probably do reflect what is contained in the Niskin bottle; they just don’t correspond to what the CTD is measuring.

When the bottom bottles and all differences >0.008 are excluded, the primary salinity is found to be low by an average of 0.0032 and the secondary salinity is high by an average of 0.0007. While the fits look flat with pressure, the standard deviation of 0.003 in both fits makes this less than convincing. The difference between the primary and secondary salinity is ~0.004 which fits the observations of previous cruises using these sensors.

When only bottles from within 5m of the seabed are included, the primary salinity is high by an average of 0.0047 and the secondary is high by 0.0094. The shallowest of those bottles is most out of line, probably due to the local gradients being higher. But even if that one is excluded, the primary is low by an average of 0.0004 and the primary is high by 0.0042. So the comparisons using bottom bottles only show the CTD to be higher by about 0.004 or 0.008 than the comparisons excluding bottles and other outliers.

None of the outliers suggest a need for quality flags as the standard deviation in the CTD salinity is high for those that are not bottom bottles.

For full details for the COMPARE run see document 2014-15-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When a fit is forced through the origin as SeaBird recommend, it looks ok in the surface layer, but is very unsatisfactory below that. Allowing an offset produces the fit:

CTD DO Corrected = CTD DO * 1.0325 + 0.1019 (**)

Using a variety of criteria to remove outliers had little effect on the fit.

Given that the bottom bottles were found to be out of line for salinity and near-bottom DO bottles from 2014-27 showed up as outliers, a fit was done excluding those:

CTD DO Corrected = CTD DO * 1.0363 + 0.0850

The bottom bottles with the largest differences were removed from the first fit as residuals anyway, so the difference between the two fits is not large. The combination of slope and offset correction amounts to a correction of 5% increase for DO~6mL/L in each case. The corrections are larger for small DO.

To see if the fit from 2015-06 works well for these data, the offset was forced to be -0.0744 and the fit was:

CTD DO Corrected = CTD DO * 1.0375 + 0.0744

Again the correction is ~5% at 6mL/L. The slope is much higher than the 1.0300 found for the 2014-06 data. Removing more outliers does not reduce the slope significantly. There is no obvious justification for choosing the 2014-06 offset since the slopes are so different.

A preliminary comparison of 2014-21 bottles versus CTD DO looked quite different from those of the earlier cruises, having a higher slope but much smaller offset, and slightly lower % corrections. That was an offshore cruise with noisy descent rate and lots of sampling and the results look more typical of what we expect from this type of sensor.  So the higher offsets seen in 2014-06 and 2014-15 do not appear to be due to changes in the sensor calibration. For both the inshore cruises the descent rate was generally very steady.

It seems likely that poor flushing of Niskin bottles is a factor, though why that should lead to a large offset is not obvious. The CTD DO is always lower than bottles, but if flushing is poor then the bottles would generally have lower DO values than usual, so we would expect the correction should be smaller. However, the response of the CTD may also differ in these conditions. We have no data with DO<2mL/L so we have no evidence of how the sensor would behave at low DO. Normally, given these problems a zero offset would be used, but such a fit looks totally unsatisfactory. DO is not as simple a variable to deal with as salinity since the error is DO dependent. With no evidence to the contrary, and no obvious reason to pick the slope from 2014-06, it appears that using the first fit (**) is the best choice. No fit is perfect and a second comparison will be run using downcast data, so we will get another look at this.

A study was made of Event #1 because it is almost completely mixed, with an average CTD DO value of 5.299mL/L and a standard deviation of 0.004mL/L. The average difference between bottles and CTD is 0.262mL/L with a standard deviation of 0.013mL/L. The precision estimate for this cruise was Sp=0.008 with the largest differences between replicates being ~0.002mL/L. The average correction to make CTD DO match bottles is ~5%. That is in line with the cumulative effect of slope and offset correction when using all the data (5.2% at 5mL/L). The correction varies from 4.5% to 5.5% if we use just one of the bottle/CTD pairs to estimate the correction. A few of the bottles were flagged because of bubbles, but removing them from the calculation does not change the range of corrections from ±0.5%. This does illustrate that bottles differ more than the CTD. A check of a few plots of DO during bottle stops suggests that this is due to poor response of the CTD, not problems with sampling or flushing of Niskin bottles.

There are no significant outliers, so no changes to flags are recommended.

For more detail see document 2014-15-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined and no outliers were found.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

From the results of cruises 2014-03, 2014-06 and 2014-27 which all used the same sensors, it appears that the dissolved oxygen channel needs to be advanced by +2.5s.

ALIGNCTD was run using a setting of +2.5s to the dissolved oxygen channel.

Dissolved oxygen concentration was derived and then plots were examined to see how well that setting worked, and the results are very good.

##### CELLTM

From the results of cruises 2014-03, 2014-06 and 2014-27 which all used the same sensors, the best CELLTM settings were found to be (α = 0.02, β=9) for the primary conductivity and (α = 0.0245, β=9.5) for the secondary conductivity. The tests were not easy to interpret, but this appears to be the best we can hope to achieve.

CELLTM was run using (α = 0.02, β=9) and (α = 0.0245, β=9.5) for the primary and secondary conductivity. After salinity was derived, plots were examined and the results do improve the data in the few sections where both up and downcast traces are sufficiently smooth to expect to see some improvement.

##### DERIVE  and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

Program DERIVE was run a second time on a selection of deeper casts to study the differences between channel pairs.

Differences from a cast from 2014-06 and one from 2014-27 are included for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-06-0043 | 300 | +0.0002 | +0.00042 | +0.0042 | Very steady |
| 2014-27-0073 | 335 | +0.0002 | +0.00045 | +0.0045 | Very steady |
| 2014-15-0032 | 400 | -0.00004 | +0.00040 | +0.0049 | Very steady |
| 2014-15-0039 | 400 | +0.00004 | +0.00045 | +0.0050 | Very steady |
| 2014-15-0071 | 400 | -0.0002 XN | +0.00040 VN | +0.0050 | Very steady |

So the differences are similar to those found for 2014-06 and 2014-27 though salinity differences are somewhat larger. These are rough estimates as the data are noisy.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. No problems were found. However, the files for events #16 and #70 have no downcast data. An attempt will be made to create files based on upcast data, but those files are likely to be heavily corrupted by stops for bottles. The station name for event #19 was changed from D20 to Doug20.

A header check was run. There are a few values that are odd, such as a negative fluorescence value, but these are likely due to near-surface spikes. No other problems were noted.

The cruise track was plotted with event #s and station names and both were added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.1db which is reasonable for inshore sampling from the Vector. One cast was examined that had a negative minimum pressure and the data show that the CTD was very close to the surface when pressure was ~0, with cases of zero fluorescence and transmissivity going suddenly from very low values to very high ones. Conductivity shows that one sensor was out of water and the other still in water. These observations suggest that the surface pressure is within 0.2db.

The altimeter and water depth readings from the headers of the CLN and BOT files were exported to a spreadsheet. Depth entries were checked against log entries. Entries were checked by noting whether the header depth minus altimetry was close to the maximum pressure (after making rough correction to get depth). Where they did not match well, or if header depths did not match log entries, altimetry was plotted. In a few cases the log depths looked more reasonable so those were selected and in a few cases the altimeter header entry looks wrong.  For 3 bottle files there was no altimeter header due to spiking, but an estimate was possible from the plots, so those were entered.

The steps between BOT and MRGCLN2 were rerun to capture these changes.

##### Shift

Fluorescence

SHIFT was run on the WetLabs ECO fluorescence channel in all casts using the usual advance of +12 records. Examination of plots after this step were hard to interpret due to high variability in both temperature and fluorescence and stops for bottles, but there appears to be slight improvement in matching the fluorescence offset to the temperature offset.

pH

As was found for fluorescence it is difficult to test these data to choose the best setting for aligning pH. They clearly need aligning and using the setting found best during 2014-06 and 2014-27 definitely improves the alignment with temperature.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but the alignment seems ok.

Conductivity

Tests were run on 3 casts using a variety of shifts and the best results were found with the settings used for 2014-06 and 2014-27.

SHIFT was run using -0.7 records for the primary conductivity and -0.6 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

Files #16 and 20 were put through REVERSE and then DELETE to produce upcast files since there were no data acquired during the downcasts. These will be examined to see if they are reasonable substitutes.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

This was the 4rd cruise using these temperature and conductivity sensors since they were calibrated at the factory. For 2 others there was salinity calibration sampling that suggested the secondary salinity was close to bottle salinity, but there were problems with the comparison. A preliminary comparison of 2 offshore cruises that followed showed the primary salinity to be closest to bottles. Based on those comparisons the secondary salinity was recalibrated for cruises 2014-03, 2014-06, 2014-27 by subtracting the difference between the 2 salinity channels, 0.004, 0.0041 and 0.0042 respectively.

Dissolved Oxygen

The sensor was calibrated at the factory in January 2013 but not used until 2014. There were 3 uses before this cruise. There was limited sampling during 2014-03 and 2014-27 with likely problems due to poor flushing of Niskin bottles. There was good sampling during 2014-06 but that cruise also had some problems with flushing.

Pressure

The sensor was recalibrated in January 2014 and it was used for 3 other cruises before this one. The factory offset was used for those cruises and the results looked reasonable.

Historic ranges – Local climatology was only available for one cast, and that one was close to shore where the 3-standard deviation climatology ranges are not appropriate. Plots with the climatology superimposed showed that the temperature was within the historic range, while salinity was a little low at the bottom of the cast, but otherwise within the range.

Repeat Casts – The only repeat casts serve only to demonstrate the variability in the region, so they are not suitable for testing repeatability of CTD data.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### CHANNEL CHOICE STUDY

An examination of T-S plots using the two channel pairs shows little difference, though the primary had a few more spikes and unstable features. Fine-tuning of the primary alignment did not help - for every feature that could be improved, another looked worse. Based on the results of cruise 2014-21 the primary salinity is believed to be more accurate than the secondary. But the secondary channels were selected since they have fewer instable features, and will be easier to edit. Recalibration will be applied later to bring the secondary into line with the primary salinity values.

##### DETAILED EDITING

The secondary temperature and salinity channels were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some surface records were removed because pumps had not yet been turned on.

All files required light editing.

All output files were copied to EDT.

NOTE: Upcast files 16 and 70 were edited heavily to remove the effect of shed wakes during of stops for bottles. Special note was made in the headers that these data are of lower quality than usual due to the rosette carrying deeper water with it as it rises.

##### Initial Recalibration

The pressure does not appear to require recalibration.

For the 3 cruises which preceded this one the secondary salinity was recalibrated by subtracting 0.0040, 0.0041 and 0.0042 in order to make the secondary match the primary salinity. For this cruise the difference between the two channels is about 0.0039 during bottle stops, and roughly 0.0045 to 0.0048 during downcasts after alignment. During the cruise that followed, the preliminary estimate is that the difference between channels during bottle stops was ~0.0054, so there does appear to be some increase with time. An estimate of 0.0045 was chosen for recalibration for recalibration of these data. While the difference during bottle stops likely relates more accurately to calibration drift, we want the two channels to match through the profiles. Smaller differences when the CTD is stopped may be due to the absence of the minor eddies affecting the flow when the package is in motion. The effect of such eddies are apt to be systematic but noisy.

CALIBRATE was run using file 2014-15-recal1.ccf to subtract 0.0045 from the Salinity:T1:C1 channel and to apply the following correction to Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files:

CTD DO Corrected = CTD DO * 1.0325 + 0.1019

COMPARE was rerun for salinity and dissolved oxygen using the recalibrated values. The results confirm that the recalibrations were done properly.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles – there were no downcast CTD data for event #16. The differences show that the downcast CTD DO is higher than the bottles by an average of about 0.037mL/L, but removal a single points makes a significant difference in the result. Given that the errors due to flushing of bottles would lead to the CTD DO looking higher, there appears to be no justification for further recalibration.

See 2014-15-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots show some unstable features near the top and/or bottom of casts, but this is not unexpected in this region. No obvious instrumental cause for the instabilities was found, so they were left unchanged.

Profiles mostly look fine though there are a few deep spikes in transmissivity and fluorescence. No editing was applied as these features are at sites with fairly strong mixing.

The PAR and Surface PAR values are very different. Checking against local time, the PAR values look reasonable, but the Surface PAR is sometimes very high and sometimes there is no signal at all. There are a few casts for which it seems possible that the SPAR signal is ok with low, noisy values when PAR is also low, but overall, this channel does not look reliable and will be removed from all casts except #1 which looked ok at least when the CTD was near the surface.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

For all casts except #1 PAR:Reference was also removed.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

Channel PAR:Reference PAR was removed from all files except for event #1 due

to the absence of a signal or unbelievable values.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 50db

±0.15 mL/L from 50 to 200db

±0.05 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2014-15_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Surface saturation values <85% were found in Johnstone Strait, an area with strong vertical mixing. Most values were between 95% and 105%, but there were 8 cases where saturation was between 105% and 110% and 1 case with saturation >110% (station DOUG 4 at ~120%). The DOUG4 casts shows a fairly high DO gradient at the surface. These values do not suggest any problem with calibration.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

Reference:PAR was removed from all casts except #1.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check was run on all files and no errors were found.

A header check was run on the final files.

No further errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

1.  2.44db at surf

16. Downcast not archived. Many bottle stops on upcast so not good substitute for profile.

22. Labels say event 23, not 22

22. Forgot to trip at 50m so went back down to get it.

29. Some bottle firings not used for analysis but still have sample #s.

34. Lots of zooplankton at surface

38. Hex file named as event 40 should be 38

47. Not archived

48. Surface bottles for microplastics, not assigned sample #s.

51. Bottles fired but no sample #s applied

56. Recovered to remove PAR sensor cap – pump was turned off then on

65. Strong current

69. Altimeter noisy

70. Forgot to start archiving – missed downcast. Many bottle stops on upcast so not good substitute for profile.

73. Fired one bottle for phyto sample, no sample #

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 31Dec2013 | Factory |  |  |
| Conductivity | 1766 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 4700 | 4Jan2014 | Factory |  |  |
| Secondary Cond. | 3321 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 21Jan2014 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | 14Jan2014 |  |  |  |
| WetLabs ECO Fluor. | 2216 | n/a |  |  |  |
| Pressure Sensor | 77511 | 30Dec2013 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |