## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| Jan. 2019 | Bottle spreadsheet converted to searchable BOT files. |

## PROCESSING NOTES

Cruise: 2014-16

Agency: SAFE

Location: Strait of Georgia & Johnstone Strait

Project: High Seas Salmon

Party Chief: Morris J. / Thiess M.

Platform: W.E. Ricker

Date: 6 March 2014 – 22 March 2014

Processed by: Germaine Gatien

Date of Processing: 28 July 2014 – 1 August 2014

Number of original HEX files:	59	Number of CTD files: 58

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1176) on the primary pump, an WET Labs ECO_AFL Fluorometer (#2215) and an SBE18  pH sensor (#0851).

The deck unit was a Seabird model 11.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book lacked a personnel list but was otherwise in good order.

Salinity calibration sampling was done, but the comparison with CTD salinity is not reliable. The depth of firing is uncertain, it is likely that the bottles did not flush completely and the CTD was on the move when it was at the sampling level. The upcast data are corrupted by the wake through which the CTD travels. The comparison with downcast data from well-mixed casts suggests that the CTD salinity is good to ±0.003.

The primary temperature and salinity channels were selected for one cast only; the primary salinity was recalibrated to bring it into line with the secondary.

The CTD ECO Fluorometer data was compared with the extracted chlorophyll samples and the usual pattern was found with the CTD fluorescence reading about 2*CHL for CHL=0.2ug/L, about 1*CHL for CHL=1ug/L and about 0.7*CHL when CHL=6ug/L.

There was no dissolved oxygen sampling so the results of a previous cruise were used to recalibrate that channel. The surface oxygen saturation values are fairly low, but this is likely due primarily to strong vertical mixing rather than calibration problems.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained. There were notes about problems including data lost from the first 4 casts and a problem at the end of one other cast.

The same configuration file was used throughout the cruise.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity, fluorometer and DO sensors were obtained.

The extracted CHL, nutrient and salinity analysis spreadsheets were obtained and the data from each were combined in spreadsheet 2014-16-bottles.xlsx.

The calibration constants were checked for all instruments. The only change made was to the dissolved oxygen factor E to reflect tests run using 2014-01 data to fine-tune the hysteresis setting. This will not affect this cruise since the hysteresis correction will not be needed, but it was changed to avoid confusion in the future.

##### Conversion of Full Files from Raw Data

All files were converted using con file 2014-16-ctd.xmlcon with hysteresis correction turned off.

A few casts were examined and all expected channels are present. File #1 is not from event #1 and was a test at JF04 – there is only surface data in the file, so it will not be processed further.

The two temperature and conductivity channels are fairly close to each other during the downcasts but are a little noisier in the upcasts. In the downcasts the secondary temperature is smoother but in upcasts there are odd excursions in both channels.

The fluorometer, pH and dissolved oxygen traces look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A setting of +4s was used during 2014-01 to advance the DO voltage. Checks of the offset between distinctive features in T and DO suggest that +4s is right for this cruise too.

ALIGNCTD was used to advance the DO Voltage by 4s relative to the pressure.

Plots were examined before and after that step and the results with +4s look good.

##### CELLTM

Tests were run on 5 casts to see what settings are best for this step. The temperature gradients were fairly low for most casts, so there is little difference among the various choices.  For the 2 casts with a higher temperature gradient the choice of (α = 0.03, β=9) looked best.

CELLTM was run using (α = 0.03, β=9) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

When these sensors were used during 2014-01 there was so much noise in the data that no reasonable comparison could be made between channels. For this cruise the casts are all shallow making the comparison less reliable, but two casts were examined and the differences were:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-16-0224 | 240 | +0.0003 | +0.00013 | +0.0011 | Fairly steady |
| 2014-16-0087 | 240 | +0.0004 | +0.0002 | +0.0017 | Fairly steady |

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  No problems were noted.

The surface check gave an average value of 3.7db which is not out of line for the Ricker. There were a few slightly negative pressures with in-water salinity values, but they were isolated so may be spikes or a case of the CTD being very briefly out of water. Cast #13 starts at about 0.8db; salinity starts with low values which then rise sharply when pressure reaches 0.9db with pumps turned off. The pressure seems reasonable.

The cross-reference check was compared with the log book and two station names were edited in the CLN files just to make the format consistent with other entries.

The cruise track was plotted and added to the end of this report. No problems were found.

##### Shift

Fluorescence

A shift of +12 is the usual setting used to align fluorescence with temperature. Tests were run on 10 casts and the results look generally ok, though it is hard to tell for many casts because fluorescence variations are very small.

SHIFT was run on all casts to advance the ECO fluorometer by +12 records.

Conductivity

Tests were run on a few casts using a variety of settings to see which shift produces reasonably stable T-S plots. The bests result for the primary salinity was the same as for cruise 2014-16, -0.7 records. But the secondary salinity was much different at -0.9 records, whereas it was best with +0.7 records for 2014-16.

Two runs of SHIFT were used to apply advance the primary by -0.7 records and the secondary by -0.9 records.

Dissolve Oxygen

Checks were made of the alignment of DO and temperature and the shift applied earlier appears to have done a good job though, as usual, for some features it was too little or too much.

pH

The pH trace appears to be quite close to temperature. Because of hysteresis in the pH signal, tests are best run on casts with distinctive features in the pH traces. This was done by finding the distance between such features in the upcast and downcast and comparing that with the temperature offset after applying a variety of shifts. The setting of +30 records produced good results.

SHIFT was run on all casts with the setting +30 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds (taken from header)

COMMENTS ON WARNINGS: There were warnings for 3 casts. There were some odd pressure jumps with no associated temperature variability. In one cast the problems were in the upcast and it was a cast that encountered problems with the winch at the end; this is of no importance for these files. For the other 2 the pressure has some odd jumps around 4db, but the data selected by DELETE look ok.

The SHFpH files were put through REVERSE and then DELETE to produce upcast files to be used in the comparison with samples.

##### Comparisons of CTD and bottle samples

Judging by plots of pressure versus scan numbers the CTD was stopped between 13.7 and 16.6db, with a median value of about 14.7db. It would be very time-consuming to select data from the right depth for each cast, and given the conditions, there is likely incomplete flushing anyway. That problem may be reduced a little by the consistent stopping of the CTD a little above the sampling level followed by sinking to the sampling level. But this comparison cannot be assumed to be more than a guide. Choosing a median value of 9.75db to match the samples is probably the best we can do.

The DEL files were bin-averaged (0.25m bins) and then thinned to values of 9.75db. The data were exported to 2 spreadsheets 2014-16-downcast-ctd.csv and 2014-16-upcast-ctd.csv.

Salinity

The data from the spreadsheets were combined with the bottle salinity data in file 2014-16-sal-comp.xlsx. The differences were calculated between the 2 CTD salinity channels and bottles for upcast and downcast. Where they were >0.1 for at least 3 of the comparisons, the data were removed. All differences>0.1 were highlighted so they would not be included in averages.

Before looking at how the CTD data compared with bottles, the differences between the salinity channels were calculated. When the differences >0.1 were excluded the median values show that the secondary salinity was higher than the primary by ~0.001 for downcasts and by 0.004 for upcasts. The average values are similar.

Four casts were identified that were well-mixed to 10m and for those the two channels were very close for both upcast and downcast, with differences of 0.0005 and 0.0006. During 2014-01 it was found that the primary and secondary salinity differed by more than that overall, but at the surface the differences were about 0.0014.

Looking at the comparison with bottles, the 4 well-mixed casts show the primary CTD reading low by an average of 0.0027 and the secondary low by 0.0023 compared to downcast CTD salinity at 9.75db. Compared to the upcast data from that depth, the CTD primary is lower than bottles by 0.0027 and the secondary is lower by an average of 0.0020. This is good agreement, but we don’t know if the difference is a measure of calibration drift or the result of bottles not flushing well so that the CTD looks lower than it really is. Any flushing error should be small for well-mixed casts, so this may be about right. It is also possible that the depth picked for the comparison is too shallow, but this should not have a large effect for well-mixed casts. The other possibility is that there was some evaporation of samples. The analysis was done fairly quickly, so this should not affect the samples much.

Looking at all the casts excluding differences >0.1, the results are quite different, with the primary CTD salinity reading lower than bottles by a median of 0.0005 for the downcast while the secondary is higher by 0.0007. The upcast salinity channels are both higher than bottles by a median 0.0122 and 0.0163. The downcast results differ from those of well-mixed casts by about 0.002, which given the noise in the comparison may not be significant. But the upcast results differ from the well-mixed by 0.01 and 0.014. Neither poor flushing nor evaporation of samples can explain the CTD reading higher than bottles. However, the CTD data come a time when it is on the move and passing through a wake created by the wire, Niskin bottle and the CTD itself, and it would have started moving just 5m below creating a disturbance of the water as it accelerated, so it is likely measuring higher salinity than if it were stopped a while. In larger salinity gradients, that difference would be much more significant than for the well-mixed casts. For the downcasts the CTD sensors move into undisturbed waters and so the data are more reliable. Plots of differences show more scatter in the upcast comparison than the downcast. There are more outliers in the differences for the upcasts and they are mostly towards the CTD looking too high. There are fewer outliers in the downcast comparison and they are in the opposite direction, which is consistent with the CTD carrying water with it during the upcasts, and perhaps to a lesser extent during the downcasts.

The 2 CTD salinity channels are close to each other and are likely within 0.003 with the secondary reading higher than the primary by roughly 0.001.

During 2014-01 when flushing was not likely to be a significant issue and there were stops for bottles, these sensors were found to be low by 0.0037 and 0.001. They were likely within 0.0015 of each other near the surface as there appeared to be some pressure dependence. There was some delay in analysis, so evaporation of samples likely occurred and a very rough estimate of that effect is 0.002, which would mean both sensors are both within 0.002.

Fluorescence

A similar comparison was done for fluorescence and extracted chlorophyll. When plots were made of the ratio of fluorescence / CHL versus CHL, the usual pattern emerged for both upcast and downcast. The ratio is about 2 for very low CHL, reaches about 1.5 at 0.5ug/L, ~1 at 1ug/L and falls off to about 0.7 at 6ug/L. Reducing the plots to only casts which had low fluorescence gradients in the top 10m made no significant difference, though only 6 casts were included.  The upcast fluorescence is slightly lower than the downcast, which is presumed to be due to the CTD moving through wakes carrying deeper water.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity and temperature sensors were both recalibrated in late January 2013 and they were used for 1 cruise since then, 2014-01. Flushing was not considered an issue for that cruise and there were stops for bottles. The primary salinity was found to be low by 0.0037 and the secondary low by 0.001. Delay in analysis of samples and problems with seals on many bottles likely lead to some evaporation, so both CTD channels were likely within 0.002. There appeared to be some pressure dependence in the primary salinity. The two salinity channels appeared to be within 0.0015 near the surface.

Dissolved Oxygen

The DO sensor was recalibrated in February 2013and this is the second use since then. From 2014-01 data hysteresis tests were done and factor E was fine-tuned.

Pressure

The pressure sensor was recalibrated in May 2013 and this is the second use since then. The factory offset was found appropriate in February 2014.

Historic ranges –All salinity data fell within the local climatology (not available for all casts.) The temperature data were mostly within the climatology but there were patches of slightly low temperatures between 40 and 70db at some casts in the southern Strait of Georgia and near Texada Island. Bottom temperatures were occasionally low in those same areas. These look like real variations. The 3-standard deviation climatology is not reliable in near-shore areas, and is just used to look for systematic excursions that might suggest problems with sensor calibration or performance.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

The first issue is to decide which sensor pair to edit. The secondary salinity might be slightly better, but there is very little difference between the noise level in the 2 channels. The secondary was chosen for most casts during 2014-01, so it was picked for this cruise too.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records.

Cast #144 had poor secondary salinity near the surface – it looks like steady flow had not been established yet. So the primary channels were edited for that cast.

All casts required some editing.

All ED1 files were copied to EDT.

##### Initial Recalibration

The pressure does not need recalibration.

The comparisons with bottles both for this cruise and for 2014-01 were subject to problems with either evaporation of samples or poor flushing of Niskin bottles. It appears that both salinity channels are providing reasonably good salinity values, so no recalibration will be applied to the secondary, but the primary salinity was recalibrated by subtracting 0.0012 to bring it roughly into line with the secondary, based on differences between the two channels during cast #144, the only cast for which the primary channels are to be archived.

The only information available about the SBE dissolved oxygen sensor since it was last recalibrated comes from 2014-01, so the correction applied to that data set will also be used for this cruise:

DOX_CTD Corrected = DOX_BOT = 1.0502* DOX_CTD +0.0142

File 2014-16-recal1.ccf was prepared to apply the corrections to dissolved oxygen.

CALIBRATE was run on the EDT files.

##### Fluorometer Processing

Special files were prepared for Dr. Peña.

A median filter, fixed size=11, was applied to reduce spikiness in both fluorescence channels. A few casts were examined before and after this step and showed that the filter worked properly but there was little spikiness in the Wetlabs fluorescence signal, so the change is not large.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except for #144 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag

REMOVE was run on cast #144 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag

Profile plots were produced at this point to check for errors. No problems were found.

T-S plots were produced; there are some small unstable features that likely reflect real conditions.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence and pH data are nominal and unedited except that some

records were removed in editing temperature and salinity.

Dissolved oxygen was calibrated using the results of cruise 2014-01 when the

same sensor was used. There was no calibration sampling from this cruise.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.  Calibration is required

for each cast to get absolute values, although general trends within a cast are likely real.

The primary CTD salinity channel was used for event #144, so it was recalibrated

to match the secondary channel which was selected  for all other casts.

For details on the processing see the report: 2014-16_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Surface Saturation

Dissolved Oxygen saturation was derived and plotted. The surface saturations varied from about 75% to 105% with values <90% being associated with very well mixed oxygen profiles, mostly in inlets and narrow passages where tidal mixing is likely. The higher values are associated with larger near-surface DO gradients and CTD fluorescence >2ug/L. The cast nearest the mouth of Juan de Fuca Strait (JF06) had a value of ~105% which is about what is expected in offshore waters when the surface layer is not well mixed.

# Particulars

1-10. JF01-JF04 - Data lost – HDR and XMLCON files created, but no HEX files.

1. There is a file 2014-16-0001.hex but it was a test at JF04 – only surface data.

13. Reboot done before cast.

87. CTD turned off at 14.7m – hydraulic leak on winch. CTD pulled in by hand.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | No | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 31Jan2013 | Factory |  |  |
| Conductivity | 2128 | 29Jan2013 | Factory |  |  |
| Secondary Temp. | 2374 | 31Jan2013 | Factory |  |  |
| Secondary Cond. | 3396 | 30Jan2013 | Factory |  |  |
| SBE 43 DO sensor | 1176 | 5Mar2013 | Factory |  |  |
| WetLabs ECO-AFL/FL | 2215 |  |  |  |  |
| SBE18 pH | 0851 | 29May2012 | Factory |  |  |
| Pressure Sensor | 75636 | 25Mar2013 | Factory |  |  |