## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2014-17

Agency: SAFE

Location: Strait of Georgia

Project: SoG Juvenile Salmon

Party Chief: Neville C.

Platform: W.E. Ricker

Date: 24 June 2014 – 8 July 2014

Processed by: Germaine Gatien

Date of Processing: 30 July 2014 – 5 August 2014

Number of original HEX files:	43	Number of CTD files: 43

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1176) on the primary pump, an WET Labs ECO_AFL Fluorometer (#2215) and an SBE18  pH sensor (#0851).

The deck unit was a Seabird model 11.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book was in good order with clear notes on problems encountered.

There was no calibration sampling, so the results of cruises 2014-01 and 2014-16 were applied since the same sensors were used.

The salinity comparisons from cruises 2014-01 and 2014-16 were both subject to some problems limiting their reliability, but it appeared that salinity was likely good to ±0003. Dissolved Oxygen calibration was only available from 2014-01and surface saturation values were very close to 100% in offshore waters, as expected. From this cruise the dissolved oxygen surface saturation values were highly variable, but this does not suggest bad calibration given the variety of areas sampled and the season.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained. There were notes about problems encountered.

No changes were made to the configuration file through the cruise.

The cruise summary sheet was completed.

The history of the pressure, conductivity, fluorometer and DO sensors were obtained.

The calibration constants were checked for all instruments. The only change made was to the dissolved oxygen factor E to reflect tests run using 2014-01 data to fine-tune the hysteresis setting. This will not affect this cruise since the hysteresis correction will not be needed, but it was changed to avoid confusion in future.

##### Conversion of Full Files from Raw Data

All files were converted using con file 2014-17-ctd.xmlcon with hysteresis correction turned off.

A few casts were examined and all expected channels are present. As usual the two temperature and conductivity channels are fairly close to each other during the downcasts but further apart in the upcasts.

The fluorometer, pH and dissolved oxygen traces look normal. The descent rate is fairly steady and high.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A setting of +4s was used during 2014-01 and 2014-16 to advance the DO voltage. Checks of the offset between distinctive features in T and DO suggest that +4s is right for this cruise too.

ALIGNCTD was used to advance the DO Voltage by 4s relative to the pressure.

Plots were examined before and after that step and the results with +4s look good.

##### CELLTM

Tests were run on 5 casts to see what settings were best for this step. The best choice was the same as that chosen for 2014-01.

CELLTM was run using (α = 0.03, β=9) for the primary and (α = 0.03, β=7) secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

When these sensors were used during 2014-01 there was so much noise in the data that no reasonable comparison could be made between channels. During 2014-16 it was steadier but all casts were <300db where differences are usually very noisy. For this cruise the casts are all shallow and differences were very noisy. Below are 2 samples from each of 2014-16 and 2014-17. While signs change all the differences are small.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-16-0224 | 240 | +0.0003 | +0.00013 | +0.0011 | Fairly steady |
| 2014-16-0087 | 240 | +0.0004 | +0.0002 | +0.0017 | Fairly steady |
| 2014-17-0155 | 240 | -0.0001 | +0.0004 | +0.0006 | High, F noisy |
| 2014-17-0170 | 240 | -0.0002 | -0.00005 | -0.0004 | High, F noisy |

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  No problems were noted.

The surface check gave an average value of 2.9db.

One cast had a few negative pressure values near the end of the cast. Those values may have been spikes or very brief periods when the CTD was very slightly above the surface; the salinity and fluorescence values are very low so the pressure appears to be reasonably accurate.

The cross-reference check was compared with the log book and no errors were found.

The cruise track was plotted and added to the end of this report. No problems were found.

##### Shift

Fluorescence

A shift of +12 is the usual setting used to align fluorescence with temperature.

SHIFT was run on all casts to advance the ECO fluorometer by +12 records.

Plots were examined after that step and the results look satisfactory.

Conductivity

Tests were run on a few casts using a variety of settings to see which shift produces reasonably stable T-S plots. The bests result for the primary salinity -0.5 records, while the secondary salinity looked best with a -1.0 records shift.

Two runs of SHIFT were used to apply those settings; salinity was recalculated.

Dissolve Oxygen

Checks were made of the alignment of DO and temperature and the shift applied earlier appears to have done a good job, though, as usual, for some features it was too little or too much.

pH

The pH trace appears to be quite close to temperature. Because of hysteresis in the pH signal, tests are best run on casts with distinctive features in the pH traces. A setting of +30 records produced good results on 2014-16 data and looks appropriate for this cruise as well.

SHIFT was run on all casts with the setting +30 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds (taken from header)

COMMENTS ON WARNINGS: The only warnings pertained to surface records at the end of two upcasts, so are of no concern in preparing the CTD files.

##### Calibration checks

Previous experience with these sensors –

Salinity:

The conductivity and temperature sensors were both recalibrated in late January 2013 and since then they were used for 2 other cruises, 2014-01 and 2014-16. For 2014-01 the primary salinity was found to be low by 0.0037 and the secondary low by 0.001. Delay in analysis of samples and problems with seals on many bottles likely allowed a little evaporation,  so the primary is likely closer to the bottles than it appears, and the secondary is perhaps a little higher than the bottles, but still likely within 0.002. 2014-16 included salinity sampling but only single bottles at ~10m. The quality of the comparison is limited by uncertain depth of sampling, incomplete flushing of bottles and the fact that the CTD was moving when data were recorded at the level of sampling. The salinity was estimated to be ±0.003.

Dissolved Oxygen

The DO sensor was recalibrated in February 2013and was used for 2014-01 when there was good DO sampling and hysteresis tests were done and factor E was fine-tuned.  For 2014-16 there was no DO sampling and the results of 2014-01 were applied.

Pressure

The pressure sensor was recalibrated in May 2013 and has been used for 2 other cruises since then. The factory offset was found appropriate for those 2 cruises.

Historic ranges – The DEL files were bin-averaged and profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. For some stations there was no local climatology available. Where it was available all temperature data fell within the local climatology, while the only salinity excursions from the climatology were slightly low values in the bottom 30m at 1 cast in Juan de Fuca Strait. That cast was very close to shore and such excursions are not unexpected in this region.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

The first issue is to decide which sensor pair to edit. The secondary salinity was used for the previous 2 cruises on which this CTD was used. For this cruise the secondary looks smoother overall, though for some features the primary looks smoother.

Secondary temperature and salinity were chosen for editing and eventual archiving.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and records corrupted by shed wakes including some surface records.

The following casts required no editing: 2, 124, 158, 189 and 201.

All other casts required light editing.

All edited files were copied to *.EDT.

##### Initial Recalibration

As was done for previous uses of this equipment, the pressure and secondary salinity will not be recalibrated.

CALIBRATE was run on the EDT files using file 2014-17-recal1.ccf  to apply the following correction to dissolved oxygen:

DOX_CTD Corrected = DOX_BOT = 1.0502* DOX_CTD +0.0142

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to reduce spikiness in both fluorescence channels. A few casts were examined before and after this step and showed that the filter worked properly but had little effect on the Wetlabs fluorescence which was already quite smooth.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was considered appropriate.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag

Profile and T-S plots were produced at this point to check for errors. No problems were found.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence and pH data are nominal and unedited except that some records were

removed in editing temperature and salinity.

Dissolved oxygen was calibrated using the results of cruise 2014-01 when the

same sensor was used. There was no calibration sampling from this cruise.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2014-17_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Surface Saturation

Dissolved Oxygen saturation was derived and plotted. There is very high variability as expected in a cruise with much sampling in areas of tidal mixing and many casts very close to shore. The surface saturations varied from about 75% to 130% with the exception of 1 very high value of 200%. The values <80% had very well-mixed near-surface oxygen profiles. Values between 80% and 90% were in the Victoria - Haro Strait area. The values closest to 100% were in the central Strait of Georgia. The highest value was from near the mouth of Juan de Fuca Strait where there were very high gradients in DO and fluorescence in the top 6m, with fluorescence rising from 5 to 65ug/L in just 1m.The variability in this study does not provide useful evidence about the dissolved oxygen sensor recalibration. During 2014-01 when the same fluorometer was used, the offshore saturation rates were close to 100%, as expected.

# Particulars

75. CTD held at 65m for ~30s due to difficulty holding ship’s position.

93. Originally saved as Set53 – changed later, so time of hex file is later than time of acquisition.

# 114. Strong winds, white caps

# 136. Surface temperature highly variable.

# 155. Salinity spike ~160m downcast.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | No | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 31Jan2013 | Factory |  |  |
| Conductivity | 2128 | 29Jan2013 | Factory |  |  |
| Secondary Temp. | 2374 | 31Jan2013 | Factory |  |  |
| Secondary Cond. | 3396 | 30Jan2013 | Factory |  |  |
| SBE 43 DO sensor | 1176 | 5Mar2013 | Factory |  |  |
| WetLabs ECO-AFL/FL | 2215 |  |  |  |  |
| SBE18 pH | 0851 | 29May2012 | Factory |  |  |
| Pressure Sensor | 75636 | 25Mar2013 | Factory |  |  |