## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2014-20

Agency: OSD

Location: WCVI & Douglas Channel

Project: WCVI Moorings and Douglas Channel

Party Chief: Spear D.

Platform: John P. Tully

Date: 25 June 2014 – 6 July 2014

Processed by: Germaine Gatien

Date of Processing: 3 September 2014 – 24 September 2014

Number of original HEX files: 	23	Number of CTD files processed: 21 (2 casts surface only)

Number of bottle casts: 	22		Number of bottle casts processed: 14 (8 fired for water for BIO)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1119) on the primary pump, a SeaPoint Fluorometer (#3642) with a 3X cable on the secondary pump and an altimeter.

The data logging computer was #3.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had personnel and equipment lists, and it was clearly indicated that the TSG was not used.

This cruise was the last of 7 that used the same sensors between March and July 2014. While there was salinity calibration sampling from all but one of them, for 4 the comparisons with the CTD are noisy and considered unreliable due to problems with poor flushing of Niskin bottles and/or evaporation of samples due to delayed analysis. For 3 cruises there was a very steady CTD ascent rate and sampling was shallow and mostly in areas with a high vertical gradient of salinity. These factors make the errors due to poor flushing significant. Cruise 2014-21 provided the best comparison because analysis was very prompt and a noisy descent rate probably improved the flushing of Niskin bottles. This showed the primary salinity to be close to bottles and the secondary high by about 0.005

For this cruise analysis of salinity samples was prompt, but flushing was likely poor. The primary salinity was found to be low by 0.0044 and the secondary high by 0.0022, but the standard deviations are ~0.005. The primary channels were selected for archiving. Post-cruise calibration results were not yet available.

During this cruise there were problems near the surface with unstable features, sometimes in the primary channels but mainly in the secondary, with salinity appearing to “overreact” to temperature changes in the top 10 to 20m. Below 20m the secondary temperature and salinity look somewhat better than the primary but not enough better to justify the severe editing that would be needed at the surface if the secondary T/S data were selected. The cause of this problem is assumed to be irregular flow in the duct so that the alignment of temperature and conductivity is variable. The alignment correction applied is based on data observed at depth. The protocol of soaking the CTD at 10m before running a cast was not used, so bubbles may be causing the problem. Another possible cause might be problems with the secondary pump, though even the primary looks somewhat affected. One cast with severe problems was the first event in Saanich Inlet. Data from the 2 previous cruises which used the same equipment and also sampled Saanich Inlet as their first events show no such features. Surface data are often noisy but these problems are of a character that we normally only see if the pumps are not turned on.  The near-surface salinity data are considered of lower quality than usual.

The fit of Oxygen:Dissolved:SBE data to titrated samples produced some oddly large offsets for 2  cruises that used this equipment between March and April 2014. There followed 2 other cruises that had normal offsets and it was thought that this might be due to them being offshore cruises with less of a problem due to poor flushing of bottles. 2014-20 was also one that likely had poor flushing of bottles yet it does not have a large offset. This may be because flushing is better than for the 2 earlier cruises or there may have been some adjustment of the equipment that affected the data quality. Whatever the reason, the fit for these data looks more normal than for the other recent inshore cruises.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8mL/L from 0db to 20db

±0.4mL/L from 20db to 100db

±0.2mL/L from 100db to 250db

±0.04mL/L below 250db

The usual pattern of CTD fluorescence reading higher than extracted CHL samples for low values of CHL and lower for high CHL values was seen for most bottles, but there were a number of near-surface exceptions where high CHL values were associated with even higher fluorescence values. Overall, fluorescence is close to chlorophyll values, with an average ratio of FL/CHL being 0.88 and a fit through the origin shows FL=0.97 * CHL.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets as well as various analysis logs were available. There were detailed notes in the log explaining cases where sampling was not done in the usual way. In many cases bottles were fired for another institution that used a different sample # protocol, so no sample numbers were assigned in the rosette sheet. Where there are no sample #s assigned, that line will be removed from the CHE file, or no CHE file will be prepared if there was no IOS sampling.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. The only corrections needed were to 2 calibration dates and to update the value of parameter E for the dissolved oxygen sensor. The latter correction is not important for this cruise since it concerns hysteresis that should not be a factor due to fairly shallow sampling, but in case others “borrow” the configuration file, it is best to fix it.

The file was saved as 2014-20-ctd.xmlcon.

The PAR sensor was included in the configuration file but was likely never mounted. After conversion this can be checked and that channel will be removed later, if not used.

##### BOTTLE FILE PREPARATION

The ROS files were converted using file 2014-20-ctd.xmlcon.

Those files were put through CLEAN to add event numbers (*.BOT).

Header Check was run on the BOT files and no problems were found.

Temperature and salinity were plotted for all BOT files. CTDEDIT was used to remove points from a section of bad secondary salinity during event 5 around 30db. The edited file was copied to BOT.

The BOT files were then averaged to enable an ADDSAMP file to be prepared. Sample numbers were added to the ADDSAMP file based on rosette log records. Event #1 contains only bottles fired at the bottom and there was no sampling, so no CHE file will be produced. Sampling during events #6, 7, 19, 25, 27 and 31 was done for another institution and no sample numbers were entered on the rosette sheet, so no CHE files will be produced. Files 1, 6, 7, 19, 25, 27 and 31 were removed from the ADDSAMP file. In many cases the first 3 bottles and the last 3 were not assigned sampled numbers and were removed from the ADDSAMP file; those bottles were fired for another institution who used their own sample numbers.

SAM files were created using the Add Sample Number routine. Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2014-20-bot-hdr.txt.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-20.xls. A simplified version of the spreadsheet was saved as 2014-20chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-20oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2014-20oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2014-20SAL.xls. The file was simplified and saved as 2014-20sal.csv. That file was converted to individual SAL files. The salinity data were analyzed 9 to 23 days after collection.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-20nuts.xls which included a report on precision.  The file was simplified, reordered on sample numbers and saved as 2014-20-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

Data from the SAMAVG files were exported to a spreadsheet and all casts checked against rosette sheets to ensure all expected bottle data were present. Errors found were:

- Event #5 – CHL sample 45 (dups ticked on rosette sheet) is missing. Note in log says not collected by mistake – should be flagged “9”. Fixed in CHL file, but spreadsheet not adjusted.

- Event #9 – there was no sampling for IOS; will not be processed further.

- Events #13, 14, 15, 16 and 21 had bottles fired with no IOS sampling and in most cases no sample #s assigned. These lines were removed from the SAMAVG files.

- Event #26 – Nutrients, Salinity and CHL have the wrong event #. It should be 26, not 25, so they were missed in the merge process. These will be fixed in the original spreadsheets and file names, as needed.

- Event #33 – Salinity was shown as taken on the rosette spreadsheet, but were missing from the MRG file. The event # was entered wrong in the analyst’s spreadsheet. This was fixed in the spreadsheet which was once again converted to SAL files.

After changes were made the SAMAVG files were exported to a spreadsheet again and no problems were found.  The MERGE process was repeated.

##### Compare

Salinity

Compare was run with pressure as reference channel. Salinity samples were analyzed within 9 to 23 days of collection.

There were many near-surface bottles and most were outliers, but the CTD salinity is very noisy, so this is not indicative of analysis problems. Outliers investigated were:

- Cast #5 –sample 34 – Standard deviation in the CTD salinity data is high and the sample is from the bottom where flushing can be an issue. Likely the bottle is ok. No flag was added.

- Cast #5 – samples 40 and 41 – Both bottles have higher salinity than the CTD by >0.02. Standard deviation in the CTD is fairly low, but a shed wake passed through early in the stop that contained high enough salinity to account for the salinity from the 2 bottles. No flag was added.

- Cast #8 – Samples #62 and 63 at 75db – The bottles have higher salinity than the CTD by ~0.06 and ~0.09. Again, a shed wake passed through early in the stop with high enough salinity to account for this difference. No flags were added.

- Cast #10 – Sample 102 at 214db – 2nd bottle fired at bottom differs markedly from other bottle at same depth and is even fresher than bottle at 200db. The first bottle at the bottom is ok.  The CTD salinity has a low standard deviation, so the in situ water appears to have settled by firing time, but flushing can be a problem at the bottom. It is possible that one Niskin carried lower salinity water from above, while the other Niskin flushed more completely. No flag was added as this is not obviously an error in collection or analysis.

- Cast #32 – Sample #201 – The CTD salinity is much lower than the duplicate samples which were in good agreement with each other. There was a shed wake that passed through at the beginning of the stop with salinity even higher than that seen in the 2 samples. The shed wake may well have entered the Niskin bottles. The CTD salinity had settled fairly well by firing time, but likely the Niskin bottles had not flushed completely after the shed wake passed by. No flag was added as the bottle values likely do reflect the contents of the Niskin bottle.

When those outliers are excluded, there is still more variability than we expect in comparisons, with the primary salinity being low by an average of 0.014 and the secondary low by 0.008; the standard deviations were ~0.021. Most of the low CTD salinity values are above 150m, which is also indicative of poor flushing as the effect is more significant in the higher salinity gradients found nearer the surface.

When only bottles below 140db are included (with the exception of 1 outlier) the primary salinity was lower than bottles by an average of 0.0044 but the standard deviation in the fit is 0.0052 reflecting the large scatter. The secondary salinity was higher than bottles by 0.0022 with a standard deviation of 0.0050. This implies a salinity difference of 0.0066, higher than during 2014-18. The difference does vary according to which data are included in the fits.

For full details for the COMPARE run see file 2014-20-sal-comp1.xlsx.

Note that file 2014-20-sal-comp2.xlsx contains the results of a comparison of salinity with bottles after a test recalibration of the secondary salinity. It was later decided to not archive the secondary, so no salinity recalibration was applied to the CHE and CTD files.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The fit of differences against DO concentration has a lot of scatter. First, large outliers were excluded and then fits were tightened by removing further outliers based on residuals. Usually fits are quite different depending on whether the offset is allowed to be non-zero or not, but in this case there was virtually no difference with the best fit being either:

DOX_BOT = 1.0404 * DOX_CTD  		R2 = 0.8336

or 	DOX_BOT = 1.0402 * DOX_CTD +0.0009  	R2 = 0.8336

We usually get a tighter fit than this, but removing more outliers produces a fit with very few points and does not look convincing. The fit with zero offset will be chosen for recalibration.

If flushing of Niskin bottles is a problem as suggested by the salinity comparison, we expect to see an effect on oxygen as well. The effect is likely reduced somewhat for this cruise due to fairly low and more complex gradients in DO, except at the surface where they are frequently very high. There are plenty of outliers, but they are mostly near the surface and they do tend to show the CTD DO being high relative to bottles, which makes sense if flushing is incomplete. Calibration drift leads to CTD values being too low. So flushing probably is an issue, but is not as easy to study because of the slow response of the CTD sensor and the complexity of near-surface DO gradients.

There are many outliers, but the only major ones were from the top 10m or had already been flagged by the analyst. No further flags were added.

Plots of Titrated DO and CTD DO against CTD salinity were examined. The only significant outlier had been flagged by the analyst.

Fluorescence

COMPARE was run using the SeaPoint fluorescence and the Extracted Chlorophyll from bottles. The ratio of SeaPoint fluorescence to Extracted CHL normally varies with the fluorescence being higher at low values of CHL and then falling to lower than CHL at values >1ug/L. In this data set, most points follow that pattern, but there are 7 outliers with the fluorescence higher than the CHL at values of CHL>1ug/L with some at very high CHL values.  These all occur in the top 10m. For the near-surface CTD fluorescence values that are not higher, most are about 60% of the extracted CHL values, which is typical of what we usually see for high CHL. Can this variability be explained, at least in part, by incomplete flushing? A fit through the origin gave FL=0.97*CHL, and an average ratio of FL/CHL was 0.88.

In areas where fluorescence has a maximum near the surface, poor flushing might lead to the CTD fluorescence looking high relative to extracted CHL samples. In areas where there is a sub-surface maximum the CTD fluorescence would look too low above the maximum and too low below it. For this particular cruise most cases with a sub-surface maximum also have the highest fluorescence, though there are some casts with high values but a maximum at the surface. So at least some of the scatter may be due to poor flushing rather than poor performance of the sensor or sampling/analysis problems. Beyond that are the complex questions about the basic difference between what is measured by the sensor compared to what is measured by extracted chlorophyll.

Several of the CHL near-surface values were flagged due to duplicates differing by >15%.

For more detail see file 2014-20-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2014-20-ctd-new.con.

A few casts were examined and all expected channels are present except PAR which was apparently not mounted. The descent rate was kept high, on average. For some casts it is very steady, while noisier for others. For at least one cast there were complete stops during the downcast. The two temperature and conductivity channels are closer during the downcasts The salinity channels differ by about 0.006 on the way down but have very similar shapes; on the upcasts there is great variability with differences from 0 to 0.1. The conductivity is similar to salinity with differences on the order of 0.0005. The 2 temperature channels are close on the downcasts but noisy on upcasts.

Altimetry looks useful though there is sometimes a lot of noise near the bottom so checks should be made that the header entries are reasonable; dissolved oxygen, fluorescence, transmissivity and PAR look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Previous uses of this equipment have shown an advance of 4s to bring the DO voltage into line with temperature. A few checks were made to ensure this setting is appropriate and the results with the 4s setting look reasonably good, though the differences between the near-surface upcast and downcast temperature traces make that judgment difficult and at depth there is little variation.

ALIGNCTD was used to advance the DO Voltage by 4.0s relative to the pressure.

##### CELLTM

Tests are also difficult for the CELLTM settings.

The upcast data are noisy so the usual tests for CELLTM settings were hard to interpret since they require a comparison of downcast and upcast TS profiles and the upcasts are very noisy. The results from earlier cruises that used these sensors were applied and the results look reasonable.

CELLTM was run using (α = 0.02, β=9) for the primary and (α = 0.03, β=9) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are always very noisy, but these data are much noisier than usual. So these estimates are extremely rough.

For comparison, the results for some casts from other recent cruises that use the same sensors are included.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-06-0053 | 350 | ~0 | 0.00044 | 0.0044 | V.Steady |
| 2014-27-0073 | 335 | +0.0002 | 0.00045 | 0.0045 | V. steady |
| 2014-21-0036 | 350 | -0.001 XN | 0.0004 | 0.005 | High, X Noisy |
| “ | 1000 | ~0 | 0.00043 | 0.005 | “ |
| “ | 1950 | +0.0001 | 0.00045 | 0.005 | “ |
| 2014-21-0128 | 350 | -0.0002 | 0.00045 | 0.0055 | High, X Noisy |
| “ | 1000 | -0.0002 | 0.00045 | Too noisy | “ |
| “ | 1950 | -0.0001 | 0.00047 | 0.006 | “ |
| 2014-18-0029 | 350 | ~0 XN | 0.00032 VN | 0.0048 VN | High, Mod |
|  | 1000 | ~0.0002 VN | 0.00034 | 0.0051 | High, VNoisy |
|  | 1950 | 0.0003 | 0.00046 | 0.0052 XN | “ |
| 2014-18-0068 | 350 | 0.0003 | 0.00042 | 0.0047 | High, Noisy |
|  | 1000 | 0.0001 | 0.00045 | 0.0053 | “ |
|  | 1950 | 0.0002 VN | 0.00045 | 0.0054 | “ |
| 2014-20-0006 | 350 | 0.0002 | 0.00056 | 0.006 VN | High, F.Steady |
| 2014-20-0021 | 350 | 0.0003 | 0.0006 | 0.0062 | High, Steady |
|  | 425 | 0.0003 | 0.00062 | 0.0064 | “ |
| 2014-20-0026 | 350 | 0.0002 | 0.00062 | 0.0065 | “ |
|  | 475 | 0.00025 | 0.00062 | 0.0064 | “ |
| 2014-20-0034 | 300 | 0.0002 | 0.0006 | 0.0061 | “ |

The differences in temperature are noisy, but the average differences are small. The conductivity differences are larger than usual and larger than in the previous cruises. The salinity differences appear to be increasing in a fairly steady fashion.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  Fluorescence did not go offscale. The speed check is fine. The track plot looks fine and was added to the end of this report.

There was a check of deck pressure at the beginning of one cast and the reading was ~0.6db. Yet, the data from 0.2db look like they are from very close to the surface, but definitely in water. If we subtract 0.6db then it would give us out of water values. The deck pressures vary considerably from cast to cast, so readings from just 1 cast are not considered a reliable pressure check. Moreover the accuracy of the sensor is ~±1db. The average surface pressure as found by the Surface Check routine was 0.9db, which is reasonable. It does not suggest that the reading is too high.

The mixed layer depth estimate confirms that no surface salinity data are likely to be useful in a bottle comparison.

The cross-reference check was compared with the log book and no problems were found.

The cruise track was plotted and added to the end of this document. No problems were found.

The altimeter readings from the headers of the CLN and MRGCLN2 files were exported to spreadsheets. A few casts did not get within 15m of the bottom so there are no header entries. Plots were made of a selection of casts and no problems were found in the altimetry headers of the CLN files.

The water depth header entries were checked against log entries; a few discrepancies were found and checked. In some cases the entry looks fine, but in 4 cases a change was made to match the log book.

The same changes were made in the SAMAVG files. The maximum pressure of the SAMAVG files were also checked to see if there was any sampling near the bottom. No casts that are to be processed had only surface samples so the altimeter heading is appropriate.

##### Shift

Fluorescence

SHIFT was run on all casts to advance the SeaPoint fluorescence by +24 records.

After this step cast #34 was checked to ensure that this setting worked well and it did. That was the only cast with no stops for bottles.

Conductivity

Based on the results of cruises 2014-20 and 2014-18 initial tests were run using values close to the settings of -0.5records for the primary conductivity and +1.4records for the secondary. The casts had little noise in the primary salinity so there was not a large difference seen, but -0.5 looked good. The secondary was much noisier before the alignment step; the setting of +1.4 did a good job of removing the noise. No other setting was found that looked better.

Two runs of SHIFT were used to apply -0.5records shift to the primary conductivity and +1.4records to the secondary conductivity. Salinity was recalculated for both channels.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The alignment looks ok.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings were that events #7 and 31 were less than 10m deep, and those were surface only casts. Only bottle files need be prepared for those events so the full profile data will not be processed further.

##### Other Comparisons

Previous experience with these sensors – All pressure, conductivity, temperature and dissolved oxygen sensors were recalibrated in late December 2013 or January 2014 and were used for 6 other cruises before this one, 4 in near-shore waters and 2 offshore.

1. Salinity:

The difference between the sensor pairs was ~0.004 during the first use after the factory visit and the difference has grown slowly since then. From the earlier cruises it appeared that the secondary salinity was closer to the bottles and the primary salinity appeared to be low. However, the sampling was shallow and conditions were such that poor flushing of Niskin bottles is suspected to have led to bottle contents coming from lower in the water column than the level of the CTD sensors. For one of the offshore casts with good salinity sampling and prompt analysis, the primary salinity was found to be very close to bottles while the secondary was high by ~0.005. Salinity analysis was somewhat delayed for another offshore cruise which might have led to some evaporation. From that cruise the primary was found to be low by ~0.001 and the secondary high by 0.004. Delayed analysis has been found to raise bottle values due to evaporation; the effect varies from bottle to bottle leading to a lot of scatter in the fits and less confidence than usual in the results. The secondary sensors were selected for archiving of the previous cruises using this equipment because there was less noise in the data, but the secondary salinity was recalibrated to bring it into line with the primary which is likely providing more accurate values. The effect of poor flushing and delayed analysis is similar with bottles reading higher than expected but not in a predictable way.

2. Dissolved Oxygen

The earlier cruises included 2 with too few bottles to enable a reasonable comparison and many bottles close to the bottom which are not reliable. Two that did provide enough bottles produced results that were very unusual with large offsets in the fit of differences against CTD DO, but fairly low slopes. This may be related to fairly shallow sampling, and the very quiet conditions under which the samples were gathered may reduce the efficiency with which the Niskin bottles were flushed. There were then 2 open ocean cruises that provided quite different offsets. The fits for the 4 most useful cruises before this one were:

DOX_BOT = 1.0286 * DOX_CTD + 0.0772 (2014-06)

DOX_BOT = 1.0325 * DOX_CTD + 0.1019 (2014-15)

DOX_BOT = 1.0400 * DOX_CTD - 0.0259 (2014-21)

DOX_BOT = 1.0417 * DOX_CTD - 0.0289 (2014-18)

3. Pressure

The factory calibration has been found satisfactory for the 6 previous cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.  Local climatology was available for only a few of the casts; those data all fell within the expected bounds.

Repeat Casts –

There were some repeat cast, but they are mostly too shallow to expect good repeatability. Casts #25 and 26 did go to 500db, so they were compared on a T-S surface. The differences suggest sections of active mixing in the top 400m, but below that the curves are close; below 475db the temperature differences along lines of constant σt are ~0.0005C° and for salinity they are ~0.0003. This is excellent correspondence.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

From this cruise the bottle comparison shows the primary salinity to be low by an average of 0.0044 and the secondary high by 0.0022. The analysis was run within 9 to 23 days after collection. As discussed in §4 there is evidence of poor flushing of Niskin bottles. The primary salinity was found to be very close to the bottles during 2014-21 when flushing was likely good and analysis was quick. The primary is likely closer to the bottles than the secondary for this cruise as well. For the previous 6 cruises, the secondary temperature and salinity looked a little more stable and easier to edit, so those channels were selected for archiving. During this cruise there were problems near the surface with unstable features, sometimes in the primary channels but mainly in the secondary with salinity appearing to “overreact” to temperature changes in the top 10 to 20m. Below 20m the secondary temperature and salinity look somewhat better than the primary but not enough better to justify the severe editing that would be needed at the surface if the secondary T/S data were selected.

The cause of this problem is assumed to be obstructed or irregular flow in the duct so that the alignment selected to match temperature and conductivity at depth does not suit the near-surface data. The protocol of soaking the CTD at 10m before running a cast was not used so bubbles may be causing the problem. Another possible cause might be problems with the secondary pump, though even the primary looks somewhat affected. One cast with severe problems was the first event in Saanich Inlet. Data from the 2 previous cruises which used the same equipment and also sampled Saanich Inlet as their first events show no such features. Surface data are often noisy but these problems are of a character that we normally only see if the pumps are not turned on.

Since the secondary does look better below the surface, it was chosen for editing and archiving.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect bin-averaged values and records corrupted by shed wakes including many near-surface records. The descent rate was generally very steady except at the top and bottom of casts, though there were a few sudden stops at mid-depths during event #16. All files required some editing.

All edited files were copied to EDT.

##### Initial Recalibration

The pressure looks ok.

The primary salinity was found to be lower than bottles by an average of ~0.0044 while the secondary was high by 0.0022, so the average difference between the two channels during bottle stops included in the fit is 0.0066. If we average all bottles, the difference is 0.0062 and if we judge from section 10 observations it is from 0.0060 to 0.0065 with 0.0064 at the deepest observations. So using a figure of 0.0064 is probably reasonable.  During 2014-21 and 2014-18 the salinity corrections applied to the secondary channel to bring it into line with the primary channel were -0.0052 and -0.0053. For this cruise it appears that there has been more drift so a correction of -0.0064 to the secondary salinity would be appropriate. However, since only the primary channels have been edited, the secondary salinity will not be recalibrated.

SBE Dissolved Oxygen data was recalibrated using file 2014-20-recal1.ccf to apply the fit described in section 4:

DOX_BOT = 1.0404 * DOX_CTD

CALIBRATE was run on the MRGCLN2, SAM and EDT files using that file.

COMPARE was rerun to check that the DO correction was done correctly and it was. The average difference after removing outliers shows the CTD DO to be within 0.0002mL/L of bottles; the standard deviation was 0.02mL/L. (For details see file 2014-20-dox-comp2.xlsx.)

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. ALIGNCTD corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a 2nd correction is sometimes found appropriate to further address response time errors determined by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. When the differences were plotted against DO concentration, the scatter was too large to allow a believable fit no matter how outliers were identified. On average, the CTD DO reads high by ~0.04mL/L with a standard deviation of 0.08mL/L. The differences are small at depth, with the CTD DO reading slightly lower than bottles below 300db. Above 150m the CTD DO is generally higher than bottles, but that may be at least partly due to poor flushing. Even without flushing as an explanation, no further calibration is justified due to the high noise level in the fit.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channels in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were a few small near-surface unstable features but they may be real, so no further editing was applied.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section, to fix many headers and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity and Fluorescence data are nominal and unedited except that some

records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

The near-surface salinity data quality is considered lower than usual.

There appear to have been problems with the flow rate near the surface

which affected the secondary salinity most severely, but the primary

salinity data were also affected.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8mL/L from   0db to  20db

±0.4mL/L from  20db to 100db

±0.2mL/L from 100db to 250db

±0.04mL/L below 250db

For details on the processing see the report: 2014-20_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots and T-S plots were examined and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged widely from ~80% to ~135% with most between 100% and 110%. All casts outside of Douglas Channel had values near 105%, as expected. Inside Douglas Channel the values vary widely. The lowest values were at Doug26; there is a subsurface DO maximum at 10m with fairly low values at the surface and surface fluorescence is low.  The highest values were at stations Doug11, KSK1 and Doug45 and the first of 2 casts at FOC1. The difference between the 2 FOC1 casts is likely due to how close to the surface the CTD data started as there is a large gradient there.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number, to fix the geographic area and to add a comment about quality flags and analysis methods and a few notes about the CTD data. The formats for nutrients were adjusted to reflect the analyst’s recommendation.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. No problems were found.

A header check was run and no problems were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and the only warnings concerned nutrient formats which had been adjusted according to the analyst’s recommendation.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

1. Test cast – Starting pressure 0.6db. Temp diff -0.00073, Sal diff -0.00073 (Note: all data are too noisy to be significant.) All bottles fired to test – no samples taken.

2. Effingham 3 blocked by log boom so samples 1-8 not taken; went to EFF04 for this cast.

5. CHLa sample #45 – not sampled by accident. Sal diff often up to 2nd decimal place.

6. Station FOC1 - 8 bottles were fired by mistake near the bottom when 8 should have been fired at the surface. Firing was continued in the normal sequence skipping the surface bottles.

7. Repeat at FOC1. The 8 surface bottles were closed. No IOS samples from this cast. No CHE file needed.

9. No sample numbers – samples for another institution using their numbers. No CHE file needed.

10-14. Sample #s out of sequence among casts 10, 13, 14 but normal order within each cast.

13-15. Some bottles fired for BIO Cooger scientists – no sample numbers.

15. Altimeter <1m off bottom – CTD raised before closing bottles.

16. Few complete stops during downcast.

19. 50m to 0m for BIO Cooger group – 8 bottles, no sample numbers.

25. All bottles for Cooger – no sample numbers. Altimeter late to kick in and lost at bottom. Don’t think it touched bottom but raised CTD before firing bottles. Bottle 5 closed by accident.

26. Overshoot on way up stopped at 143m and went back down to 150m.

27. Confusion due to inexperienced LARS operator – some confusion in sampling – log book contains detailed notes, but there were no sample numbers and no samples for IOS.

31. Surface only, no sample numbers, no IOS samples.

32. Attached an SBE37 ODO to test with CTD DO sensor. Rosette held near bottom for 5 min. to test sensor.

33. Test of 2nd SBE37 DO sensor. Overshot 100m bottle. Stopped at 94m and went back down to 100 to close bottle. Mixup in sample #s - samples begin at 217 (207-216 do not exist)

34. CTD cast only, no rosette firing.

CTD SUMMARY

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 31Dec2013 | Factory |  |  |
| Conductivity | 1766 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 4700 | 4Jan2014 | Factory |  |  |
| Secondary Cond. | 3321 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 21Jan2014 | Factory |  |  |
| SeaPoint Fluor | 3642 | Jan. 2014 | Factory |  |  |
| Pressure Sensor | 77511 | 30Dec2013 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |