## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 27 Mar 2025 | Updated channel names & formats in TOB files. G.G. |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 16 August 2020 | ### Added HPLC data. S.H. |
| 1March2017 | Correction to CHL in file 2014-21-0006.che. G.G. |
| 24 Sept 2015 | Correction to file 2014-21-0008.che. Samples #22 and 23 had been entered in wrong order. Niskin bottles were fired out of order. G.G. |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |
| 19 Nov. 2014 | Added note to Summary section on post-cruise sensor calibrations. GG |

## PROCESSING NOTES

Cruise: 2014-21

Agency: OSD

Location: WCVI

Project: La Perouse / WCVI

Party Chief: Yelland D.

Platform: John P. Tully

Date: 27 May 2014 – 8 June 2014

Processed by: Germaine Gatien

Date of Processing: 14 July 2014 – 4 September 2014

Number of original HEX files: 99	Number of CTD files:

Number of bottle files: 56		Number of bottle casts processed: 53 (3 not sampled)

Number of original TSG files:   7	Number of processed TSG files:  7

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1119) on the primary pump, a SeaPoint Fluorometer (#3642) with a 3X cable on the secondary pump, a Biospherical QSP-200 PAR sensor (#4601) and an altimeter.

A thermosalinograph (Seacat 21 S/N 3363 ) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor and a flow meter.

The data logging computer was #3.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log had personnel and equipment lists, but the section for the TSG was not completed. Since the raw TSG data are often not downloaded with the raw CTD data, a note in the log helps to ensure that someone looks for it.

This cruise was the 5th of 6 that used the same sensors between March and June 2014. While there was salinity calibration sampling from all but 1 of them, for 4 the comparisons with the CTD are noisy and considered unreliable due to problems with poor flushing of Niskin bottles and/or evaporation of samples due to delayed analysis. For 3 cruises there was a very steady CTD ascent rate and sampling was shallow and mostly in areas with a high vertical gradient of salinity; these factors make the errors due to poor flushing significant. Cruise 2014-21 provides the best comparison of the 6 because analysis was very prompt and a noisy descent rate probably improved the flushing of Niskin bottles. This showed the primary salinity to be close to bottles and the secondary high by about 0.005. Since the secondary salinity is less spiky, the secondary temperature and salinity channels were selected for archiving. The secondary salinity was recalibrated by subtracting the average difference between the two channel pairs.

NOTE: 19 November 2019: Post-cruise factory calibration shows that the secondary salinity was high by about 0.005 by mid-August. Based on observations during a series of cruises, it appears that the calibration applied to these data in July was appropriate, though there remains some doubt about when drift occurred. G. Gatien

During offshore cruises during 2013 there were unusual excursions in salinity that appeared to be due to varying alignment of T and C sensors and were associated with high deceleration rates of the CTD. During this cruise there were a few minor cases of salinity looking out of line during rapid deceleration, but the problem was not severe, probably because the range of descent rate was generally lower, rarely going exceeding 2m/s.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8mL/L from   0db to  50db

±0.6mL/L from  50db to 100db

±0.3mL/L from 100db to 150db

±0.2mL/L from 150db to 500db

±0.05mL/L below 500db

As usual, the ratio of SeaPoint fluorescence to Extracted CHL varies as CHL increases. There is a lot of variability for CHL<1ug/L with the fluorescence being up to 4 times the CHL though the average is closer to 1; the ratio falls to about 0.5 when CHL reaches 3ug/L. and stays at about that level for higher values. By mistake no chlorophyll sample was collected at 5m during event #123, so a loop sample was taken as a substitute for sample #453.

The TSG data quality was much improved from other uses over the past few years. A watchful eye was kept on the system during the cruise. The change to mounting 2 filters allowed a quick switch from one to the other, enabling cleaning without much data loss, though one filter proved better than the other. The intake temperature was available throughout the record. Salinity data were much less noisy suggesting that the debubbler worked well. The fluorescence compared well with the CTD fluorometer and reasonably well with loop chlorophyll; this was a different fluorometer than used in the past. There was one short patch of noisy salinity, but even that was not as bad as generally seen in recent years.

The TSG flow rate was raised several times, but between those adjustments it was very steady. The temperature is closest to the CTD using a flow rate of 0.95 and the same rate produces a salinity difference closest to that found in the loop comparison. However, to make the TSG intake temperature become closer to the CTD temperature suggests that the higher flow draws water from deeper waters, closer to the 5m bottle stop. But for salinity the differences get larger as though the water came from shallower waters. This contradiction may just be the result of too much variability in all the data sets, especially the loop samples. Furthermore, what works best when stopped may not be best underway. The fluorescence did not vary significantly with varying flow rate.

Further flow rate studies with lots of loop sampling (while moving and while stopped) might help establish what flow rate is best. LaPerouse is an excellent project for such tests with many CTD casts to enable studies while stopped, but there are many variables to consider, so robust conclusions may not emerge.

This was the first use of the TSG since the last factory service, so there is cause to doubt the evidence that suggests that the salinity is low by from 0.02 to 0.04. There is a lot of noise in all the salinity data comparisons. No recalibration was applied.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Files from the initial cast were sent to IOS for initial processing. No errors were found in the configuration file, but it was noted that there was a lot of noise in the primary conductivity.

##### Preliminary Steps

The Log Book and rosette log sheets as well as various analysis logs were available.

A few errors in station names and depths in the file headers were noted in the log.

Extracted chlorophyll, NH4, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts. The draw temperature was recorded for DO sampling so concentration can be calculated in mass units as well as mL/L.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments. The only corrections needed were to 2 calibration dates. A new SeaPoint fluorometer was used with a 3X cable throughout.

The file was saved as 2014-21-ctd.xmlcon.

The PAR sensor was not always mounted. Based on log notes cast lists were prepared with and without PAR so that it will be easy to remove PAR as appropriate.

PAR on: 1-28, 39-69, 79-95, 109-117, 138-180.

PAR off: 29-36, 71-77, 96-107, 119-136.

In preparing the lists above it was discovered that the cast saved as #118 should be #117; event #118 was a net cast. The file name was changed after conversion.

##### Hysteresis Study

##### Since this is the first cruise sampling deep water since this DO sensor was last calibrated at the factory, the hysteresis parameter E should be fine-tuned. However, while there are some deep stations, there are not enough data to do this well. Moreover, an initial look at the comparison between CTD DO and titrated samples showed there was a serious problem with the analysis of some DO samples. So processing was suspended while hysteresis tests were done on data from cruise 2014-18 which followed this one and had much deeper sampling and used the same sensors. Based on that cruise parameter E in the DO calibration was set to 0.0031.

##### BOTTLE FILE PREPARATION

The ROS files were converted using file 2014-21-ctd.xmlcon.

Those files were put through CLEAN to add event numbers (*.BOT).

Header Check was run on the BOT files and no problems were found.

Temperature and salinity were plotted for all BOT files. A few spiky points were removed from the primary salinity in casts #4, 26, 35, 56, 64, 72 and 86 and from the secondary salinity in casts #26 and 144 using CTDEDIT. The edited file was copied to BOT.

Cast #36 looks odd but this was a case of bottles being fired on the move just to get them all closed. This file will not be processed further.

The BOT files were then averaged to enable an ADDSAMP file to be prepared. Sample numbers were added to the ADDSAMP file based on rosette log records.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2014-21-bot-hdr.txt.

Dates of creation were added to the names of spreadsheets from analysts.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-21.xls. Event numbers were missing, so those were added. A simplified version of the spreadsheet was saved as 2014-21chl.csv which was then converted to individual CHL files. Loop data were removed from that file but were added to file 2014-21-loop.csv.

Note: CHL sample #453 was not collected as intended, so a loop sample was taken as a substitute (loop126).

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-21oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2014-21oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2014-21SAL.xls. The file was simplified and saved as 2014-21sal.csv. That file was converted to individual SAL files. Loop data were removed, but were first added to file 2014-21-loop.csv. The salinity data were analyzed 3 to 14 days after collection.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-21nuts.xls which included a report on precision.  The file was simplified, reordered on sample numbers and saved as 2014-21-nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2014-21_NH4.xls which included a report on precision. The file was simplified and saved as 2014-21-NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

Data from the SAMAVG files were exported to a spreadsheet and all casts checked against rosette sheets to ensure all expected bottle data were present. One error was found and corrected in the CHL spreadsheet and in the CHL files. The MERGE process was repeated for 2 casts to correct the errors.

Files 2014-21-0039 and 2014-21-0114 were removed from the cast list since there was no sampling; they will not be processed further.

##### Compare

Salinity

Compare was run with pressure as reference channel. Salinity samples were analyzed within 2 to 3 weeks of collection.

There were no major outliers. When one minor outlier was excluded the primary salinity was found to be low by an average of 0.0003 and the secondary was found to be high by an average of 0.0052. The standard deviations were 0.0007 and 0.0008 respectively. The scatter in the fits is ~±0.0015 while the precision estimate from duplicates is Sp=0.0002.

For full details for the COMPARE run see file 2014-21-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The outliers were investigated:

- Event #4 – The bottle seems a little high – based on CTD data would expect this bottle value around 20db. CTD upcast and downcast values close. (An error was later found in assignment of sample numbers. Once corrected this looked fine.)

- Event #19 – CTD data looks a little odd and local gradient is high, so probably bottle ok.

- Event #28 – Bottle at 41db looks like the CTD at 25db, but the CTD data is pretty noisy and local gradients high, so probably bottle was ok. Similarly at 20db difference is high but local gradients high.

- Event #41 – At 64db the CTD is odd and at 9db there might be poor flushing of the Niskin, so likely no analysis problem. No flags justified.

- Event #44 – CTD data noisy.

- Event #53 – Bottle at 10db already flagged 4 and surface sample a little out of line, but not unusual at surface.

- Event #72 - 8 bottles from above 125db. The CTD data looks fine with upcast and downcast in good agreement. The bottle values are very high and no such values are found at other casts in that area. This looks like the problem noted during 2014-06 that was determined to be an analysis problem. Flagged 5 by analyst.

- Event #89 – CTD data noisy.

- Event #90 - 4 bottles above 30db. While this looks like the event #72 outliers, the bottle values are not unreasonably high and a look at the CTD DO profile shows that the CTD data are the problem here. The upcast DO data look bad. The primary salinity also looks quite different from the secondary salinity and from both downcast salinity channels above 50db, so the problem most likely was due to the primary plumbing. There is a note in the log to say that a rinse was done with a syringe after the cast, though it is not specific about what was rinsed. The primary salinity and dissolved oxygen data were replaced with pad values in the SAMAVG file which was then remerged with the MRGCLN1a file.

The fit of differences against DO concentration is close to linear if an offset is allowed. When low DO values are allowed and outliers are removed based on residuals:

DOX_BOT = 1.04 * DOX_CTD -0.0259  R2 = 0.96

It makes very little difference whether DO values <0.5mL/L are included or not, if an offset is allowed. If the offset is set to 0, then it does make a more significant difference with a fit that looks ok at intermediate values but not at the high or low end of the DO range.

A preliminary comparison from 2013-18 which followed this cruise shows a fit of:

DOX_BOT = 1.0431 * DOX_CTD -0.0338  R2 = 0.93

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were detected.

NOTE: An error was later found in the Event #4 data; when COMPARE was rerun with that correction the fit was almost identical, with the same slope but an offset of 0.0256mL/L.  There is no need to redo the recalibrations as the change would be negligible.

Fluorescence

COMPARE was run using the SeaPoint fluorescence and the Extracted Chlorophyll from bottles. The ratio of SeaPoint fluorescence to Extracted CHL varies with the fluorescence being up to 4 times the CHL for CHL<1ug/L; the ratio falls to about 0.5 when CHL reaches 3 or 4ug/L. and stays at about that level for higher values.

A linear fit through the origin shows the CTD fluorescence being 55% of the extracted CHL, but the average is 92% and median 82% because there are many low values where the ratio FL/CHL is much higher.

For more detail see file 2014-21-fl-chl-comp1.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using 2014-21-ctd-new.con. The file names and depths were fixed in files noted as having errors in the log book. It was also noted that the format of station names for the EFF stations were not the same as for most other station names, so those were changed from Eff** to EFF** for neatness.

A few casts were examined and all expected channels are present. The descent rate was kept high, on average, but is extremely noisy with obvious evidence of shed wake corruption. The two temperature and conductivity channels are closer during the downcasts

Altimetry looks useful; dissolved oxygen, fluorescence, transmissivity and PAR look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on 4 casts to determine the offset between the DO channel and the primary temperature. All settings tested improved the alignment but a 4s advance looked best. That is the advance that was found best for all previous uses of this sensor since its last servicing.

ALIGNCTD was used to advance the DO Voltage by 4.0s relative to the pressure.

##### CELLTM

The upcast data are noisy so the usual tests for CELLTM settings were a little hard to interpret. Tests were run on many few casts using a variety of settings to see which produced the closest upcast/downcast on a T-S plot. The differences among settings are slight, but overall the setting (α = 0.02, β=9) is the best for the primary conductivity for most casts and that was the finding for other cruises using these sensors. For the secondary, the best setting was using (α = 0.03, β=9) but for some the setting using (α = 0.0245, β=9.5) was best. The latter was used for other recent uses for this equipment. One cast that had a very steady descent rate had better results with (α = 0.03, β=9) so that will be selected.

CELLTM was run using (α = 0.02, β=9) for the primary and (α = 0.03, β=9) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are always very noisy, but these data are much noisier than usual. So these estimates are extremely rough.

For comparison, the results for one cast each from 2014-06 and 2014-27 since the same sensors were used. While descent rates were steadier for ht earlier casts, there was no deep sampling..

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-06-0053 | 350 | ~0 | 0.00044 | 0.0044 | V.Steady |
| 2014-27-0073 | 335 | +0.0002 | 0.00045 | 0.0045 | V. steady |
| 2014-21-0036 | 350 | -0.001 XN | 0.0004 | 0.005 | High, X Noisy |
| “ | 1000 | ~0 | 0.00043 | 0.005 | “ |
| “ | 1950 | +0.0001 | 0.00045 | 0.005 | “ |
| 2014-21-0073 | 350 | -0.001 | 0.0004 | 0.005 | High, X Noisy |
| “ | 1000 | ~0 | 0.0004 | 0.0053 | “ |
| “ | 1950 | -0.0001 | 0.00042 | 0.0054 | “ |
| 2014-21-0105 | 350 | ~0 | 0.00044 | 0.0055 | High, X Noisy |
| “ | 1000 | +0.0001 | 0.00048 | 0.0056 | “ |
| “ | 1950 | -0.0001 | 0.00048 | 0.0057 | “ |
| 2014-21-0128 | 350 | -0.0002 | 0.00045 | 0.0055 | High, X Noisy |
| “ | 1000 | -0.0002 | 0.00045 | Too noisy | “ |
| “ | 1950 | -0.0001 | 0.00047 | 0.006 | “ |

The differences in temperature are very noisy, but the average differences are small. The conductivity differences are fairly small but appear to be increasing, though it is not completely clear. The salinity differences appear to be increasing in a steadier fashion, which fits the slow drift noted in bottle comparisons through 4 cruises that used this equipment before this one.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  There are no off-scale or negative fluorescence values.

There was a check of deck pressure at the beginning and end of one cast and values were both ~0.5db. The deck pressures vary considerably from cast to cast, so readings from just 1 cast are not considered a reliable pressure check.

Cast #82 includes some records at the end when pumps were on but rosette was out of the water. The CTD pressure is ~0.45db, which is about right as the top of the rosette is ~1.5mb above the sensor. It is possible that the pressure is slightly too high, but not by much. The CTD was still reading “in-water” conductivity values and fluorescence, but it must have been very close as the transmissivity has a few sections with very high values, close to what is expected in air.

The average surface pressure as found by the Surface Check routine was 2.6db, which is reasonable. It does not suggest that the reading is too high.

The mixed-layer was shallow so surface salinity bottles are not useful for calibration purposes except at a very few sites.

The cross-reference check was compared with the log book and a few files had errors in the station name or the format of station names. Those were corrected.

The cruise track was plotted and added to the end of this report. No problems were found.

The altimeter readings from the headers of the CLN and MRGCLN2 files were exported to spreadsheets. Some casts did not get within 15m of the bottom so there are no header entries. Plots were made of a selection of casts and no problems were found in the altimetry headers of the CLN files.

The water depth header entries were checked against log entries; discrepancies of <4m were ignored and others were investigated. Using an estimate based on maximum depth sampled plus minimum altimetry, it was decided whether the log or header entry best represented water depth when the CTD was at the bottom. The header in the IOS file was changed where the log entry appeared a better choice. For cast #75 neither seems right, so there was probably a lot of variation in depth through the cast; an estimate was entered. CLEAN was rerun for the affected casts.

The same changes were made in the SAMAVG files. The maximum pressure of the SAMAVG files were also checked to see if there was any sampling near the bottom. For all casts that had only surface samples, the altimeter heading was missing, as appropriate. The MERGE with MRGCLN1s was repeated.

##### Shift

Fluorescence

SHIFT was run on all casts to advance the SeaPoint fluorescence by +24 records.

After this step a few plots were checked to ensure that this setting worked well and it did.

Conductivity

An initial test on one cast showed that a negative shift is needed for the primary and positive for the secondary. Tests were run on a group of casts.

For the primary conductivity the best results overall were with -0.5 records, though the data were extremely noisy so it is hard to judge.

For the secondary conductivity there was even more noise. The best result was with a setting of +1.3 records.

Two runs of SHIFT were used to apply those settings.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The alignment looks ok, though the noisy signal makes this very hard to judge.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors – All sensors were recalibrated in late December 2013 or January 2014 and were used for 5 other cruises since then, 4 before this one in near-shore waters and 1 after this well offshore.

1. Salinity:

The differences between the sensor pairs was 0.004 during the first use after the factory visit and the difference has grown slowly since then. From the earlier cruises it appeared that the secondary salinity was closer to the bottles and the primary salinity appears to be low. However, the sampling was shallow and conditions were such that poor flushing of Niskin bottles is suspected to have led to bottle contents coming from lower in the water column than the level of the CTD sensors. For the cruise that followed this one a preliminary comparison shows the primary salinity closer to bottles than the secondary, but not as close as observed during this cruise. Analysis was somewhat delayed for the later cruise which might have led to some evaporation and hence, slightly higher bottle values, on average.

2. Dissolved Oxygen

The earlier cruises included 2 with too few bottles to enable a reasonable comparison and many bottles close to the bottom which are not reliable. The two that did provide enough bottles produced results that were very unusual with large offsets in the fit of differences against CTD DO, but fairly low slopes. This may be related to fairly shallow sampling and the very quiet conditions under which the samples were gathered may reduce the efficiency with which the Niskin bottles were flushed. The fits for the 2 casts before this one and the one that followed were:

DOX_BOT = 1.0286 * DOX_CTD + 0.0772 (2014-06)

DOX_BOT = 1.0325 * DOX_CTD + 0.1019 (2014-15)

DOX_BOT = 1.04     * DOX_CTD - 0.0259 (2014-21)

3. Pressure

The factory calibration has been found satisfactory for the 4 previous cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. These are not appropriate for the casts closest to shore. The only excursions were some high temperature and low salinity values right at the surface of offshore casts and some high temperature values at the bottom of 1 offshore cast and high temperature values at the bottom of 2 casts in the Strait of Georgia. These were very minor outliers and not systematic, so not indicative of calibration problems.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

There were no post-cruise calibrations available.

##### DETAILED EDITING

From this cruise the bottle comparison shows the primary salinity to be low by an average of 0.0003 and the secondary high by 0.0052. The analysis was run within 3 to 14 days after collection. There is little difference between the two sensor pairs in plots that were checked, but there was trouble with the primary for one upcast. The secondary channels will be selected for archiving because they have been chosen for other cruises due to more stable T-S traces. The salinity will be recalibrated to match the primary.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records. There were some bad salinity records in areas of rapid deceleration though this was less of a problem than seen in 2013. The descent rate was often noisy. Records were often removed from near the top and bottom of casts.

All edited files were copied to EDT.

##### Initial Recalibration

The pressure looks ok.

The primary salinity was found to be lower than bottles by an average of ~0.0002 which the secondary was high by 0.0053, so the average difference between the two channels during bottle stops is 0.0055. Examining downcasts we find differences increase through the cruise from about 0.005 to 0.0055 at 2000m. However, in the bottle comparison the temporal drift appears to be in opposite directions for the two channels. The noise level is high enough to cast doubt on these comparisons. During 2014-18 the difference was 0.0053 in motion and 0.0051 during bottle stops. Subtracting 0.0053 from the secondary salinity for this cruise looks reasonable.

SBE Dissolved Oxygen data was recalibrated using file 2014-21-recal1.ccf to subtract 0.0053 from the secondary salinity and to apply the fit described in section 5:

DOX_BOT = 1.04 * DOX_CTD -0.0259

CALIBRATE was run on the MRGCLN2, SAM and EDT files using that file.

COMPARE was rerun to check that the DO correction was done correctly and it was. The average difference after removing a few outliers shows the average difference shows the secondary salinity to be within 0.0001 of bottles and  CTD DO to be within 0.001mL/L of the bottles on average, though the standard deviation is 0.02mL/L. (See file 2014-21-sal-comp2.xls.)

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. ALIGNCTD corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time errors found by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. When the differences were plotted against DO concentration, the scatter was too large to allow a believable fit no matter how outliers were identified. Overall the differences look small at the bottom and, on average, small at the top of the range. In the middle of the range the values look a little high, but that could be due to the bottle data coming from a little lower in the water column due to incomplete flushing; this factor would be most notable in the higher gradients in the mid-depths. When outliers were excluded based on residuals, the average difference indicated that the CTD DO sensor read higher than titrated DO samples by ~0.02mL/L but the standard deviation was 0.08mL/L.

A check was made to find what change in depth is needed to find a change in CTD DO of 0.02mL/L and for two casts checked between 100m and 300m an offset of from 0.1m to 2m was found; this would vary greatly with local gradient but we are mostly interested in the high gradient zones where the CTD appears to read high. The descent rate of the CTD may also be a factor in how well the Niskin flushes. Fortunately, the effect of incomplete flushing is reduced a little because bottles are mounted ~1.3m above the CTD. Nonetheless, minor inefficiency in flushing would easily account for the CTD DO being higher than bottles by 0.02mL/L in waters with a large vertical DO gradient.

Even without flushing as an explanation, no further calibration is justified due to the high noise level in the fit.

##### Special Fluorometer Processing

Special files were prepared for Dr. Peña by clipping the COR1 files to 150db. The files were bin-averaged (0.25db bins) and put through HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAM files were put through REMOVE and named *.BOF and saved.

A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channels in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were a few small near-surface unstable features but they may be real, so no further editing was applied.

##### Final CTD File Steps (REMOVE and HEADEDIT)

A check was made of the cast lists with and without PAR sensor by plotting PAR.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

Profile and T-S plots were produced at this point to check for errors. There were small unstable features that were investigated at the editing stage – these may be real.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section, to fix many headers and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8mL/L from   0db to  50db

±0.6mL/L from  50db to 100db

±0.3mL/L from 100db to 150db

±0.2mL/L from 150db to 500db

±0.05mL/L below 500db

For details on the processing see the report: 2014-21_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values ranged widely from ~80% to ~130% with most between 105% and 110%. The lowest values were just offshore of Effingham Inlet and near Brooks Peninsula where the surface temperature and oxygen gradients were low and fluorescence mostly quite low. The highest values were at the head of Effingham Inlet and at stations LG01 and LG02 near shore and in the Strait of Georgia and the surface DO and temperature gradients were high and fluorescence high. Surface DO samples are close to the CTD DO values and often a little higher than the CTD DO values. So the high saturation is likely due to high productivity and not poor calibration of the DO sensor.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number, to fix the geographic area and to add a comment about quality flags and analysis methods and a few notes about the CTD data. The formats for nutrients were adjusted to reflect the analyst’s recommendation.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. Two problems were found:

- Sample #466 Nutrients is shown on the rosette sheet, but missing from the nutrient spreadsheet. However, there were 2 samples with #464 and one of these looks like the missing sample – it makes sense in profile. The analyst confirmed this explanation and corrected the nutrient spreadsheet.

- Sample #478 CHL was analyzed but not shown on the rosette sheet. The label is clear, so this must have been an extra sample taken, just not noted in the log.

A header check was run and no problems were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and the only warnings concerned nutrient formats which had been adjusted according to the analyst’s recommendation.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found.

##### Thermosalinograph Data

Data were provided in 7 hex files.

The external thermistor was connected.

There were 7 loop salinity samples taken, of which just one was taken while the ship was stopped.

a.) Checking calibrations

The configuration files for the 7 files are identical and the parameters are correct. One of the files was saved as 2014-21-tsg.con.

b.) Conversion of Files

The files were converted to CNV files using configuration file 2014-21-tsg.con. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

Time-series plots were produced. The data look good with just one section of salinity that is rather noisy. The flow rate was 0.5 until June 2nd at about 1500 hours when it was raised to 0.65. The rate rose again to 0.95 on June 5 at about 0500. The flow rate was unusually steady in each of the 3 sections, although there was a brief drop to 0.3 during file #1. A new configuration was used for the TSG during this cruise with seawater flowing through a Rusco 1.5” filter (mesh size 30),  through a MSRC VDB-1 VORTEX debubbler and then into the thermosalinograph. There were 2 filters so that they could be quickly switched for cleaning. The noisy salinity is seen a few hours after a switch of filters and shortly before the flow rate was increased.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2014-21-ctd-tsg-comp.xls.

87 CTD casts with data from 4m overlapped with TSG data. The 7 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The average differences were 0.0000° for both, with the largest difference being 0.0008°.

This shows both the times and positions are reliable for both systems

This spreadsheet will also be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Rosette samples and TSG and CTD data

- T1 vs T2 The average difference between the intake and lab temperature was 0.33C° and the median was 0.27C° when all data are included. The flow rate varied from 0.48 to 0.97. When the differences are plotted against flow rate, there is no obvious pattern, with many outliers at all rates. When plotted against intake temperature there is a fairly clear pattern of decreased heating as intake temperatures rise. This makes the plot against flow rate hard to interpret because the highest flow rates were from cast #134 onwards where the surface temperatures were generally lower. It is likely that low flow rates do increase heating in the loop. When only casts with heating between 0.17C° and 0.37C° are included, the average heating in the loop is 0.290C° and the median is 0.255C°.

A check was made of file #6 to see if the heating in the loop was different when the ship was in motion. Early in the file the flow rate was about 0.65 and the heating had a median value of 0.287C°. The flow rate was increased to 0.95 for most of the file which we would expect to lead to less heating, but the temperatures were slightly lower, on average, than in the early section which would raise the heating a little. The median difference was 0.246C°. During ship stops with the higher flow rate, the median difference was 0.295C° or 0.281C° if outliers were excluded. So heating in the loop does appear to be lower when the ship is moving. This may be because water is drawn from a little higher in the water column where the temperature is slightly higher and salinity lower, or it may be just applicable to this section of data. The heading of the ship and sea state may all be relevant. It does suggest that recalibration of data based on times when the ship is stopped may be justified if differences are very large, but in most cases is not advisable.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

1. Intake Temperature The TSG intake temperature was higher than the CTD temperature by a median value of 0.090C°. When a few outliers (mostly from Effingham Inlet) are excluded the median difference is 0.086C°.

2. LAB TEMP Using the same data as in the analysis of the intake temperature the lab temperature is found to be higher than the CTD temperature by 0.328C° or 0.324C° when outliers are excluded. Most, but not all, outliers are associated with high standard deviation in TSG temperature.

3. SALINITY Excluding the same outliers as used in the analysis of the intake temperature plus a few other outliers, the TSG salinity is lower than the CTD salinity by a median value of 0.0.0267. When the data are separated for low, medium and high flow rates, the medians are 0.022, 0.027 and 0.040, but there are many fewer data available for medium and high rates. This suggests that the higher the flow rate, the shallower the water entering the loop, but it might also be due to more bubbles in the loop at high flow rates. Most, but not all, outliers are associated with high standard deviation in the TSG salinity. Based on the mixed layer depth as calculated in the surface report, only a few casts were well-mixed to below 10m (salinity varying by <0.005 compared to 4db value). A check was made of those casts and the differences vary widely.

4. FLUORESCENCE Again, using the data from part 1 of this section, the average and median ratio of TSG fluorescence to CTD fluorescence is 1. The same ratio is found using all the data as found with outliers excluded. A plot of TSG FL versus CTD FL shows that there is good correspondence for FL<2.5ug/L but a lot of scatter at higher values. Only 20 casts out of 87 had CTD fluorescence >0.25ug/L. At higher values there is more scatter, but the overall agreement is good. The ratio was higher at a medium flow rate than at low or high, but the difference is not significant. Outliers do not appear to be proportional to the standard deviation in the TSG fluorescence.

5. WELL_MIXED CASTS Based on the mixed layer depth as calculated in the surface report, only 9 casts were well-mixed to at least 10m (salinity varying by <0.005 compared to 4db value). The median value for those indicates that the intake temperature is high by 0.086, the salinity is low by 0.03 and the ratio of TSG fluorescence to CTD fluorescence is ~1.

(See 2014-21-ctd-tsg-comp.xls.)

Loop Bottle - TSG Comparisons There were 8 loop samples for both salinity and extracted chlorophyll, but all but 1 of them came from the period when the flow rate was 0.65.  The TSG salinity was lower than the loop salinity by an average of 0.044 and a median of 0.036. For the two samples with very low standard deviation in the TSG salinity, the TSG is low by ~0.018 and 0.028. The TSG fluorescence is was 0.79% of extracted CHL on average (median 72%). For the highest CHL values the ratio was 0.75for CHL=10ug/L and 0.63 for CHL=17ug/L.

(See 2014-215-loop-tsg-comp.xlsx.)

- Surface rosette CHL vs TSG fluorescence

There were no near-surface salinity samples during this cruise. CHL samples were available.

A comparison was made between the TSG fluorescence and the Niskin bottle CHL data. There were 42 cases of samples taken during TSG acquisition. The match is not expected to be great because the times for the TSG data are from the start of casts, not the time of the rosette sampling, and variability is high in both and there is likely some mismatch in depths that may be systematic. The TSG fluorescence is found to be about 68% of the extracted CHL on average, but it tends to read higher for CHL<1ug/L. At higher values there is a lot of variability and a generally lower ratio.

(See 2014-21-rosette-tsg-comp.xls.)

- Calibration History

The TSG temperature and conductivity were recalibrated in December 2013 and this appears to have been the first use since then.

The fluorometer was used during 2013-18 during which the TSG fluorometer was about 1.1 to 1.2 times the SeaBird fluorometer on the CTD when the flow rate was >1.

Conclusions

1. The TSG clock appears to have worked well.

2. The TSG flow rate was raised several times, but between those adjustments it was very steady.

3. The temperature in the loop increases by about 0.25Cº between the intake and lab based on comparisons during stops for CTD casts. The heating in the loop varies according to intake temperature and flow rate.

4. The TSG intake temperature appears to be higher than the CTD temperature by ~0.086Cº when outliers are excluded but that varies from ~0.095 Cº when the flow rate is low to ~0.02Cº when the flow rate is high. Given some mismatch in time and depth of sampling between the two sources, recalibration is not justified.

5. The TSG Salinity is lower than the CTD salinity by about 0.027, but that varies with flow rate; for low, medium and high flow rate the differences are 0.022, 0.027 and 0.040 which may suggest that the water is being drawn from shallower waters when the rate is high, or it may be that there are more bubbles at higher flow rates. The TSG salinity is lower than loop samples by a median value of 0.036 – most of these samples had flow rate 0.65. This might imply that the water in the loop comes from higher in the water column when the ship is moving, or that there are more bubbles in the loop water when underway.  Confidence in the comparison of TSG and loop salinity is low because the standard deviation in the differences is 0.049, and given the mismatches in time and depth of sampling for the TSG-CTD comparison, recalibration is not justified.

6. The TSG fluorescence is close to the CTD fluorescence. It is an average of 79% of the loop CHL samples. The TSG fluorescence is about 65% of the extracted CHL samples from the 5m rosette bottle. There is too much scatter to detect a pattern, which is not surprising given mismatches in time and depth.

7. A rough analysis suggests that the high flow rate produces a better match to the CTD, with temperature differences reducing as the rate increases. The salinity differences increase, but at rate 0.95 they reach the difference found in the comparison with the loop salinity samples, though that data set had a very high standard deviation so may not be reliable. There is likely too little data to say that a high flow rate is best, and it is hard to understand why the high rate works well for both T and S. If at low flow rate the TSG is measuring water above 4.5m and then as the rate increases it draws water from deeper, that would reduce the difference in temperature. But it should also decrease the difference in salinity, which it does not. There is likely too little data and too much variability to reach a conclusion about this, especially given that the standard deviation in the comparison with loop samples was the same size as the median of differences. The fluorescence did not vary significantly with varying flow rate.

8. Now that the TSG is performing better, we could run tests to determine the best flow rate. A good test for salinity would be to try a few different flow rates and take many loop salinity samples. We can only test the temperature while stopped. La Perouse is an excellent cruise for such tests since there are many CTD casts so we could compare the salinity results underway with those while stopped. However, there are many sources of error and a well-mixed surface layer would limit what we learn.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT to see if there was a need to remove any bad data. There were only small spikes in salinity and fluorescence which were left unedited. The only cases of low flow rate were so brief that they had little effect on the data.

g.) Recalibration

The loop comparison suggests that the salinity is low by 0.04 but the standard deviation in the comparison is 0.05, so this result is not robust enough to justify recalibration.

The temperature and conductivity sensors were calibrated in December 2013 and there is no record of them having been used since then. So a drift of +0.04 seems unlikely.

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

The flow channel was not removed since it may be useful for further studies of optimal flow rate.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

Par On: 1-28, 39-69, 79-95, 109-117, 138-180.

Par Off: 29-36, 71-77, 96-107, 119-136.

12. Stop at 8db for 2 min

36. All Niskins closed just to have them closed, no samples. Not processed.

39. Cast for sounder calibration only – no samples. Not processed.

41. Station name should be EFF02 – wrong in header

52. Wrong depth in header – should be 86.

82. Acquisition not stopped until the rosette was out of the water.

88. Pressure on deck 0.53db, 0.49db.

89. Computer froze on upcast; brought on board, rebooted computer, cast restarted.

90. Oxygen trace very strange from 40db on the upcast. Did a rinse with syringes after the cast.

97. Wrong header name and depth.

107. Wrong depth in file.

114. 1 bottle fired but not sampled. Not processed.

117. Saved as 118, but log shows rosette cast at this time and place as #117; #118 was a net cast.

126. Loop collected 2xCHL, 1xSal instead of sample #453 which should have been taken during 123.

130. Bottle 21 closed to test but not sampled.

22:23 on 4 June: TSG file 6 started and switched to other filter.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 31Dec2013 | Factory |  |  |
| Conductivity | 1766 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 4700 | 4Jan2014 | Factory |  |  |
| Secondary Cond. | 3321 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 21Jan2014 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| SeaPoint Fluor | 3642 | Jan. 2014 | Factory |  |  |
| Pressure Sensor | 77511 | 30Dec2013 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2014-21

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 28Mar13 | Factory |  |  |
| Conductivity | 3363 | 28Mar13 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | Aug12 | IOS |  |  |
| Temperature:Secondary | 0603 | 7Mar12 | “ |  |  |
| Flow meter | ? | n/a |  |  |  |