## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 27 Mar 2025 | Updated channel names & formats in TOB files. G.G. |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 02 August 2020 | ### HPLC data added. S.H. |
| ?? March 2015 | 1. Post-cruise calibration reports showed that the recalibration applied to temperature and recalculation of salinity were inappropriate. Calibration was redone. For details see note at end of this report after §25. 2. Header notes about salinity bottle inserts were corrected in CHE files. |

## PROCESSING NOTES

Cruise: 2014-22

Agency: OSD

Location: WCVI

Project: La Perouse – Zooplankton Monitoring

Party Chief: Yelland D.

Platform: John P. Tully

Date: 5 September 2014 – 15 September 2014

Processed by: Germaine Gatien

Date of Processing: 18 December 2014– 25 February 2015

Number of original HEX files:	99	Number of CTD files: 98 (1 file upcast only)

Number of bottle files: 48	Number of bottle casts processed: 44 (2 joined & 3 not sampled)

Number of original TSG files:  	8	Number of processed TSG files: 8

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (# 0443) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#3642) with a 3X cable on the primary pump, a Biospherical QSP-200 PAR sensor (#4615) and an altimeter (#1204). This CTD was used for cast #7 and all casts thereafter except for some test casts that are not to be processed.

A thermosalinograph (Seacat 21 S/N 3363) was mounted with a Wetlabs/Wetstar fluorometer (WS3S-713P), remote temperature sensor and a flow meter.

The data logging computer was #3. Seasave version V7.22.4 was used for acquisition.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette sampling log sheets were generally in good order, but there was no entry for the TSG. Sometimes TSG data do not get downloaded with the CTD data and it is a good idea to make it clear that the system was used.

Sampling notes were provided and were helpful in processing.

The pressure was recalibrated by applying an offset of +0.6db.

CTD salinity samples were analyzed within 38 to 47 days. The salinometer malfunctioned during the analysis of samples from event #121. There was a lot of scatter in the comparison between CTD and bottle salinity with both CTD salinity channels appearing to be lower than during the 2 previous cruises that used the same sensors. This is likely due to evaporation of samples. The results of cruise 2014-19 are considered more reliable; at that time the primary salinity was lower than bottles by an average of about 0.003 while the secondary salinity was very close to bottles. As found for 2014-19 and 2014-60, the temperature channels were much further apart than usual and evidence suggested that the primary sensor was the one out of line. As was found for 2014-19, recalibrating the primary temperature to bring it closer to the secondary, followed by re-computing salinity, brought the two salinity channels into good correspondence.

The fit of the Oxygen:Dissolved:SBE data against dissolved oxygen bottles is unusual, but similar fits were found during the 2 previous cruises when the same sensor was used. During 2014-19 there was a sudden change in the fit partway through the cruise. The fit for this cruise is very similar to that for 2014-60 which immediately preceded it. During 2014-19 there were many 2-sided spikes in the SBE dissolved oxygen profiles; there are a few small ones in these data, but the problem is not as severe.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8mL/L from 0db to 75db

±0.3mL/L from 75db to 200db

±0.1mL/L from 200db to 500db

±0.05mL/L below 500db

The SeaPoint fluorescence data were frequently bad below 500db until the gain cable was changed before Event #73. All fluorescence data below 510db were removed from bottle and CTD files for events #1-#72.

For this cruise and the one that preceded it, a new TSG setup was used with water flowing through a pond filter before entering the debubbler. The flow rate was quite steady and the TSG salinity traces were mostly free of the large spikes that have been plaguing this system for some time. The salinity data were recalibrated based on the results of cruise 2014-19 because there was good calibration sampling from both CTD casts and underway loops and samples were analyzed promptly. The data from the Wetlabs fluorometer were mostly bad for this cruise though they looked fine for the previous cruise and for the last file from this cruise. Comparisons with the fluorometer on the CTD and with CTD CHL samples all indicated that the TSG fluorescence was out of line (usually much too high but occasionally very low) and the traces were frequently unnaturally smooth. The fluorescence was removed from all files except 2014-22-0008.tob which had believable fluorescence values and compared well with loop samples.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets as well as various analysis logs were obtained.

Sampling notes were reviewed and the headers of raw files for casts #77 and 79 were corrected based on those. (Those changes were later found to be wrong and were changed back to the originals.)

Notes were added to the particulars section at the end of the report based on the sampling notes and log book notes.

Extracted chlorophyll, DMS, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. They had been used for the 2 previous cruises, 2014-60 and 2014-22.

The calibration constants were checked for all instruments. A few small errors in spelling and the date of the pressure sensor calibration were fixed in the configuration file, which was then saved as 2014-22-ctd.xmlcon.

The PAR sensor was not always mounted; cast lists PAR and NO PAR were prepared based on log notes.

##### BOTTLE FILE PREPARATION

The ROS files were converted using files 2014-22-ctd.xmlcon with Tau and Hysteresis corrections. selected. Those files were put through CLEAN to add event numbers (*.BOT).

Temperature and salinity were plotted for all BOT files. There were a few spikes in salinity which were examined with CTDEDIT and a few points were removed or smoothed in casts #63, 72, 74 and 127.

Header Check was run on the BOT files and no problems were noted.

The BOT files were then averaged to enable an ADDSAMP file to be prepared. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

Sort was used to ensure that the ADDSAMP file was in sample number order.

Event #43 was interrupted during the upcast and continued as event #44. The BOT file for the first file was renamed as BOTx, while the second file was renamed with the number of the first file plus BOTy. It is easier to combine such files if the bottles fired during the 1st part are fired again at the beginning of the 2nd file. That didn’t happen this time. The bottle positions are correct in all files, but the bottle firing numbers needed to be fixed in the second file of each pair. The bottle numbers were adjusted in the 2nd files using Ultraedit and then JOIN was used to create new files, 2014-19-0069.bot and 2014-19-109.bot.

The following files will not be processed further:

- Cast #54 - run with vents open and was then rerun as 56

- Cast #74 - run with vents open and was then rerun as 76

- Cast #77 - a single bottle was fired by accident and there was no sampling

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

It was found that the fluorescence data were bad in many casts below 500db, with unbelievably high values for that depth, or 0 values. CLEAN was run on casts 1-72 to replace fluorescence values below 510db with pad values. At cast #73 the fluorometer cable was replaced and the problem was resolved.

Altimetry and water depth entries were corrected in the MRG and SAM files as described in section 12.

Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was ordered on sample number and converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2014-22-bot-hdr.txt.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-22CHL.xls. A simplified version of the spreadsheet was saved as 2014-22chl.csv which was then converted to individual CHL files. Loop data were copied to file 2014-22-loops-tsg-rosette-comp.xlsx.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-22oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2014-22oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2014-22SAL.xls. The file was simplified, loop samples were removed and the file was saved as 2014-22sal.csv. That file was converted to individual SAL files. The salinity data were analyzed 38 to 47 days after collection. Loop samples were saved to 2014-22-loops-tsg-rosette-comp.xlsx.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-22nuts.xls which included a report on precision.  The loop samples were copied to file 2014-22-loop.csv. The file was simplified, reordered on sample numbers and the file was saved as 2014-22nuts.csv. The file was converted to individual NUT files. The loop data were copied to file 2014-22-loops-tsg-rosette-comp.xlsx.

NH4

NH4 data were obtained in file 2014-22_NH4.xls which included a report on precision. The file was simplified and saved as 2014-22NH4.csv and converted to individual NH4 files.

For some of the samples one event number was listed as 43/44 because the cast was split. This was changed to 43 since the two will be joined in a single file, 2014-22-0043.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The MRG files were exported to a spreadsheet and compared to the rosette sheets. A few inconsistencies were noted:

- The dissolved oxygen samples from LB08 were listed as being from event 13, but should be from event 17. This was changed in the working files and the analyst adjusted the DO spreadsheet.

- Chlorophyll samples from event #83 were not marked on the rosette sheet, but the label makes it clear that they were taken. However, the label gives the station name as LB12 and it should be LB13. The spreadsheet is correct. The label was written in pencil – so obviously this was a last-minute arrangement.

- Some Chlorophyll samples had not been converted properly to CHL files, so the process was repeated.

- Event #76 was a repeat of event #74 and samples came from #76. But the rosette sheet said #74, so the samples were mostly labelled as from #74. Only dissolved oxygen was entered correctly in analysis spreadsheets. The analysts were informed and corrected the spreadsheets.

- The rosette log shows samples originally intended to come from Niskin #5 actually came from #11 for event 76. The CHL labels make it clear that samples came from bottle 11 (#5 had been found to be leaky.)  The ADSAMP, SAM and SAMAVG files were adjusted to reflect this.

- The chlorophyll sample #315 was entered as being from event #92; this should be event #90.

The draw temperatures were compared with the CTD temperature. As expected, there is a large difference for deep samples and smaller ones near the surface. There were some outliers that almost all came from cast #38. A check was made to see if there might be a problem with bottles firing during cast #38, but bottle samples show that was not the case.

An initial header check did not turn up any further problems.

##### COMPARE

Dissolved Oxygen

COMPARE was run using DO from the MRG and SAM files.

There is insufficient data to do hysteresis tests and they were run on a recent cruise, 2014-19.

Differences were plotted against CTD DO for all casts and the fit when outliers are excluded based on residuals is:

DOX_BOT = 1.0325 * DOX_CTD - 0.1960

Or when outliers were based on standard deviation in the SBE DO data, the fit was:

DOX_BOT = 1.0320 * DOX_CTD - 0.1956

The 1st fit had less scatter so may be more reliable, but there is little difference. Fits using different combinations of casts showed some variation, but this seems most dependent on the depth of the casts. For very shallow casts there are smaller ranges of values and more outliers because they often come from areas of high vertical gradients. All fits had much larger offsets than usual for this type of sensor. For SBE DO values <6mL/L the sensor read higher than the bottles, on average, something we sometimes see at very low DO values but not above 1mL/L.

Tests were done by varying the fits by time or depth. Excluding points based on depth had no significant effect on the offset; the slope varied but not in a systematic way. There was variation in the different cast ranges but again it did not seem systematic. Even choosing using DO values >2 had no significant effect. However, when the casts that were shallower than 100db were excluded there was some difference, with a higher slope and a slightly smaller offset. This is likely due to a lesser influence from bottles in higher vertical DO gradients. This probably explains the differences among different cast ranges since some contain more shallow casts than others. The variations that were found may not be significant since they are based on fewer bottles, so the results vary greatly depending on how outliers are identified. But in all cases the offsets were higher than normal.

There have been other cases of larger offsets in 2014 including 2 inshore April cruises using a different sensor (#1117); those fits had offsets of about 0.07 and 0.10. That sensor was later used on 2 offshore cruises and those fits against bottles produced much smaller offsets. It is possible those odd fits were due to poor flushing of Niskin bottles in protected waters. It was noted that the gradients were fairly low and the descent rate very steady. The correction was similar in size to normal fits.

The other cases of large offsets came from the previous 2 cruises on which this sensor was used, 2014-19 and 2014-60. There was a sudden shift in the fits of the earlier cruise with a normal offset for the early casts and a sudden change to a higher offset that lasted for the rest of the trip. The offset was about 0.10 for the latter part of that cruise and was about 0.20 for 2014-60 which was the next use of the sensor. There was little variability in the slope for these fits. Poor flushing is not very likely in an offshore cruise and the corrections are much smaller than usual, so the behavior does not look like what happened in the April cruises. There was an observation of the DO traces shifting during one cast at about the time the fits change. There were also many 2-sided spikes in the DO traces from 2014-19 which is not as notable in these data though there are some. This may be because 2014-19 had much deeper sampling and the spiking is mostly seen in very deep water.

For 2014-19 there was no evidence of analysis or sampling problems and none were found for this cruise. The sensor was freshly calibrated at the factory and behaved normally for a while.

There is a lot of scatter in the fit. Most outliers are associated with high standard deviation in the CTD DO.  The other noteworthy outliers are associated with large DO gradients and often there is evidence of a large shed wake at the beginning of the bottle stop. The differences between CTD and bottles are more likely due to the bottles containing water from a different level than problems with analysis or sampling; the titrated values are likely representative of the bottle contents. No further flags are justified.

For more details see 2014-22-dox-comp1.xls.

Salinity

Compare was run with pressure as reference channel. Salinity samples were analyzed within 38 to 47 days of collection so we expect some evaporation of samples.

An initial comparison was done that showed that when only bottles below 200db were included and some outliers excluded, the primary salinity was lower than bottles by about 0.0057 and the secondary salinity lower by 0.0031 with standard deviations of 0.002. The outliers come primarily from cast #121, and all cases of bottle values being lower than CTD salinity by more than 0.001 come from that cast. Only 1 bottle from that cast was higher than the CTD and the CTD salinity was very noisy for that case.

An examination of the salinity analysis log sheets shows that the samples from cast #121 were the last ones to be analyzed and that the analyst encountered a problem shortly before doing these samples. He found the salinity had dropped when he ran a standard, so he worked on raising the temperature in the lab. The next standard was ok so that he continued to run samples, but a malfunction in the salinometer was later found. It appears that the malfunction first appeared at the end of the 2014-22 run.

Even when cast #121 is excluded there is a lot of scatter as expected when samples wait a long time before being analyzed. Above 200m the bottle values are all low which could be due to evaporation and/or incomplete flushing of Niskin bottles.

There is one outlier below 200m, sample 331 (event #108) which was already flagged 3 by the analyst due to the screw cap being cracked though the insert was fine. The analyst felt the sample was probably ok. The value is out of line at a depth where poor flushing is unlikely, so the bottle value is probably high due to evaporation. While the cracked cap in itself might not have had a large effect, the fairly long wait for analysis likely led to significant evaporation. A change to flag 4 looks appropriate with the addition of a comment that it is a significant outlier in comparison to CTD data. The difference is not so large that pad values are needed.

All bottles from cast #121 were flagged 5 and replaced with pad values; the salinometer is known to have malfunctioned at some point and it looks like it started shortly before these samples were run.

For full details for the COMPARE run see file 2014-22-sal-comp1.xls.

Fluorescence

COMPARE was run to see how the SeaPoint fluorescence channel compared with extracted chlorophyll.

The SeaPoint Fluorescence has the usual general pattern of being higher than extracted CHL at low values and then falling to lower ratios, but there is a split in the data with a lower FL/CHL ratio closer to shore and a higher ratio well offshore.  This is not just due to higher CHL being found near-shore, since the ratio is higher near shore even when CHL < 1ug/L.

For more detail see file 2014-22-fl-chl-comp.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using con files 2014-22-ctd.xmlcon with hysteresis and Tau corrections chosen.

A few casts were examined and all expected channels are present.

The two temperature and conductivity channels track fairly well during the downcasts, but upcast data are much noisier.

PAR, SeaPoint fluorescence and SBE dissolved oxygen data look normal.

The altimetry has a useful signal, but there are spikes near the bottom of the some casts, so results of the header algorithm will need to be checked.

The fluorescence data were bad below 500db for casts #30, 31, 35, 38, 39, 57 and 60 with some high values and large sections of 0 values. The gain cable was replaced for event #73 which probably explains the fact that later deep casts look fine. Rather than go through the slow process of using a graphical editor to remove the bad Fluorescence data, the CLEAN routine will be used to remove all fluorescence values below 510db for all casts up to #72.

For the bottle files the SAM files were edited in the same way.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

DO sensor #0997 was used during 2014-19 and 2014-60 and a setting of +2.5s was found to improve the results.

As was found for the earlier cruises, the high variability in temperature makes upcasts and downcasts hard to compare. It is thus difficult to determine the best alignment setting to use, but a choice of +2.5s produces reasonable results.

ALIGNCTD was used to advance the DO Voltage by +2.5s for all casts.

##### CELLTM

The usual test for CELLTM settings is to check which ones produce the best correspondence between downcast and upcast traces on a T-S plot. Such tests are hard to interpret when the temperature data are noisy, but all choices improved the data. For 2014-19 the choice of (α = 0.03, β=9) produced the best results overall for both conductivity sensors and the same choice looked reasonable for 2014-60.

CELLTM was run using (α = 0.03, β=9) for the primary and secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences were very noisy.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-19-0079 | 400 | -0.0037 N | -0.0003 N | ~0 N | High, Noisy |
|  | 1000 | -0.0030 | ~0 N | 0.0024 XN | High, V Noisy |
|  | 1980 | -0.0033 | ~0 N | 0.0030 N | “ |
|  | 2980 | -0.0035 | ~0 N | 0.0035 N | “ |
|  | 3900 | -0.0036 N | ~0 N | 0.0036 VN | “ |
| 2014-60-0024 | 400 | -0.0043 | -0.00022 | 0.0017 | High, V Steady |
| 2014-22-0035 | 400 | -0.0029 XN | -0.00015 XN | +0.0002 XN | High, F.Steady |
|  | 1000 | -0.0036 VN | -0.00009 N | +0.0026 N | “ |
|  | 1980 | -0.0034 VN | -0.00028 N | +0.0032 N | High, VNoisy |
| 2014-22-0084 | 400 | -0.0035 XN | -0.00012 N | +0.0005VN | High, Noisy |
|  | 1000 | -0.0030 XN | -0.00005 | +0.0026 | “ |
|  | 1980 | -0.0030 XN | -0.00001 N | +0.0031 | “ |
| 2014-22-0132 | 400 | -0.0037 VN | -0.0002 | +0.0006 XN | High, VNoisy |
|  | 1000 | -0.0032 VN | -0.0006 VN | +0.0023 VN | “ |
|  | 1980 | -0.0034 VN | -0.00003 VN | +0.0032 VN | “ |

The descent rate of the CTD was mostly noisy though not extremely so.

A cast from each of 2014-19 and 2014-60 are included for comparison, but there was no deep sampling for the latter.

The differences in temperature are consistent with no obvious temporal or pressure dependence.

There is a lot of noise in the conductivity differences but they are all very small.

The salinity differences are pressure dependent which is not unusual and may be related to the effect of minor misalignment of primary and secondary sensors having more influence in shallow waters. There is no temporal variation.

The salinity differences at 1000db are close to those found in section 4.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. For Events #1-72, the fluorescence values below 510db were replaced with pad values.

##### Checking Headers

The header check was run.  Pressure got as low as -0.5db during event #140 and there was a negative DO value for event #125; there may be more – these are the extreme values. The maximum/minimum descent rates were quite high.

The surface check gave an average value of 2.0db which if fairly low for offshore sampling from the Tully. There were data from Event #140 with “in water” salinity values when the pressure was about -0.45db. The transmissivity decreased rapidly at about -0.4db and was near zero at-0.45db. During 2014-60 the pressure was believed to be low by about 0.6db based on many observations of negative pressures with “in water” non-zero conductivity/salinity. That is likely a good estimate for this cruise too.

The cross-reference check was compared with the log book and the only discrepancies were 3 station names – casts #77, 79 and 159. The first 2 had been changed earlier due to an entry in the sampling notes that must be incorrect judging by the track plot. The 3rd case was just a typo in the entry. The 3 entries were corrected in the CLN files.

The altimeter, maximum pressure and water depth were exported from the headers of the CLN and MRG files to spreadsheets. The water depths were checked against the log book and many discrepancies were found, chiefly cases where the log entry had obviously been changed. The new log entries look closer to the maximum pressure (converted to depth) plus altimeter value than the header values, so the header entries were adjusted. In a few cases there was no header entry so the log entry was used. In one case the altimeter had no signal and the log entry was 779-830; given the maximum depth sampled was ~810m with no altimetry registering, 830m looks like the best estimate. In another case the log has 2 different values, presumably the bottom was changing through the cast; the entry for the NET cast that followed looks closer to what the altimeter suggests so that was used.

Plots were made to check a selection of altimetry entries and all look appropriate when the CTD got within 20m of the bottom. No problems were found.

The same changes made to the CLN files were also made to the SAM files. This included removing the altimetry header for 3 files that had only 5m samples – the CTD casts were much deeper, but the bottle files do not contain any entries that were within 15m of the bottom. The altimetry header was removed to avoid giving the impression the samples were from near the bottom. After the corrections to the SAM files they were bin-averaged and put through the final merge again and CLEAN was rerun.

The cruise track was plotted and added to the end of this report. No problems were found.

##### 13. Shift

Fluorescence

Normally a shift of +24 records is used to align SeaPoint fluorescence with temperature. Tests were run to see if that setting improves the alignment and it does.

SHIFT was run on all casts to advance the SeaPoint fluorescence by +24 records.

Conductivity

Tests run on 4 casts. The T-S plots were very noisy, especially for the secondary channels, so the choice of the best setting is not completely clear, with results varying from feature to feature. Overall the best results were the same as those for 2014-19 and 2014-60 when the choices were -0.7 for the primary conductivity and +0.7 for the secondary.

SHIFT was run twice on all casts applying a shift of -0.7 records to the primary conductivity and +0.7records to the secondary conductivity.

SALINITY was recalculated after each of the conductivity shifts.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The alignment looks as good as we are likely to achieve.

##### 14. DELETE

Before running DELETE some data were removed from cast #111 because the initial soak period was accidentally archived, including the first drop to 10m and the return to the surface.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds (taken from header)

COMMENTS ON WARNINGS: The only warning concerned file #44 which contains upcast data only.

##### Other Comparisons

Previous experience with these sensors – These sensors were used during 2014-19 and 2014-60.

1. PRESSURE For the 1st cruise the pressure was not recalibrated since tests suggested it was within 0.2db. But during 2014-60 pressure was adjusted by adding 0.6db.

2. TEMPERATURE/SALINITY There was a much larger difference between the temperature channels than usual, ~0.0035C°. Because the secondary salinity was very close to bottles it was presumed that the primary temperature needed recalibration. It was recalibrated by subtracting 0.003C°; this brought the primary salinity into good correspondence with bottle salinity. The same method was used for 2014-60.

3. DISSOLVED OXYGEN The calibration of the DO sensor changed abruptly during the cruise 2014-19, so 2 different corrections were applied:

DOX_Corrected = 1.0245 * DOX_CTD -0.0036 (early)

DOX_Corrected = 1.0326 * DOX_CTD -0.1050 (late)

During 2014-60 the results were similar to the latter part of 2014-19 but the offset was larger. The recalibration applied was:  DOX_Corrected = 1.0329 * DOX_CTD -0.1963

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. There was only one excursion in the temperature data with values between 380 and 420m at station JI22 being slightly lower than the climatology minimum. The deep salinity was near the maximum at many casts and was slightly low at the bottom of station LC04. The few excursions are slight and not systematic, so they do not suggest a calibration problem.

Repeat Casts – The only repeat casts were too shallow for a comparison to be useful.

Post-Cruise Calibration - There were no post-cruise calibrations available at the time of processing.

##### DETAILED EDITING

The first issue is to decide which sensor pair to edit. In processing 2014-19 and 2014-60 (when the same sensors were used) it appeared that the calibration of the secondary temperature and conductivity sensors were likely better than for the primary. However, there was a lot of noise in the secondary data making it harder to edit. For this cruise the delayed salinity analysis makes a judgment about calibration less certain, but the differences between the sensor pairs are close to those from the 2 earlier cruises. And the secondary T-S plots are much noisier. So the primary sensors were selected for editing and archiving.

CTDEDIT was used to clean salinity, and to remove some bad records near the top and bottom of many casts and some records corrupted by shed wakes. Some bad salinity points were removed; these are associated with high deceleration rates, so mostly at deeper casts where the descent rate frequently reached values >1.5m/s.

All casts required some editing.

T-S plots were examined and a little more editing was applied to 3 casts. There remain a few small unstable features, but they look believable. No further editing was applied.

The edited files were copied to *.EDT.

##### Initial Recalibration

Pressure will be corrected by adding 0.6db to all casts.

Based on the bottle comparison described in section 4 the dissolved oxygen channel will be recalibrated using the fit:

DOX_BOT = 1.0325 * DOX_CTD - 0.1960

During 2014-19 it was found that the primary salinity was low by about 0.0033 but that this was likely due to drift in the primary temperature sensor rather than in the primary conductivity sensor. A recalibration to lower the temperature produced salinity that was close to bottle samples; there were trial runs that brought both primary temperature and salinity closer to the secondary channels. The best settings for temperature differed from the best for salinity, so a compromise value was found that had acceptable results for both. For 2014-60 there were fewer bottles and most came from areas with fairly high vertical gradients, so there was a lot of scatter; the results of 2014-19 were used for temperature recalibration and salinity recalculation.

For this cruise the difference between primary and secondary salinity and temperature channels are close to those found for 2014-19. Given the limitations of the salinity comparison from 2014-22 the most reliable information is from 2014-19 so the same correction will be applied to these data.

File 2014-22-recal1.ccf was prepared to add 0.6db to pressure, to subtract 0.0033C° from the primary temperature and to apply the correction

DOX_BOT = 1.0325 * DOX_CTD -0.1960

to the Oxygen:Dissolved:SBE channel. Salinity was recalculated.

CALIBRATE was run on the MRGCLN2, SAM and EDT files using that file.

COMPARE was rerun to check that the corrections were done correctly.

The salinity was found to be low by an average of 0.0028 which matches the secondary salinity.

The dissolved oxygen was found to be high by an average of 0.001mL/L.

(See files 2014-22-dox-comp2.xls and 2014-22-sal-comp2.xlsx.)

30 March 2015 – Corrections made to calibration See note at end of §25

##### Final Calibration of DO

Downcast files were bin-averaged to 0.5db bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

COMPARE was run again. When the differences were plotted against DO concentration, the fit was very noisy. Differences were small for low DO. At the high end of the DO range the SBE DO looks slightly low, on average, but there is a lot of noise. In the middle range, 2mL/L to 5mL/L, the CTD DO is high, by about 0.1mL/L when outliers are removed. When differences are displayed versus depth they are small below 500db. The differences above 500m may be due to incomplete flushing in the presence of high DO gradients. This does not imply a large problem with flushing as a vertical offset of <1m can easily account for such DO differences in the maximum gradients. Given the good comparison below 500m, the rough method used and the scatter in the comparison, no further recalibration is justified.

A plot of differences against pressure was used to make a final study of outliers and to make a rough estimate of the accuracy of downcast CTD dissolved oxygen.

##### Fluorometer Processing and special files for Angelica Peña

##### The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2014-50-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, fixed size=11, was applied to the COR1 files to reduce spikiness in the fluorescence channel. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. A little further editing was applied to a few casts and the Calibrate, Filter and Bin-average steps were repeated.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

Profile and T-S plots were produced at this point to check for errors. No problems were found.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity and

fluorescence data were removed below 510db for events #1-72.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that an offset in

the fit was allowed.

The Oxygen:Dissolved:SBE data had an unusual fit against titrated samples, but

it was similar to that found during the latter half of 2014-19 and all

of 2014-60 when the same equipment was used.

There was an unusually large difference between the CTD temperature channels.

A correction applied to the primary temperature followed by

recalculation of the primary salinity brought the latter into good

agreement with the secondary salinity for both this cruise and the

two which preceded it and used the same equipment. For one of those

cruises there was prompt analysis of salinity bottles, and the primary

salinity was close to bottles after the correction. For this cruise

both salinity channels were lower than bottles, but analysis was

delayed and there was likely some evaporation of samples.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8mL/L from 0db to    75db

±0.3mL/L from 75db to  200db

±0.1mL/L from 200db to 500db

±0.05mL/L below 500db

For details on the processing see the report: 2014-22_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values varied from ~75% to ~150% with most between 90% and 130%. Most of the higher values were near shore in areas with fairly high near-surface DO gradients. Most offshore casts had DO saturations between 90 and 110%, as expected. The near-surface DO gradients are often complex with sub-surface maxima. A few of the casts with the highest and lowest saturations were checked to see how the SBE DO compared with bottles and the two values were close.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts with a PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

REMOVE was run on all casts with no PAR sensor mounted to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. There was no sampling from bottle #5 during event #76 and bottle #3 from event #81, so those lines were removed from the MRG files and the steps that follow that were rerun. A few other discrepancies were found but a check suggests the problems are with the log, not the analysis spreadsheets. For example there were DO duplicates for sample #261, not #260 as shown on the rosette sampling log sheet; the raw data makes this clear.

A header check was run and no errors were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found.

##### Thermosalinograph Data

There were loop nutrients, extracted chlorophyll and salinity samples taken, 1 while stopped and 3 while underway.

a.) Checking calibrations

The configuration files for the 8 casts are identical. One file was renamed as 2014-22-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The 8 files were converted to CNV files using configuration file 2014-22-tsg.con. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. The flow rate was about 1 and mostly steady, though there were a few sudden changes, but they mostly stay within a range of 0.95 to 1.05. The ends of files sometimes have a section of zero flow. The traces look good except for the fluorescence which has some unbelievably high values in files #2-6 and values that are increasing steadily with time.

The salinity signal looks good, with only occasional spikes.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2014-22-ctd-tsg-comp.xls. There were 94 casts which overlapped with TSG files.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found when times were matched. They were all ≤0.0007° and the average for each was 0.0000°. This shows both the times and positions are reliable for both systems.

The loop samples were combined in file 2014-22-loop-tsg-rosette-comp.xlsx. There are a number of problems with the loop samples.

- For loop sample 101 there are nutrients, but no CHL or Salinity. There are samples labelled Loop #1 for Salinity and CHL, but there is no Loop 1 noted in the log book. Checks do not support the Loop 1 samples being from either event #1 or #101.

- The salinity samples were analyzed when the salinometer is known to have malfunctioned, so values are likely significantly too low.

Then 5m rosette salinity and CHL bottle data were added from the 1 cast which occurred at the same time as loop sampling (Loop 2, Event #121).

The two spreadsheets will be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature worked throughout the cruise. The differences between the intake and lab temperatures during CTD casts ranged from 0.31 to 0.58C°, with a median value of 0.17C°. As expected, the heating in the loop increases as the intake temperature decreases.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

1. Intake Temperature The intake temperature is higher than the CTD temperature by an average of 0.08 but the median difference is only 0.01C° and the median of the 11 casts with the best mixed surface waters is 0.005C°. This is likely as close to the CTD as we can possibly expect.

2. Salinity The TSG salinity data are lower than the CTD salinity by an average of 0.068 and a median of 0.064. For the well-mixed casts the difference was only slightly lower at 0.062.

3. Fluorescence The TSG fluorescence was generally much higher than the CTD fluorescence, often by an order of magnitude. This varies greatly from one file to another, with 2 cases of the TSG FL being lower than the CTD: file #6 has TSG fluorescence values about 20% of the CTD fluorescence while those from file #8 are about 50%. All other files have the TSG fluorescence much higher than the CTD sensor values. File #5 has constant fluorescence at 35.529ug/L after the first 160 records. File #6 starts close to the CTD fluorescence but drops quickly with a smooth trace; the last few records have zero flow rate. File #7 had no flow at the end of the cast.

| File # | CTD Fl med | TSG Fl med | Tint-Tctd | Tlab-Tctd | SALtsg-SALctd | FLtsg/ FLctd | # of CTD casts |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 5.792 | 15.938 | 0.039 | 0.326 | -0.071 | 2.761 | 9 |
| 2 | 1.757 | 15.991 | 0.209 | 0.197 | -0.048 | 9.101 | 13 |
| 3 | 2.905 | 29.753 | 0.006 | 0.152 | -0.057 | 10.664 | 15 |
| 4 | 2.401 | 33.521 | 0.008 | 0.168 | -0.064 | 14.005 | 19 |
| 5 | 3.346 | 35.529 | 0.013 | 0.191 | -0.082 | 10.759* | 2 |
| 6 | 2.053 | 0.363 | 0.018 | 0.190 | -0.074 | 0.212 | 7 |
| 7 | 3.183 | 32.068 | 0.079 | 0.224 | -0.097 | 10.099 | 7 |
| 8 | 2.251 | 1.261 | 0.009 | 0.205 | -0.060 | 0.522 | 21 |

* Note that for file #5 the TSG fluorescence was constant after the first 80 minutes; the ratio FLtsg / FLctd would be even larger if the signal had not gone off-scale.

(See 2014-22-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons

There were data from 4 loop CHL, SAL and Nutrient samples but one CHL and one SAL sample could not be compared with TSG data because the source is unknown. In those cases the label said LOOP 1, but there is no such entry in the log book. The label does not indicate depth or Niskin number and is in pencil. The data do not compare well with CTD or TSG data from event #1. There are nutrient samples from event #101 so it seemed reasonable to assume that the CHL and SAL might belong there but the comparisons with the TSG data look poor.

The loop salinity samples were all run at a time when the salinometer is known to have malfunctioned, giving lower values by at least 0.02.

TSG salinity was found to be lower than the loop samples by from 0.03, 0.06 and 0.07.

The TSG fluorescence was equal to 80%, 97% and 57% of the Loop CHL. This seems reasonable. All the loop samples correspond to TSG File #8.

(See 2014-22-loop-tsg-rosette-comp.xlsx.)

- Surface rosette vs TSG

There was only 1 case where we have both rosette TSG fluorescence and a CHL loop sample; we also have a rosette CHL sample for that case. The TSG and the Loop sample were very close; both were about 60% of the rosette sample. There was no rosette salinity sample at 5m for that cast.

- Calibration History

The TSG temperature and conductivity were recalibrated in December 2013 and this appears to be the 4th use since then. During 2014-21 the TSG salinity was found to be lower than loop samples and CTD salinity by ~0.03 but the difference varied with flow rate which was highly variable.  No recalibration was applied due to the variability in the comparisons and the fact that such a large drift in calibration on its first use seemed unlikely. During 2014-18 the salinity was found to be low by 0.014 but due to very noisy salinity traces the comparison was not trusted. For 2014-19 comparisons to CTD casts and many loop samples led to a recalibration of salinity by adding 0.02.

The fluorometer was used during 2013-18, 2014-21 and 2014-18 and 2014-19. The TSG fluorometer was about 1.1 to 1.2 times the SeaBird fluorometer on the CTD when the flow rate was >1 during the first of those cruises and very close to the CTD fluorometer on the second.  During 2014-21 the TSG fluorometer was ~79% of the loop CHL and 65% of the rosette CHL samples. During 2014-19 the TSG fluorescence was about 80% of the CTD fluorescence and perhaps a bit lower for high values. The comparison with loop samples looked typical of this type of sensor with values too high at low CHL, very close in the 1-5ug/L range and too low for higher CHL values.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was generally steady at about 1.0 with a few small shifts and two sections of no flow at the end of files #6 and 7.

3. The temperature in the loop increases by an average of ~0.17Cº between the intake and lab based on comparisons during stops for CTD casts. The heating in the loop varies according to intake temperature.

4. The TSG intake temperature appears to be higher than the CTD temperature by ~0.005Cº No recalibration is justified.

5. The TSG Salinity is lower than the CTD salinity by a median of ~0.064 whether all casts are used or only the casts that are best mixed at the surface. Similar results were found from the 3 comparisons with the loop salinity, but there was a known problem in the analysis of the loop samples so the TSG may be lower than they suggest. There is a lot of noise in the salinity comparison. If we assume that the loop salinity values are too low by about 0.02, then the TSG salinity looks low by from 0.05 to 0.09. The only evidence we have while the ship was underway are two of these questionable loop samples. The “on-station” observations appear to have more salinity spikes; since the spikes are all towards low TSG salinity, there is some reason to suspect that the TSG underway salinity is not as low as the comparison with CTD data suggests. It is likely best to depend on the results of 2014-19 when there were more underway loop samples and more reliable salinity analysis. For that cruise the TSG salinity was recalibrated by adding 0.02.

6. The TSG fluorescence data look bad most of the time with values steadily climbing through many of the files, or being constant (off-scale) and in one case falling steadily with a smooth trace. The source of the problem does not appear to be calibration error as there are sudden changes in the comparison to CTD fluorescence when files change, and in file #8 the data look somewhat low but reasonable. The TSG fluorescence is quite close to the loop chlorophyll but those loops all came from file #8.  Fluorescence should be removed for files #1-7.

f.) Editing –

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT and the following editing was applied:

File #1-4, 6-8 – Salinity was cleaned lightly.

File #6 – The fluorescence looks ok at the beginning of the file but it drops quickly and has a very smooth trace during the drop. There is a short section near the end when it look better but then the flow goes to zero; The temperature and salinity were replaced with pad values for the no-flow section at the end of the file.

File #7 – At the end when there was no flow; temperature and salinity values were replaced with pad values for the no-flow section.

g.) Recalibration

Recalibration of salinity was based on the results of 2014-19 when there were more underway loop samples. File 2014-22-tsg-recal1.ccf was used to add 0.02 to the Salinity:T0:C0 channel.

h.) Preparing Final Files

REMOVE was used to remove the following channels from casts #1-7: Scan Number, Temperature:Difference, Fluorescence:URU:Wetlabs, Conductivity:Primary, Flag and Position:New channels.

REMOVE was used to remove the following channels from cast #8: Scan Number, Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

The flow channel was not removed since it may be useful for further studies of optimal flow rate.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and geographic area and to change channel names to standard names and formats.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series plots and they look fine.

The cruise plot was added to the end of this report.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

CORRECTIONS – March 30, 2015: A post-cruise report on the drift in conductivity and temperature sensors showed that it was the secondary temperature sensor that was out of line, not the primary. So the corrections made in the original processing (subtracting 0.003 from the temperature and recalculating salinity) were inappropriate. The steps from calibration (§17) to the end were rerun to produce new CTD and CHE files. The calibration step was changed to include only a correction for the pressure and Oxygen:Dissolved:SBE channels.

Also, errors in the header comments about salinity bottle inserts were corrected in the CHE files.

# Particulars

8. Bottle at 50m missed

17. CHL sample meant to be from bottle 8 may have come from bottle 9.

43/44 – cast split at 100m point of upcast.

54, 56 – vents were open so cast was repeated.

73. Cable changed on fluorometer (still 3X).

74, 76 - vents were open so cast was repeated.

77. Bottle 5 fired for test only.

81. Bottle 3 tripped by mistake at 750m, Bottle 20 tripped at 500m to replace.

111 – No sample bottle 11 (175m). Initial drop for soak archived by mistake.

PAR ON: 1-29, 40-44, 95-107, 138-168.

# CRUISE SUMMARY  2014-22   CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1. | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #0443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 7May2014 | Factory |  |  |
| Conductivity | 2280 | 7May2014 | Factory |  |  |
| Secondary Temp. | 2663 | 4Jun2014 | Factory |  |  |
| Secondary Cond. | 2754 | 7May2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 997 | 3May2014 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| SeaPoint Fluor | 3642 | Jan. 2014 |  |  |  |
| Pressure Sensor | 77511 | 14Jul2014 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2014-22

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 28Dec13 | Factory |  |  |
| Conductivity | 3363 | 28Dec13 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | Aug12 | IOS |  |  |
| Temperature:Secondary | 0603 | ? | “ |  |  |
| Flow meter | ? | n/a |  |  |  |