## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2014-23

Agency: Ocean Sciences Division

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: June 15, 2014 – June 19, 2014

Processed by: Germaine Gatien

Date of Processing: 30 October 2014 – 12 November 2014

Number of original HEX files: 64

Number of CTD files: 64	Number of bottle casts:	21

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0724) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#1483), a SeaPoint Fluorometer (#3640), a Biospherical QSP-400 PAR sensor (#4656), a surface PAR (#16504), a pH sensor (#0691) and an altimeter (#43281).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order. Photos of the setup were available.

The differences between the primary and secondary temperature and conductivity channels (and hence salinity channels) were noisier than usual, especially above 50m. This looks like it is due to variable alignment with the significance depending on vertical gradients.

Both salinity channels were found to have values lower than bottle samples. Those samples were analyzed about 7 weeks after collection, so there was likely some evaporation leading to higher bottle values. There is also some evidence of poor flushing of Niskin bottles which also leads to higher bottle salinity than CTD, a factor that is most significant in the top 100m where the vertical salinity gradient tends to be higher. These two factors result in the comparison not being reliable enough to justify recalibration. Both T and C sensors had been recalibrated at the factory in late 2013 and this appears to have been their first use since then, so little calibration drift is expected. The secondary channels were selected because they were less spiky; no recalibration was applied to the salinity.

The fluorescence values were all lower than extracted chlorophyll samples, being much lower at high values of CHL. These results are unusual. This is a new sensor; the CTD technician was informed that it should be tested before another deployment.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1    mL/L from 0 to 20db

±0.3 mL/L from 20db to 50db

±0.15 mL/L from 50db to 200db

±0.10 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real. A few clearly bad values were replaced with pad values.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen, extracted chlorophyll and salinity.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained, but only the pressure sensor has been used for other cruises since its last factory calibration.

The XMLCON files did not change through the cruise.

The calibration constants were checked for all instruments and the DO sensor calibration parameters were out of date. They were replaced with calibrations from 17 January 2014. The date of calibration was added to the pH entry. The corrected file was saved as 2014-23-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2014-23-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. The temperature and salinity channels are farther apart than usual on downcasts down to about 50m. The upcast traces are noisier but closer. In areas with a low vertical gradient the downcast traces are close. The problem is less notable below 50m but since gradients tend to be lower at those levels it is a little hard to be certain about that. Later, when conductivity channels have been aligned, T-S plots may make it clearer which sensor pair is behaving better. Unfortunately, the problem will not be addressed by the bottle comparison because the upcast data do not seem to be affected and there was only one bottle above 50m.

The altimetry looks useful and fluorescence, transmissivity, PAR and SPAR look normal though PAR is usually much lower than SPAR. The pH signal has a section of 0 values in cast #9, so all casts should be checked later.

The descent rate of the CTD was fairly high and steady for most casts, but there is a fine “jitter” in the traces. There were notes in the log about noise from the winch.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2014-23-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files to check for outliers. No editing was found to be necessary.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2014-23-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-23chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2014-23chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-23oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2014-23oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2014-23SAL.xls. The analysis was done within 41-45 days of collection. The files were simplified and saved as 2014-23sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-23nuts.xls. This includes a precision study.

Then the file was simplified, reordered on sample numbers and saved as 2014-23-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A few problems were found and corrected:

- Event #76 has a CHL sample at the bottom (sample #246) that is not shown on the rosette sheet, has no duplicate and has a high value (~9ug/L) not likely to be seen at 177m. The analyst found this entry was in error, a duplication of another sample with the wrong sample number.

- Event #39 has 3 CHL samples but only 1 shown on the rosette log sheet. Since that is the only cast with only 1 box ticked, it seems most likely that the rosette sheet is wrong. The CHL samples look reasonable.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure shows a lot of scatter with the CTD generally being lower than the bottles. The two cases of bottles being lower are associated with very noisy CTD salinity. When outliers are removed based on differences >0.01 and standard deviation in the CTD data being >0.0008, the primary salinity was found to be low by an average of 0.0033 and the secondary salinity low by 0.0042.

There are 3 reasonable explanations for why CTD salinity would be reading lower than bottles:

- Calibration drift – It is unlikely that both sensor pairs would have drifted this much since recalibration. There is no record of them being used on another cruise.

- Poor flushing of Niskin bottles – Poor flushing means that bottles will contain water from deeper than the firing level, with higher salinity than the ambient value. Flushing problems are worst when there is little motion of bottles during bottle stops and that was the case for many bottles during this cruise, though as noted earlier there is a fine-scale “jitter” in the pressure. The stops for the bottles with the lowest differences are associated with a little more CTD motion than those with the largest differences. This shouldn’t cause large errors at depth, but errors on the order of 0.005 are probably easily explained this way. Larger errors are expected for shallow bottles because the gradients tend to be higher and in this comparison the largest errors are near the surface.

- The samples were analyzed about 7 weeks after collection and previous experience suggests that this can lead to evaporation and hence bottle salinity values that are slightly higher than they should be. This effect tends to be random rather than depth or time dependent. The comparison does show some randomness.

It is likely that the problem is due primarily to poor flushing of Niskin bottles with some evaporation of samples adding to the effect.

Whatever the cause, there is no justification for recalibration of salinity.

For full details for the COMPARE run see file 2014-23-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter. When a few large outliers are removed and then further data are removed based on residuals, the trendline found are:

CTD DO Corrected = CTD DO * 1.0409

or

CTD DO Corrected = CTD DO * 1.04 + 0.0037

depending on whether an offset is allowed or not.

The R2 values for those fits are the same. The fits tend to remove a lot of the near-surface data where the CTD values are higher than the titrated samples. Given the nature of these sensors we expect the CTD DO to be lower. While local reversals in DO gradients accounts for some scatter, a likely cause of outliers with CTD looking higher than expected is poor flushing of Niskin bottles in the presence of a high DO gradient. So using a fit that ignores most of these near-surface bottles is quite reasonable.

The outliers are not extreme with the largest being the surface sample from cast #73. The analyst flagged this sample as 3 with the comment “Niskin top cap did not seal.” It might be justified to change this to 4 except that the CTD data is very noisy and flushing is likely a problem. So the flag was left unchanged. None of the samples flagged 4 look notably out of line, but given the limitations of the comparison there is no justification in changing them.

Plots of Titrated DO and CTD DO against CTD salinity were examined. The only outliers were minor and associated with high DO gradients.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

The fluorescence values were all lower than extracted chlorophyll, being much lower at high values of CHL. These results are most unusual. Having the wrong cable would not explain this result as higher fluorescence would require a 1X cable and we don’t have one of those.

The dark value of the fluorometer is <0.1ug/L.

For full details of the comparison see file 2014-23-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests show that a +3s alignment brings the downcast DO and primary temperature into reasonable alignment. If the vertical offset between the upcast and downcast DO and Temperature are compared a higher shift might be selected, but this step is of importance to the downcast, not the upcast.

ALIGNCTD was run using +3s. (Note the alignment was adjusted later as described in section 12.)

##### CELLTM

The upcast and downcasts differed so much that the usual tests are unlikely to assist in choosing the best parameters. Tests were run using the default settings of (α = 0.0245, β=9.5) on both channels, and the results were compared with a run with no CELLTM correction. The results show that the correction produced much better correspondence between upcast and downcast T-S traces.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE  and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 2 of the deeper casts to examine differences between sensor pairs.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-23-0020 | 320 | -0.003 VN | -0.0005 | -0.002 | Mod, F. steady |
| 2014-23-0037 | 400 | -0.002 XN | -0.0003 | -0.001 | High, V.steady |

Even for casts this shallow the differences are noisier than expected when the descent rate is very steady. The temperature differences are very high near the surface and gradually approach 0, but they are very noisy so that the values listed are only very rough estimates. As was noted earlier the large differences largely disappear around 50m but that now appears likely to be due to reduced gradients rather than the problem clearing up. The differences never settle down and the sign changes as the gradients change sign. This does suggest an alignment problem.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book and a few problems in station names were corrected in the CLN files (Events #1, 31, 48) and the MRG files (Event #1). It was also discovered that cast #5 was missing because there was no HEX file. A search of the shipboard computer failed to find it, so it likely was never saved.

A header check was run. No problems were detected.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 1.5db which is reasonable for the Vector.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for most casts and the headers look appropriate. Depth readings were checked against the log and there were many discrepancies. Checks were made using bottom pressure (converted to depth) plus altimeter readings to make an estimate. In general the log entry looks more accurate for near-bottom readings; the log is probably accurate for the beginning of the cast. Since the value from the bottom is more useful, the header entries were adjusted in the CLN and SAMAVG files, if the difference was >2db (20 casts). The SAMAVG files were merged again with the MRGCLN1s files, and CLEAN was then run to remove SeaBird headers and comments from the secondary file.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +40 records and +60 records and +50 looks like the best choice overall.

SHIFT was run on the pH:SBE channel using a setting of +50 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The results are a little confusing, but it looks overall that the original alignment went a little too far. So SHIFT was run using a setting of -12 records. The alignment looks better after this step though it does vary with depth. The cumulative shift is +60records or 2.5s.

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results were found with -0.6 records for the primary and -0.3 records for the secondary.

SHIFT was run using -0.6 records for the primary conductivity and -0.3 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Temperature, conductivity and dissolved oxygen sensors were all recalibrated recently and there is no record of them having been used since then. The pressure sensor was recalibrated in 2009 and may have been used for other cruises since then.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.

Temperature values were near the minimum below 150m on the western side of northern Strait of Georgia. In only a few cases was it lower than the climatology minimum. Salinity had only 2 slight excursions from the climatology, in one case at the bottom of a cast and in the other near the mouth of Juan de Fuca which are both areas of high variability. None of these excursions are considered evidence of calibration drift. The 3-standard deviation climatology is too severe for this region.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### CHANNEL CHOICE STUDY

COMPARE shows that the primary salinity is slightly closer to the bottles than the secondary salinity, but problems with bottle flushing and evaporation of samples are likely to mean the bottle values are too high. So this does not make it clear which sensor pair have the more accurate values.

It is likely that this was the first use of these sensors since there last factory calibration so we do not expect much drift in either pair.

The vertical offset between the primary and secondary traces was investigated. It looks fairly consistent on the downcasts with the primary leading the secondary by 5 or 10 scans. But during the upcast the secondary leads. This might suggest an actual physical offset in the sensors. Whichever pair lead also appear to have noisier data, though this is hard to judge for the upcast since that data is always rather noisy. An examination of downcast T-S plots using the two channel pairs shows that the primary pair is usually noisier and more data would be lost in editing. The primary noise sometimes looks like an unsteady flow rate so perhaps there was a plumbing problem. The worst sections are marked by spikes in the primary salinity. At times the descent rate, while high and fairly steady, does have a jitter in it of varying size. How any of these observations are linked to any of the others is not obvious, but the secondary sensors look like the better choice for archiving.

##### DETAILED EDITING

Since the secondary temperature and salinity channels seem a little less spikey, those channels were selected for archiving, and hence, editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts.

All files required some editing.

##### Initial Recalibration

There is no evidence of a problem with the pressure calibration.

While the comparison between the two salinity channels suggests both are low, the bottles data are likely affected by poor flushing of Niskin bottles and possibly by evaporation of samples. So the CTD salinity channels are likely higher with respect to bottles than they appear. Since the temperature and salinity sensors were recalibrated recently and the two salinity channels differ by only 0.001 it is likely there has been minimal calibration drift. There is certainly no reliable evidence on which to base a recalibration of salinity.

CALIBRATE was run using file 2014-23-recal1.ccf to apply the following correction to DO in the SAM and MRGCLN2 files:

CTD DO Corrected = CTD DO * 1.04 + 0.0037

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was +0.0003mL/L but the standard deviation is 0.03mL/L. (See file 2014-23-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.03mL/L, but the results vary with depth with lower differences at greater depth. The descent rate was quite steady for many of these stations which may decrease the efficiency of flushing. The effect of poor flushing (in profiles where DO is decreasing with depth) is that CTD DO looks higher than the titrated samples, except at the bottom where it makes the CTD DO look lower than bottles. For this cruise some of the bottom bottles do show the CTD DO looking slightly low. So it is not clear that the differences are due to errors in the CTD calibration.

No further calibration was applied.

See 2014-23-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2014-23-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some unstable features, but that is normal for this region in near-shore and well-mixed casts. All profiles look fine except for pH which has some odd spikes. A text editor was used to replace bad values with pad values in casts 3, 6, 9, 15 and 23. The transmissivity was very low in the top 5m of cast #48 but the log mentioned that the water looked very brown so these values are reasonable. Nighttime PAR:Reference values are sometimes high, likely due to lights shining on the sensor.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

There were sections of bad pH:SBE:Nominal data, so values that were

obviously out of line were replaced with pad values. These data

are otherwise unedited.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1 mL/L from 0 to 20db

±0.3 mL/L from 20db to 50db

±0.15 mL/L from 50db to 200db

±0.10 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2014-23-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok except that there are some unexpectedly high nighttime PAR:Reference values; this is likely due to lights shining on the sensor. The data are nominal so were left unedited.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values varied from ~70% to 140%. The lowest values were in the area of Haro Strait and the eastern part of Juan de Fuca Strait, both areas of active mixing. The highest values were in the northern and middle parts of the Strait of Georgia where the DO gradients in the top 20m were higher. The saturation values do not suggest any problem with calibration.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars (Notes from Log Book)

1. Test cast in Saanich Inlet to 50m.

2. Strong currents Descent/Ascent rate <0.1m/s below 100m. Stopped descent at 230m, winch making strange noises on recovery.

4. Strong tide – drift through station to reduce wire angle.

5. HEX file not saved.

12. Label printer died.

15. Ship lab printer networked to CTD laptop used for labels.

28. Ship surrounded by tideline flotsam in Baynes Sound.

44. Switched winch.

48. Rosette out of sight at surface; very brown water.

62. Drifting not keeping station due to currents, traffic.

71. Position change due to heavy ship traffic.

76. Chains on.

77. Chains off.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0724 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4397 | 28Dec2013 | Factory |  |  |
| Conductivity | 2992 | 31Dec2013 | Factory |  |  |
| Secondary Temp. | 4402 | 28Dec2013 | Factory |  |  |
| Secondary Cond. | 2984 | 31Dec2013 | Factory |  |  |
| Transmissometer | 983DR | 12June2014 | IOS |  |  |
| SBE 43 DO sensor | 1483 | 17Jan2014 | Factory |  |  |
| PAR | 4656 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0691 | 29Dec2010 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 90559 | 27May2009 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |