## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2014-26

Agency: SAFE

Location: West Coast Vancouver Island

Project: Strait of Georgia Juvenile Salmon

Party Chief: Neville C.

Platform: W.E. Ricker

Date: 15 September 2014 – 6 October 2014 & 5 November 2014 – 7 November 2014

Processed by: Germaine Gatien

Date of Processing: 21 April 2015 – 27 April 2015

Number of original CTD HEX files: 42		Number of CTD files: 41 (+4 upcast files)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. Attached were an SBE 43 DO sensor (#1176) and a WET Labs ECO_AFL Fluorometer (#2215) on the primary pump and an unpumped SBE18 pH sensor (#0852).

The deck unit was a Seabird model 11 (S/N 0242).

The data logging computer was OSD Field Computer PAC02511. SeaSave version 7.22.4 was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log included a list of participants and an equipment list although the pH sensor was not mentioned. There are notes about problems encountered including thee fact that bottom depths were recorded at deployment time but frequently changed because of drift during the cast.

File 2014-26-0002 contained only a little surface data and was not processed.

Data acquisition started late for many casts; where it started below 10db the upcasts were used to produce a full profile. The upcast files have names with the format 2014-26-9***.ctd. The quality of upcast data is lower than from downcasts, so the partial downcast files are also provided.

There was no salinity calibration sampling; the primary salinity was recalibrated by adding 0.002 based on an assumption that most of the drift implied by the post-cruise calibration had occurred at the time of this cruise.

The fluorometry traces show little spiking so the sampling rate was presumed to have been set low. This rate is adjusted within the sensor itself and is set before the cruise. Normally we have a lot more detail in this channel and use a filter to reduce it. No filter was applied to these data.

There was dissolved oxygen sampling for only the 1st of the 8 cruises that have used this sensor between factory visits. The results of cruise 2014-01 were used to recalibrate DO. The post-cruise factory calibration suggests a drift that is reasonably close to the correction applied.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained and problems noted. There are some notes that are not clear, but may refer to the soak period. There is no evidence of acquisition during the soak period, so no action is needed. See the Particulars section at the end for notes about the log comments.

The configuration stayed the same throughout the cruise; one file was saved as 2014-26-ctd.xmlcon.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity, fluorometer and DO sensors were obtained.

The calibration constants were checked for all instruments. The only change made was to the dissolved oxygen factor E (from 3.6 to 3.75) to reflect tests run using 2014-01 data to fine-tune the hysteresis setting. This will not affect this cruise since the hysteresis correction will not be needed, but it was changed to avoid confusion in the future.

##### Conversion of Full Files from Raw Data

All files were converted using con file 2014-26-ctd.xmlcon with the hysteresis correction turned off.

A few casts were examined and all expected channels are present.

There are many pressure spikes at the end of casts.

The two temperature and conductivity channels track reasonably well during the downcasts though they are rather far apart; on upcasts they are noisier.

The fluorometer, pH and dissolved oxygen traces look normal.

The descent rate varies from very quiet to noisy.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A setting of +4s was found appropriate to advance the DO voltage for 7 other 2014 cruises using this sensor, both before and after this one. Checks of the offset between distinctive features in T and DO suggest that +4s is right for this cruise too.

ALIGNCTD was used to advance the DO Voltage by 4s relative to the pressure.

Plots were examined before and after that step and the results look reasonable.

##### CELLTM

A setting of (α = 0.03, β=9) was found appropriate for the thermal mass correction for other 2014 cruises using these sensors, both before and after this one.

CELLTM was run using (α = 0.03, β=9) for both the primary and secondary conductivity.

Tests were run on a few casts but were hard to interpret because the upcast and downcast traces were often very different, but in areas where the profiles were similar CELLTM did improve the correspondence on a T-S curve.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

When these sensors were used during 2014-01 there was so much noise in the data that no reasonable comparison could be made between channels. For cruises 2014-16, 2014-17, 2016-36, 2014-08, 2014-51, 2014-66 and this cruise, the casts are all fairly shallow, making comparison less reliable than if we were looking at very low gradient zones as found in deeper water. The deepest sampling was about 250m for most of the other cruises, so some casts from each of the other cruises were compared with 2 from this cruise at 240db. The differences were:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-16-0224 | 240 | +0.0003 | +0.00013 | +0.0011 | Fairly steady |
| 2014-16-0087 | 240 | +0.0004 | +0.0002 | +0.0017 | Fairly steady |
| 2014-17-0155 | 240 | -0.0001 | +0.0004 | +0.0006 |  |
| 2014-17-0170 | 240 | -0.0002 | -0.00005 | -0.0004 |  |
| 2014-36-0016 | 240 | -0.0003 | -0.00003 | ~0 noisy | Fairly steady |
| 2014-36-0141 | 240 | -0.0002 | -0.0001 | -0.0013 | Fairly steady |
| 2014-36-0174 | 240 | -0.0002 | -0.0002 | -0.0017 N | Fairly steady |
| 2014-08-0005 | 240 | -0.0006 VN | -0.00025 N | -0.0024 N | Noisy,moderate |
| 2014-51-0003 | 240 | -0.0003 | -0.00058 | -0.0061 | Noisy,moderate |
| 2014-51-0013 | 240 | -0.0006 | -0.00057 | -0.0060 | Noisy,moderate |
| 2014-26-0021 | 240 | -0.0003 | -0.0005 | -0.005 | Noisy,f high |
| 2014-26-0083 | 240 | -0.0003VN | -0.0004N | -0.004 VN | VNoisy,high |
| 2014-26-0162 | 240 | -0.0002N | -0.0004N | -0.0042 N | VNoisy,high |
| 2014-66-0045 | 240 | ~0 N | -0.00043 | -0.0044 | Fairly steady |
| 2014-66-0159 | 240 | -0.0001 | -0.00058 | -0.0062 | Fairly steady |
| 2014-66-0181 | 240 | ~0 | -0.00064 | -0.0072 | Fairly noisy |

The differences from this cruise are similar to those of 2014-66 which followed, a little lower than the previous cruise and mostly higher than the first 4 cruises. They all used mostly the same equipment. It will be seen later that post-cruise calibrations show that the 2 salinity channels differed by only 0.002 after 2014-66. There may be differences in the conductivity channels while moving that are due to alignment variations rather than calibration drift.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. As is typical of this type of fluorometer, there were some negative values, so Fluorescence values <0ug/L were replaced with pad values.

##### Checking Headers

The header check was run and the only problems noted were pressure spikes, but they are likely at the end of the files right at the surface. It is also noted that there are some low pH values, but these may be related to the pressure spikes.

The surface check gave an average value of 3.8db which is reasonable for the Ricker. There were many negative pressures at the end of most casts with a lot of spiking in pressure. Where the pressure is steady and very low (~0.5db), the associated conductivity values were very low. They are seen after the pumps were turned off. There is no evidence of problems in the pressure calibration.

The cross-reference check was compared with the log book. There were a few discrepancies:

- Event #2 – the time is about 2 hours later than the time in the log book. It is possible this was just an issue with the computer time, but the file contains only a little surface data, so it is possible that the original data got overwritten during a test run later. This cast will not be processed.

- The station name in the headers was incorrect for events #16, 207, 208 and 210. These were corrected to match the log.

- The times in the log and header differ by roughly 1 hour for the casts in November. It is likely that due to the change from PDT to PST between the two legs there was some confusion in converting local times to UTC. No change was made to the headers as they are most likely correct.

The cruise track was plotted and added to the end of this report. No problems were found.

##### Shift - Alignment

Fluorescence

A shift of +12 is the usual setting used to align the Wetlabs ECO fluorescence with temperature and it was used for 2014-36, 2014-51 and 2014-66 when the same sensor was used. Tests were run on a few casts and the results look good.

SHIFT was run on all casts to advance the ECO fluorometer by +12 records.

Conductivity

For the 6 other 2014 Ricker cruises using these sensors, the best results were found using -0.7records for the primary conductivity. For the secondary -0.9 was best for the first 3 cruises but -1.3records worked better for 2014-51 and 2014-66. Tests were run to see which of the secondary settings worked better for the current cruise and the 2014-66 settings did.

Two runs of SHIFT were used to advance the primary by -0.7 records and the secondary by -1.3 records. Salinity was recalculated after the shifts.

pH

Because of hysteresis in the pH signal, tests are best run on casts with distinctive features in the pH traces. This was done by finding the distance between such features in the upcast and downcast and comparing that with the temperature offset after applying a variety of shifts. Settings between +50 and +60 records produced the best results overall during 2014-36 with the same equipment, and the setting +55 looked good for 2014-51 and 2014-66 and these data as well.

SHIFT was run on all casts with the setting +55 records.

Dissolved Oxygen

Checks were made of the alignment of DO and temperature to see if the shift applied earlier (§5) had done a good job; it worked very well for some features but was too little or too much for others. The earlier adjustment appears to be the best we can achieve overall.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds (taken from header)

COMMENTS ON WARNINGS: There were many warnings but all were from the upcast near the surface so will have no impact on the DEL files.

For files #16, 70, 104 REVERSE was run and then DELETE to get a REVDEL file that may be useful to get the top sections that were missed due to data acquisition starting late.

The *.revdel files were renamed as 2014-26-9016.del, 2014-26-9070.del, 2014-26-9104.del and 2014-26-9119.del. The 9 indicates upcast data. Both downcast and upcast will be archived – the former are of higher quality and the latter more complete.

##### Comparisons of CTD and bottle salinity samples

There were no salinity calibration samples.

##### Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity and temperature sensors were both recalibrated in late January 2013 and they were used for 8 cruises before the next factory visit. The comparisons with bottles, where there was sampling, were not fully trusted. The best sampling was from 2014-01 when the primary salinity was found to be low by about 0.004, but even for that cruise there was some concern that there might be some evaporation of samples and some small effect from poor flushing especially in light of that being the first use of the sensors since the previous factory visit. The primary sensors were only used for one cast so were recalibrated to match the secondary which was not corrected.

Dissolved Oxygen:

The DO sensor was used for 8 cruises between factory calibrations, but only 2014-01 included DO sampling, so the fit from that cruise was used for all the others (slope/offset = 1.0502/+0.0142). During 2014-01 data hysteresis tests were done and factor E was fine-tuned.

Pressure:

The pressure sensor was recalibrated in May 2013 and the factory offset was found appropriate for the 7 other cruises since then.

Historic ranges – Where local climatologies were available, all temperature and salinity data fell within them, except for one cast with temperatures slightly high and salinity slightly low at the surface.

Repeat Casts –

There were no known repeat casts though some may be very close. However, being shallow and in areas of active mixing, comparing them would not be helpful.

Post-Cruise Calibration

Post-cruise calibrations were obtained from March 2015 for the primary conductivity and from December 2014 for the other temperature and conductivity sensors. They show that the primary and secondary temperature sensors were reading slightly low, by 0.0008C° and 0.0007°, respectively. The conductivity sensors had drifted so that the primary salinity was low by 0.0026 and the secondary low by 0.0046 salinity units. The combined effect produces primary and secondary salinity low by about 0.002 and 0.004psu. Since this was the 7th of 8 cruises between calibrations, it is assumed that most of the drift occurred by the time of this cruise. So salinity is likely low by ~0.002.

There are post-cruise calibrations for both pH sensors but the data provided are from after modifications to the sensors, so not appropriate for data collected earlier.

The pressure post-calibration shows the pressure may be slightly low by no more than 0.8db, but this calibration was done as part of some equipment modifications that might possibly affect the offset.

- DETAILED EDITING

The first issue is to decide which sensor pair to edit. The primary T-S plots look a little less spiky and the primary temperature and salinity have been selected for the other 2014 Ricker cruises with these sensors. So the primary channels were selected for archiving for this cruise too.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records.

All casts required some editing except for events #153 and 207. Notes were entered in the headers about editing applied.

Because downcast acquisition began late for cast #16, 70, 104 and 119, both the downcast and upcast files were edited. A comparison of the upcast and downcast files shows there are significant differences, but there is a lot of variability near the surface, so some differences may be due to real change. It is expected that the upcast data are of lower quality due to the CTD carrying deeper water with it.

- Initial Recalibration

Using the post-cruise pressure calibration from December 2014 would produce pressure higher by about 0.8db, but it is possible it was affected by modifications to equipment that were carried out at that time and there may have been drift between this cruise and December 2014. There was no evidence from these data nor from the cruise that followed that the pressure was reading low. No recalibration will be applied.

The post-cruise calibration indicates that the primary salinity was low by about 0.002 in March 2015. This was the 7th of 8 cruises between factory visits, so it is reasonable to estimate that most of the drift had occurred by the time of this cruise. The primary salinity will be recalibrated by adding 0.002.

The only information available about the SBE dissolved oxygen sensor since it was last recalibrated comes from 2014-01, so the correction applied to that data set will also be used for this cruise:

DOX_CTD Corrected = DOX_BOT = 1.0502* DOX_CTD +0.0142

File 2014-26-recal1.ccf was prepared to apply the above correction to Oxygen:Dissolved:SBE and to add 0.001 to the primary salinity. CALIBRATE was run on the EDT files.

##### Fluorometer Processing

At this stage a median filter, fixed size=11, is usually applied to reduce spikiness in fluorescence channels, but examination of all casts shows little spikiness, so this step was skipped. It looks like the sampling rate was set low on the sensor just as it was for other 2014 cruises using this sensor.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag

Profile plots were produced at this point to check for errors. No problems were found.

T-S plots were produced; the only noteworthy unstable features are seen in the 3 casts (#190, 200 and 204) in the Discovery Channel area that likely reflect real conditions.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

----------------------------------------------------------------------------------

Data Processing Notes:

----------------------

Fluorescence and pH data are nominal and unedited except that some

records were removed in editing temperature and salinity.

Dissolved oxygen was calibrated using the results of cruise 2014-01 when the

same sensor was used. There was no calibration sampling from this cruise.

Recalibration of the salinity channel was based on a report from the factory on

drift in the temperature and conductivity sensors as of 23 December 2014 and

3 March 2015, respectively. That report shows the primary salinity had drifted

low by about 0.002. This was the 7th of 8 cruises between the factory

checks so it was estimated that about the CTD salinity was reading low by

about 0.002.

For casts that have no data above 10m due to acquisition starting late, upcast

files were created with names with the format 2014-26-9***.ctd; the upcast

data are considered to be of lower quality due to the CTD package

entraining some deeper water through which the sensors must past.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were collected.  Calibration is required for each cast to

get absolute values, although general trends within a cast are likely real.

For details on the processing see the report: 2014-26_Processing_Report.doc.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

The track plot looks fine.

The sensor history files were updated.

##### Dissolved Oxygen Surface Saturation

Dissolved Oxygen saturation was derived and plotted. The surface saturations varied from about 60% to 170%. The values in the more open parts of the Strait of Georgia were mostly between 95% and 115%. Both the highest and lowest values were close to shore, with the lower ones mostly in Puget Sound or inlets or areas of tidal mixing. We have no DO sampling to confirm these results but they look reasonable.

Particulars

2. Surface data only. Not processed.

16. Log says archiving started at 20m depth –actually 70m.

60. Note in log says “started archiving during calibration, missed the 250 max note, went to 280” - not clear what this means and the cast looks normal with no data from soak period.

70. Started archiving at 30m.

104. Archiving started at 20m.

146. Went to 10m twice before calibration of 1min. Not sure what this means – cast looks normal.

200. Log Note: Communication with CTD interrupted at about 5m. Restarted at 5m and continued. The first 5m are in file 199. PROCESSING NOTE: No file 199 found.

General: Fluctuating bottom depth through some casts – log record generally is depth at deployment time.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | No | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 31Jan2013 | Factory | Mar2015 | Factory |
| Conductivity | 2128 | 29Jan2013 | Factory | Dec2014 | Factory |
| Secondary Temp. | 2374 | 31Jan2013 | Factory | Dec2014 | Factory |
| Secondary Cond. | 3396 | 30Jan2013 | Factory | Dec2014 | Factory |
| SBE 43 DO sensor | 1176 | 5Mar2013 | Factory | 23dec14 | Factory |
| WetLabs ECO-AFL/FL | 2215 |  |  |  |  |
| SBE18 pH | 0852 | 21April2011 | Factory | 31Dec14 | Factory |
| Pressure Sensor | 75636 | 25Mar2013 | Factory | 17Dec14 | factory |