## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 12Feb. 2025 | Removed duplicate Salinity:bottle channel, fixed units in salinity channel. |
| 1 April 2015 | Correction to header comment about salinity bottles in CHE files. G.G. |
| 19 November 2014 | Note added to end of report on factory post-cruise sensor calibrations. GG |
| 21 July 2014 | CTD salinity in CTD & CHE files recalibrated. See note at end of report. GG |

## PROCESSING NOTES

Cruise: 2014-27

Agency: WVL, Marine Environment and Habitat Science, Vancouver, B.C.

Location: Strait of Georgia

Project: PARR Shellfish Carrying Capacity

Party Chief: Sutherland T.

Platform: Vector

Date: April 7, 2014 – April 11, 2014

Processed by: Germaine Gatien

Date of Processing: 26 June 2014 – 14 July 2014

Number of original HEX files: 47 (1 with pumps off)

Number of CTD files: 46	Number of bottle casts:	23

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0585) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#1119), an ECO Fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#4615), a surface PAR (#16504), a pH sensor (#0692) and an altimeter (#1204).

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log was missing a personnel list. The rosette log sheets and analysis logs were in good order. No cruise report was available at the time of processing. There were no photos of the setup, but it is believed to be identical to that of cruise 2014-06.

CTD #0585 was used for 4 cruises in spring 2014; this was the 3rd in the series.

The salinity analysis was completed within 2 weeks of the beginning of the cruise. The two salinity channels differed by about 0.004. Comparisons between CTD salinity and bottle samples were hampered by the fact that most samples came from within 4m of the bottom where flushing problems render comparison untrustworthy. Using the few bottles that look reliable, it appears that the secondary salinity is close to the bottles. It is recommended that calibration samples be taken at least 10m above the bottom.

Pumps were not turned on for event #3. A CHE file was prepared with pumped CTD channels replaced by pad values. No CTD file was prepared.

The comparison of CTD dissolved oxygen and titrated bottle DO was noisy, likely due to flushing problems near the surface and near the bottom of casts. When outliers are removed the results were reasonably close to the fit found for 2014-06, so the results of that cruise were used to recalibrate the CTD DO channel. The stops for bottles at the surface were often too short for shed wake corruption to settle down; some were as short as 10s towards the end of the cruise.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8 mL/L from 0 to 10db

±0.4 mL/L from 10 to 50db

±0.1 mL/L from 50 to 100db

±0.06 mL/L below 100db

There was no chlorophyll sampling during this cruise. During 2014-06 it was found that the ECO fluorometer had values that were ~60% of extracted CHL values for CHL>3ug/L. At the lowest values of the CHL range the ECO tends to read higher than or close to CHL.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets and final spreadsheets for dissolved oxygen, nutrients, extracted chlorophyll and salinity. DIC samples were not yet available.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained.

The calibration constants were checked for all instruments and the only errors found were in some dates. One corrected file was saved as 2014-27-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2014-27-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. The temperature and conductivity channels are close on downcasts, but the upcast traces are noisy. The altimetry looks extremely noisy at the bottom, though the trend above may enable an estimate if the header reading algorithm fails.

The fluorescence, transmissivity, DO, pH and PAR look normal. SPAR mostly looks ok but first few casts look like there is no signal.

The pumps were never turned on for cast #3.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2014-27-ctd.xmlcon.

The files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. There was some noise in a few points of the secondary temperature at the bottom of cast #3, but the pumps were off, so these data will be removed later anyway. A bottle file will still be prepared as the pressure is fine and there was sampling.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2014-27-bot-hdr.txt; it may need further editing to reflect problems found during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-27oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2014-27oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2014-27SAL.xls. The analysis was done within 4-9 days of collection. The files were simplified and saved as 2014-27sal.csv, which was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-27nuts.xls. This includes a precision study.

Then the file was simplified, reordered on sample numbers and saved as 2014-27-nuts.csv. The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files, to produce MRG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions and it was discovered that the nutrients were missing from the first file because the event number had been entered as 0. That is what the rosette log sheet said, but the daily science log shows it as 1, as expected. When that was corrected and the merge process repeated, the exported spreadsheet looked fine.

NOTE: The CTD pumps were off for event #3, so the pumped CTD channels were replaced with pad values in the MRG file and a note of explanation was added to the headers.

The MRG files were put through CLEAN to produce MRGCLN2 files.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure shows a lot of scatter. If the problem were due to evaporation or poor flushing, we would expect then the bottle values would be higher than expected, making the CTD values look lower than they really are. We don’t expect evaporation problems because the analysis was done quickly. Incomplete flushing could well lead to poor values, especially near the surface, but again we would mostly expect the bottle values to be too high. For this cruise most of the outliers are in the opposite direction with the CTD values looking higher than bottles. The explanation for this is likely that the bottles were taken too close to the bottom. The general rule we use is to not take bottles at the bottom but come up roughly 10m. In this case, the bottle sampled for salinity was the 2nd bottle fired. But for many of the shallow casts, that was only 2m above the bottom. Incomplete flushing close to the bottom appears to have quite a different character, with bottles mostly reading too low, so that the CTD looks high. This is a pattern that has been seen in the past when samples are taken at or close to the bottom.

The last 6 casts had salinity bottles taken a little higher in the water column. When only those bottles are included, the 3 below 160m have differences in line with what was observed during 2014-06. The other 3 show the CTD being lower than the bottles by more than expected based on 2014-06. This is consistent with incomplete flushing which tends to be significant closer to the surface due to higher salinity gradients. The 6 casts were examined in detail and the following was observed during the stops for bottles with salinity sampling:

- Event #65 – There was a shed wake at the beginning of the stop that includes salinity equal to the bottle salinity. This stop was about 5m above the bottom. Difference is much larger than the average from the previous cruise.

- Event #67 – There was a shed wake at the beginning of the stop that includes salinity equal to the bottle salinity. This stop was well above bottom. Difference is much larger than the average from previous cruise.

- Event #73 – There was a shed wake at the beginning of the stop that includes salinity equal to the bottle salinity, but the shed wake salinity is not very different from the value after the salinity settled down. This stop was well above bottom. Difference similar to the previous cruise.

- Event #74 – There was variation of salinity on the order of 0.0025 just before the stop. This cast was only 4m above the bottom. Difference similar to the previous cruise.

- Event #75 – This was the stop with the steadiest salinity. There was no shed wake. The stop was well above the bottom. Difference similar to the previous cruise.

- Event #76 – The bottle value was seen in a shed wake at the beginning of the stop. The stop was about 4m above the bottom. Difference is much larger than average from the previous cruise.

The average difference using the 3 bottles with no shed wake corruption or very little, indicates that the primary salinity is low by 0.0040 and the secondary is high by 0.00001. The 3 casts are also the deepest of those that are at least 4m off the bottom. These results are close to those of 2014-06.

There is one outlier in the secondary differences that is not notable in the primary, but that came from the cast with the pumps turned off. The only bottle below 60m that suggests the CTD salinity was higher than the bottles came from only 2m above bottom, so flushing problems can explain it. None of the outliers are severe and all are likely explained by incomplete flushing of Niskin bottles. No flags are justified.

For full details for the COMPARE run see file 2014-27-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter, with no obvious grouping. The pumps were turned off for cast #3, so that was excluded from the comparison. The choice of outliers determines the fits, many of which do not look like those we usually see.

When just the 6 deep casts at the end are included and bottles within 3m of the bottom and some surface outliers are excluded, and the trendline is forced to have a zero offset, the fits do not look convincing. But if the offset is forced to be -0.0744 (the offset found for 2014-06), the slope is 1.0281 which is close to the slope of 1.0300 from 2014-06. So the 2014-06 correction looks appropriate for the deep casts.

For the shallow casts the results are not as good. Removing outliers in the same way as was done for the deep casts, produces a similar fit (slope 1.0265), but with so much scatter that it is not convincing. But these data are from a very narrow DO range, especially after removing the near-bottom bottles.

The best we can do is to apply the correction from 2014-06 which included a wide range of DO using the same equipment:

CTD DO Corrected = CTD DO * 1.03 + 0.0744

There are no obvious outliers. Event #3 stands out in a plot versus file pair number. This looks a little like the problem noted in 2014-06 when 2 casts had very different fits from all other casts in plots of differences between Titrated DO and CTD DO versus CTD DO. Could those 2 odd casts have been due to pumps not being on, even though the pump status channel indicates they were on? A plot was made of Event #3 and the slope and offset were only a little different from the other casts, not at all like the cases seen in 2014-06. So pumps were likely not the problem for the earlier cruise.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on 2 previous cruises and an advance of +2.5s brought the vertical offset between the downcast and upcast DO traces into the best agreement with the temperature traces.

ALIGNCTD was run using a setting of +2.5s to the dissolved oxygen channel.

Examination of the deeper casts after this step shows satisfactory alignment.

##### CELLTM

Tests are not helpful for these data since there are so many shallow casts with high variability and the deeper casts have many stops for bottles. Whether any setting improved the data was not clear, but in other uses of this type of equipment, there is always improvement. So the results of cruise 2014-06 were applied.

CELLTM was run using (α = 0.02, β=9) and (α = 0.0245, β=9.5) for the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The differences between sensor pairs were studied only on cast #73 since it was the only cast deeper than 300m. The temperature differences were extraordinarily noisy, so the only level at which the readings were recorded were close to the bottom. The differences from 2014-06 are included for comparison.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-06-0043 | 300 | +0.0002 | +0.00042 | +0.0042 | Very steady |
| 2014-06-0053 | 300 | ~0 VN | +0.00041 | +0.0043 | Very steady |
| 2014-06-0053 | 350 | ~0 VN | +0.00044 | +0.0044 | Very steady |
| 2014-06-0073 | 300 | +0.0013 VN | +0.0005 | +0.0044 | Very steady |
| 2014-27-0073 | 335 | +0.0002 | +0.00045 | +0.0045 | Very steady |

So the cast #73 differences are similar to those found for 2014-06.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. The only discrepancy was the station name for cast #3 which the log indicates should have been changed from 22 to 23. That was fixed in the CNV, IOS and CLN files.

A header check was run. There are a few values that are odd, such as a negative fluorescence value, but these are likely due to near-surface spikes. No other problems were noted.

The cruise track was plotted with event #s and station names and both were added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.3db which is reasonable for the Vector.

The altimeter and water depth readings from the headers of the CLN and SAM files were exported to a spreadsheet. The depths sometimes differed from the one entered in the log especially at shallow casts, but in most cases both were lower than the maximum depth plus the altimeter readings. This could be due to drift during the cast, though it always seems to be drift to deeper water, or the sounder may have been reading a little low. The depths were not adjusted as those do vary through a cast and the altimetry varies somewhat at the bottom.

Plots were made of altimetry for all casts and the headers mostly look appropriate despite some severe noise near the bottom. Exceptions are entries for:

- Event #2 looks - somewhat too high in the CLN file, but the entry in the SAM file looks better, so that value was entered in the CLN file.

- Event #21 was changed in both the CLN and SAM files from 3.5 to 2.5m.

- Events #75 and 76 had no entry for the SAM files, but there are entries for the CLN files at the same depth, so those entries were used for the SAM files.

The SAM files were bin-averaged and merged with the MRGCLN1s files again and the MRG files were then put through CLEAN again to capture these changes.

##### Shift

Fluorescence

SHIFT was run on the WetLabs ECO fluorescence channel in all casts using the usual advance of +12 records. Examination of plots after this step were hard to interpret due to noisy fluorescence and many stops for bottles, but there appears to be slight improvement in matching the fluorescence offset to the temperature offset.

pH

As was found for fluorescence it is difficult to test these data to choose the best setting for aligning pH. They clearly need aligning and using the setting found best during 2014-06 definitely improves the alignment with temperature. SHIFT was run on the pH:SBE channel using a setting of +70 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but the alignment seems ok.

Conductivity

Tests were run on 2 casts using a variety of shifts and the best results were found with the settings used for 2014-06. SHIFT was run using -0.7 records for the primary conductivity and -0.6 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings. But plots show that a lot of data were removed from cast #21 because the descent rate was very low (and very steady) throughout the cast. DELETE was rerun on that cast with the low descent rate feature turned off.

##### Other Comparisons

Previous experience with these sensors –

- Salinity:

This was the 3rd cruise using these temperature and conductivity sensors since they were calibrated at the factory. Primary salinity was found to be lower than bottles by ~0.004 and the secondary salinity was very close to bottles.

- Dissolved Oxygen

The sensor was calibrated at the factory in January 2013 but not used until 2014. There were 2 uses before this cruise. There was limited sampling during 2014-03, but good sampling during 2014-06.

- Pressure

The sensor was recalibrated in January 2014 and it was used for 2 other cruises before this one. The factory offset was used for those cruises and the results looked reasonable.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature and salinity data fell within the climatology.

Repeat Casts – There were many repeat casts at station 22, but they were very shallow in a region of high variability, so they are not suitable for testing repeatability of data.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### CHANNEL CHOICE STUDY

Since the secondary channels were selected for the previous 2 cruises using this equipment and the secondary salinity appears to be more reliable than the secondary, the secondary channels were selected for editing. The secondary T/S curves appear to be a little more stable, though there is not a great deal of difference.

##### DETAILED EDITING

The secondary temperature and salinity channels were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts.

Events #10 and 12 required no editing.

All other files were edited lightly.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure does not appear to require recalibration.

From 2014-06 which preceded this cruise, the primary salinity was lower than bottles by an average of 0.003 and the primary high by 0.0011. When only the 4 deepest bottles (below 275m) were used the primary was low by 0.0035 and the secondary low by 0.0002. From this cruise there are only 3 bottles that seem reliable, and the average of those shows the primary to be low by 0.004 and the secondary high by 0.00001. The difference between these two estimates is close to that observed in section 9. The secondary appears to be very close to bottles and will not be recalibrated.

For dissolved oxygen the results of 2014-06 will be applied to these data since the earlier cruise had more bottles and 2014-27 has too much scatter to determine a convincing correction.

CALIBRATE was run using file 2014-27-recal1.ccf to apply the following correction to DO in the SAM and MRGCLN2 files:

CTD DO Corrected = CTD DO * 1.03 + 0.0744

COMPARE was rerun for dissolved oxygen using the recalibrated values, and plots confirm that the recalibrations improve the comparison. The CTD DO look a little high near the surface which is probably due to incomplete flushing of Niskin bottles so they contain water of lower DO than ambient conditions. Towards the bottom the CTD DO looks low, but the low values come primarily from bottles within 3m of the bottom where there is also likely to be poor flushing of Niskin bottles, but with the opposite effect of bottles containing higher DO water.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles – there were no downcast CTD data available sufficiently close to some bottles. When the bottles within 3m of the bottom and Event #3 (pumps off) were excluded, the differences are small at depth but increase with decreasing depth. This is likely due to the effects of incomplete flushing of Niskin bottles being greater in the higher DO gradient zones near the surface. The bottom bottles that were excluded mostly show the CTD reading a little low, which is likely also due to incomplete flushing, but the effects are smaller in the lower DO gradient zone near the bottom. The results do not suggest a need for further recalibration of DO.

See 2014-27-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some unstable features, but those are expected in this region.

Profiles mostly look fine though but there is no PAR:Reference signal for casts #1-2, so it will be removed. The signal is odd for casts #55, 65 and 67 but there is some signal, though it doesn’t vary much, if at all.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For casts #1-2 channel PAR:Reference was also removed because there was no signal.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH, PAR and Reference PAR data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8 mL/L from 0 to 10db

±0.4 mL/L from 10 to 50db

±0.1 mL/L from 50 to 100db

±0.06 mL/L below 100db

The PAR:Reference channel was removed from casts #1-2 as there was no

signal; there are very low signals from other casts, which are suspicious,

but they were left in place.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2014-27_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. In Baynes Sound the surface saturation values were mostly between 90 and 110% but were a little higher towards the north. Values were also between 90% and 110% in the open waters to the north. In the open waters towards the south values were much higher, reaching as high as 180% at station 22. This may be due to geography and/or time since the values >130% were all late in the cruise.

Checks were made against bottle DO values to see if these high surface saturation values are realistic. The downcast CTD DO values are higher than the CTD or bottle DO from the upcasts in some cases, but there is also a big difference between downcast and upcast profiles in those cases with salinity and temperature also varying a lot between the down and upcast. Plots show that the surface bottle stops were often short, as little as 10s so that shed wakes would be bringing lower DO water past the CTD and bottles. Most stops were long enough, just not the ones at the surface where the gradients are largest so that the shed wake problem is most severe.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For casts #1-3 channel PAR:Reference was also removed because there was no signal.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check was run on all files and no errors were found.

A header check was run on the final files.

No further errors were found.

The track plot looks ok.

Plots of each file were examined and no further problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

3. Station 22 in file - should be station 23. Pumps off throughout cast.

13. Bottles fired out of sequence. DIC cast.

21. “Dead Slow Cast” – GYO SONDE attached.

37. DIC cast

NOTE: RECALIBRATION 21 July 2014

When these data were originally processed, it was believed that the secondary salinity did not need recalibration based on comparisons of CTD and bottle salinity during cruises 2014-06. There were some doubts about that comparison because of incomplete flushing of Niskin bottles, but it was the best available. Later, comparisons from 2014-21 and 2014-18 showed that the primary salinity was close to bottles, while the secondary was high. So, the primary CTD salinity is presumed to be reliable. The secondary salinity data from this cruise were recalibrated by subtracting 0.004, which was roughly the difference between the two channels.

The two conductivity sensors will be returned to the factory for a post-cruise calibration. When the results are available, further recalibration will be applied, if needed.

G. Gatien

Note: 19 November 2014

Post-cruise factory calibrations of T and S sensors show that the secondary salinity was high by about 0.005 by mid-August. Based on observations during a series of cruises, it appears that the calibration applied to these data in July was appropriate, though there remains some doubt about when drift occurred.

G. Gatien

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 31Dec2013 | Factory |  |  |
| Conductivity | 1766 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 4700 | 4Jan2014 | Factory |  |  |
| Secondary Cond. | 3321 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 21Jan2014 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0692 | 14Jan2014 |  |  |  |
| WetLabs ECO Fluor. | 2216 | n/a |  |  |  |
| Pressure Sensor | 77511 | 30Dec2013 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |