## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2014-28

Agency: WVL, Marine Environment and Habitat Science, Vancouver, B.C.

Location: Bayne Sound

Project: PARR Shellfish Carrying Capacity

Party Chief: Sutherland T.

Platform: Vector

Date: June 19, 2014 – June 24, 2014

Processed by: Germaine Gatien

Date of Processing: 25 February 2015 – 16 March 2015

Number of original HEX files: 37

Number of CTD files: 37	Number of bottle casts:	18

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (Arctic #1189) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), an SBE 43 DO sensor (#1483), a SeaPoint Fluorometer (#3640), a Biospherical QSP-400 PAR sensor (#4656), a surface PAR (#16504), a pH sensor (#0691) and an altimeter (#43281).

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log was missing a personnel list. The rosette log sheets and analysis logs were in good order. No cruise report was available at the time of processing. There were no photos of the setup.

An offset included in the pressure calibration was originally presumed to be based on some test done before the cruise, but it turned out to be a mistake. So a correction was added later to remove that offset.

The rosette sheet indicates that 1 salinity sample was taken but no analysis of that sample has been found. While a single sample is not of great use in calibration, losing samples is a concern. The same CTD (with different external sensors) was used during a later cruise, 2014-56, which did include salinity sampling and there was a post-cruise calibration after 2014-56. The secondary temperature and salinity channels were archived for most casts, but for 2 casts the primary channels were selected. So salinity channels were recalibrated principally to bring the two channels close.

The comparison of CTD dissolved oxygen and the titrated bottle DO was noisy, likely due to flushing problems near the surface and near the bottom of casts. However, the differences were not very large. The range of DO values was relatively small and response appeared to be relatively quick.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 10db

±0.25 mL/L from 10 to 100db

±0.15 mL/L from 100 to 200db

±0.05 mL/L below 200db

The pH values at 4 near-shore casts in Baynes Sound were very much lower than for all other casts. The pH sensor has not been calibrated since 2010.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Many files had names with an extra 0 in them, so program Bulk Rename Utility was used to change them to standard format.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets and final spreadsheets for dissolved oxygen and nutrients.

The cruise summary sheet was completed.

The history of the DO sensor was obtained; no history was available for temperature and conductivity sensors, as they had not been used before this cruise.

The calibration constants were checked for all instruments. The DO calibration data was out of date so the information was replaced with values from a more recent calibration. The date for the pH sensor calibration was missing, so that was added. The corrected file was saved as 2014-28-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2014-28-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. The temperature and conductivity channels are further apart on downcasts than usual but this is likely an issue with alignment; both look quite smooth. The upcast traces are sometimes closer but often noisy. The altimetry sometimes has spikes at the bottom, though the trend above that may enable an estimate if the header reading algorithm fails.

The fluorescence, transmissivity, DO, PAR and SPAR look normal. pH values are low for some casts.

The descent rates are quite steady and kept high at mid-depths.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2014-28-ctd.xmlcon.

The files were converted to IOS format. They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. There were a few bad values in the primary salinity in file #56 (170m) and the secondary salinity in file #6 (25m); 3 points were removed from each of these casts. The edited file was copied to *.BOT.

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2014-28-bot-hdr.txt; it may need further editing to reflect problems found during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-28oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2014-28oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

The rosette log shows that 1 salinity sample was taken but it was apparently lost.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-28nuts.xls. This includes a precision study. There were some bucket samples that did not come from Niskin bottles, so will not be included in the rosette files. The file was simplified, reordered on sample numbers and saved as 2014-28-nuts.csv and the bucket samples were removed. The file was converted to individual NUT files.

The OXY and NUT files were merged with CST files in 2 steps.

After the 2nd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files, to produce MRG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions and it was discovered that bottle data were missing from event #34. The rosette sheet shows the event number for station BS17 as #33 while the Daily Science Log Book indicates it was #34. There was an event #33 but there were no bottles fired. So the OXY and NUT file names were changed to #34; the merge process was repeated. The exported spreadsheet data then looked fine.

The MRG files were put through CLEAN to produce MRGCLN2 files.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is so much scatter that it looks impossible to find a convincing trendline.

If all data are included, the average of the (CTD DO – Bottle DO) / CTD DO is -0.034; in other words the CTD is low by ~3.4% which is fairly typical for this type of instrument given one other use since factory recalibration. During 2014-23 it was considered low by ~4.4%.

To the eye, the differences look quite flat as though there might be a constant offset, so a fit was forced through the average difference, -0.145, but the R2 value was extremely low at 0.0019.

If a fit is forced through the zero then it looks more like the type of fit we expect and the R2 value is a little better at 0.129. Removing a few outliers based on expected performance, the R2 value rises slightly and the slope is about the same. The fit is far from convincing. Sampling was mostly shallow, so a subset of 4 casts that went below 60m was studied. When data are removed that are obvious outliers (assuming a normal fit) and the fit is forced through the origin, the slope is 3.9%.

The fact that the fits don’t look normal is likely due to inefficiency in the flushing of the Niskin bottles which becomes important in high gradients. Previous cruises in this area and in other well-protected waters have shown the same character, presumably because the Niskin does not bounce up and down before closing so it doesn’t flush well. All samples that read higher than the CTD DO came from the top 25m and most are from the top 10m. The water in the bottle sample likely comes from lower in the profile than the water measured by the CTD and hence is reading lower than the CTD. So ignoring shallow samples that are reading higher than the CTD is justified. Unfortunately, this does reduce the range of DO values in the fit. Another problem is that there were no DO values <2.8mL/L so we don’t have the ability to make a good estimate of the offset. With that in mind a fit through the origin is appropriate.

Three of the notable outliers with CTD DO much lower than bottles came from the bottom of casts where poor flushing has the opposite effect to that seen at the surface because the Niskin bottle contains water from above it, not below, so DO values are higher than expected in the samples. So much of the scatter is easily accounted for by poor flushing.

When data are included only if differences were between -0.5 and -0.03mL/L and standard deviation in the CTD DO <0.15, the fit was:

CTD DO Corrected = CTD DO * 1.0348

None of the outliers are severe enough to justify downgrading the quality flags, especially given the fact that the cause of large differences is more likely poor flushing, not analysis or sampling.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

One way to correct the DO signal for errors due to transit time and slow response time is to compare upcast and downcast temperature and DO traces; this is difficult when there are so many stops for bottles and so many very shallow casts. The offset is due to both the CTD carrying water with it on the upcast and the response time of the DO sensor. Since the temperature is only affected by the former, the difference is assumed to be due to transit time and slow response of the DO sensor. However, for these data there are very few casts deeper than 50m and only 2 of them had no stops for bottles. Those 2 suggest an alignment of about 2.2s, but the best choice varies greatly from feature to feature. (NOTE: At a later stage this setting was found to be too high, so SHIFT was run to apply a -12 records adjustment, equivalent to -0.5s, so the net setting was 1.7s.)

Another method is to align downcast DO and temperature traces by matching features, but while there are many notable reversals in DO there are few striking features in temperature. A setting of +2.2s does appear to improve the alignment.

ALIGNCTD was run using a setting of +2.2s to the dissolved oxygen channel.

##### CELLTM

The usual test for choosing the parameters for CELLTM involves comparing the results on a TS curve to see which produces upcast and downcast traces that are closest. This is difficult for these data because the casts are shallow and there are many stops for most upcasts and the upcast temperature tends to be rather noisy. All the settings tested improved the data, with the best results appearing to be with (α = 0.02, β=9) and (α = 0.03, β=9.5) for the primary and secondary conductivity, respectively.

CELLTM was run using (α = 0.02, β=9) and (α = 0.03, β=9) for the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

Differences between sensor pairs were studied only on the deepest cast, #98, since it was the only one deeper than 300m. The temperature differences were extraordinarily noisy, so the only level at which the readings were recorded were close to the bottom.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-28-0098 | 310 | -0.0003 VN | +0.00015 VN | +0.0016 VN | Very steady |

It is difficult to tell much from such shallow data but there is no indication of a problem.

However, examination of plots show large differences in areas of high gradients; it looks like the primary temperature and conductivity are slower to respond to change and they look noisier in those regions and fail to resolve steps seen by the secondary. On the other hand there are some bad secondary salinity data during events #29, 50 and 54 in the top 10m, 4m and 8m, respectively.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book.

- The first cast had different latitudes in the log book and the file header; it is assumed that the file header is correct since there appears to be insufficient time to travel from event #1 to #2 if the log is correct.

- Event #18 is listed in the log but there was no HEX file; there is no note of explanation but the whole cast took only 2 minutes so was presumably either a test or abaorted.

- Event #38 has different latitudes in the log and file header, but the file header agrees with the log entry for the net cast at the same station, so the file header is assumed to be correct.

- Event #92 had the wrong station name – it should be BS38 – that is clear from several entries in the log. The entry was fixed in the full file and in the hydro files for that cast.

A header check was run. The only problem noted was a minimum pressure of about -0.9db during event #2. An examination of the file shows that this data is from the upcast. The primary conductivity shows “in water” values until the pressure is ~-0.7db and the secondary does so until -0.8db.

Surface check was run and shows an average surface pressure for the cruise was 0.9db which is low for the Vector.

The pressure offset is -0.69db. This does not match the information from the Arctic group dated 27 May 2014. It was assumed that the change was deliberate due to drift as measured by tests at the beginning of the cruise. However, it appears it was entered in error. Adding 0.7db to values looks appropriate.

The cruise track was plotted with event #s and station names and both were added to the end of this report.

The altimeter and water depth readings from the headers of the CLN and SAM files were exported to a spreadsheet. The depths were checked against the log entries and an error was found for cast #40; that was fixed in the IOS and CLN files; it was not a rosette cast. There was another entry with a significant difference between the header and the log, but the header looks more reasonable.

Plots were made of altimetry for all casts and the headers mostly look appropriate despite some severe noise near the bottom. The reading for the bottle file for event #56 had a reading that was too high, likely due to a spike. It was changed from 4.2 to 2.5m in the bottle files. The full profile files had a better reading.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step were hard to interpret due to noisy fluorescence and many stops for bottles, but there appears to be slight improvement in matching the fluorescence vertical offset to the temperature offset.

pH

The pH sensor clearly was not pumped. While the upcast and downcast traces are further apart than for the temperature, the downcast data look well aligned with major features in the temperature.

No shift was applied.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier, but the check of some casts suggests the correction went too far. SHIFT was used with a setting of -12 records (~0.5s) so the net advance was 1.7s.

Conductivity

Tests were run on 4 casts using a variety of shifts. There were few unstable features so the results are not very different, but overall a setting of +0.2 records improved the primary salinity slightly and -0.5 records improved the secondary salinity.

SHIFT was run using +0.2 records for the primary conductivity and -0.5 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Salinity: The temperature and conductivity sensors were new in April or early May. This was probably the first use.

Dissolved Oxygen: The DO sensor was calibrated at the factory in January 2014 and was used shortly before this cruise during 2014-23 when a linear correction was applied with slope 1.04 and offset +0.0037.

Pressure: The sensor was new in December.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. This climatology is not really appropriate to data collected so close to the coast. All salinity data fell within the climatology. Temperature values were a little above the climatology maximum for events #68, 71, 74 & 77 between 25m and 35m. This does not look like a calibration error as there are other cases where temperature is very close to the minimum (events #89 & 98 around 140-150m.)

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were post-cruise calibrations in mid-December 2014 for the temperature and conductivity channels. The factory reported drift since the previous calibration of -0.0009C°/year and -0.0010C°/year for the primary and secondary temperature channels, and -0.0011psu/month and -0.0007psu/month for the primary and secondary conductivity channels.

##### CHANNEL CHOICE STUDY

The choice of channels to archive is difficult because the primary data looks poor in areas of high gradients, but the secondary is clearly bad near the surface for a few casts. The secondary temperature and conductivity will be selected for editing except for casts #29 and 54. For cast #51 the secondary salinity channel was poor right at the surface but there were problems in the primary as well, so removing some surface data from the secondary channels look best.

##### DETAILED EDITING

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts.

All casts required light editing.

##### Initial Recalibration

The pressure will be adjusted by adding 0.7db, based on information in section 11.

The dissolved oxygen data will be adjusted based on the information in sections 5 and 14. The bottle comparisons for 2014-23 and 2014-28 suggest corrections should be:

CTD DO Corrected = CTD DO * 1.04 + 0.0037   	 (2014-23) Mid June 2014

CTD DO Corrected = CTD DO * 1.0348     		 (2014-28) Late June 2014

The correction based on the comparison for this cruise is slightly lower but the range of DO values was smaller which may have allowed better response of the sensor, so it will be used.

For salinity we have 4 pieces of information:

- The sensors were calibrated at the factory shortly before this cruise and this is the first use on record.

- The differences between the salinity channels are ~0.0015 in areas of low salinity gradients.

- The same T and C sensors were used during 2014-56, a long cruise. There was salinity sampling that indicated that the primary salinity was reading low by ~0.014 and the secondary low by ~0.011. The standard deviations in the fits were large and the casts were mainly shallow so that there were likely Niskin bottle flushing problems leading to higher bottle values, making the CTD salinity channels look lower than they really were. There was also delay in analysis that could possibly have led to some evaporation of samples. Both factors would affect both channels in the same way, so the difference between channels is likely reliable.

- A post-cruise calibration indicates that by December 2014, after 2014-56, the drift in conductivity and temperature sensors would lead to primary salinity reading low by ~0.008 and secondary reading low by ~0.005.

The differences found between primary and secondary salinity channels in the 2014-56 bottle comparison are consistent with the differences based on the drifts found during the post-cruise calibrations. In both cases the secondary was higher than the primary by about 0.003. The difference between salinity channels was lower for this cruise, at ~0.0015. The secondary had the smallest drift according to the factory report. Assuming a little drift by the time of this cruise and a difference between the two of 0.0015, then a reasonable guess is that the primary was reading low by ~0.0025 and the secondary was reading low by ~0.001. If we assume there was 3 times as much drift by the time of the factory visit, then we would have the primary low by about ~0.008 and the secondary by ~0.003, as was reported. The important thing to achieve in the recalibration of these data is to have the primary and secondary reading close together, since we chose different T/C pairs for different casts. So recalibration by adding 0.0025 to the primary and 0.0010 to the secondary will achieve that and make a reasonable estimate for each. While this is a rough recalibration it is likely better than ±0.002.

CALIBRATE was run using file 2014-28-recal1.ccf to add 0.7db to the pressure, +0.0025 to the primary salinity, +0.0010 to the secondary salinity and to apply the following corrections to the SAM and MRGCLN2 files:

CTD DO Corrected = CTD DO * 1.0348

COMPARE was rerun for dissolved oxygen using the recalibrated values, and plots confirm that the recalibrations improve the comparison. The CTD DO look a little high near the surface which is probably due to incomplete flushing of Niskin bottles so they contain water of lower DO than ambient conditions.

The 2 salinity channels are closer after the recalibration.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the upcast bottles. The results do not suggest a need for further recalibration of DO. See 2014-28-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined and look good.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except #29 and 54 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For casts #29 and 54 the following channels were removed:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH, PAR and Reference PAR data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 10db

±0.25 mL/L from 10 to 100db

±0.15 mL/L from 100 to 200db

±0.05 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2014-28_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made. The PAR:Reference channel had some very low and almost constant values. A closer look at the files before Delete and Bin Average were run shows some variation at the beginning of the files. The surface PAR is subject to errors due to shading or ship lights, so this is likely not a sensor problem. It is nominal data so was left as is. pH values were very low at 4 stations in Baynes Sound but this may well be right.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. In Baynes Sound the surface saturation values were mostly between 95% and 110%. At the far ends of the sound some lower surface saturations (2 casts with 65% and 1 cast at 85%) were found. Outside of Baynes Sound values occasionally got as high as 120%.

Where possible checks were made against bottle DO values to see if the highest and lowest surface saturation values were realistic and the matches with SBE DO were good in the bottle files.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check was run on all files and no errors were found.

A header check was run on the final files.

No further errors were found.

The track plot looks ok.

Plots of each file were examined and no further problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

18. No HEX file. Entered in log book as normal excepet that the whole cast lasted 2 minutes and no bottom depth is shown. Assume there was no cast.

# CRUISE SUMMARY

# CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1189 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5830 | 1May2014 | Factory | 14Dec14 | Factory |
| Conductivity | 5831 | 10April2014 | Factory | 17Dec14 | Factory |
| Secondary Temp. | 4700 | 29April2014 | Factory | 18Dec14 | Factory |
| Secondary Cond. | 3321 | 9April2014 | Factory | 18Dec14 | Factory |
| Transmissometer | 983DR | 12June2014 | IOS |  |  |
| SBE 43 DO sensor | 1483 | 16Feb2013 | Factory |  |  |
| PAR | 4656 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0691 | 24Dec2010 | Factory |  |  |
| Seapoint Fluor. | 3640 | Jan2014 |  |  |  |
| Pressure Sensor | 1189 | 14May2014 | Factory |  |  |
| Altimeter | 43281 | n/a |  |  |  |