## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 25 March 2025 | ### Changed channel name Silicate:Flag to Flag:Silicate G.G. |
| ### 16 March 2021 | ### Added DIC and ALK data to events 2, 5, 7, 14, 31, 34, 37. S.H. |
| 4July2018 | Added DIC and ALK data to 2014-50-0014.CHE |
| 12May2015 | Sample #208 – values replaced with pad values – premature closing of Niskin bottle. |

## PROCESSING NOTES

Cruise: 2014-50

Agency: Ocean Sciences Division

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: October 26, 2014 – October 30, 2014

Processed by: Germaine Gatien

Date of Processing: 13 November 2014 – 9 February 2015

Number of original HEX files: 31 (2 shallow casts for videographer – not to be processed)

Number of CTD files: 29	Number of bottle casts:	22

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1438), a SeaPoint Fluorometer (#3642), a Biospherical QSP-400 PAR sensor (#4615), a surface PAR (#16504), a pH sensor (#0691) and an altimeter (#43281).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave V7.

The CTD deck unit was an SBF model 11+, serial number 0619.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order.

During the downcast of event #31 there is a small section of water near 210db with lower transmissivity than above and below it. On the way up the section of odd data is much more extensive (~120db-240db) with transmissivity and dissolved oxygen much lower than during the downcast and fluorescence higher. (up to 0.4ug/L). Above 120db transmissivity and fluorescence are close to the downcast values, but the dissolved oxygen stayed low. When the equipment was brought aboard biological fouling of the DO sensor pump inlet was discovered. The affected DO data were replaced with pad values in the bottle file. The other variables look reasonable during bottle stops, so were left unchanged.

The Oxygen:Dissolved:SBE voltage data are of lower quality than usual with frequent shifts to higher values followed by shifts to lower values. When converted to oxygen concentration this produces two-sided spikes (~±0.1mL/L). No such shifts were seen in other channels. Filtering and de-spiking routines were ineffective at removing these excursions. While the data are noisy even during bottle stops, the sensor values compare reasonably well with titrated samples. This looks like a flow rate problem. There are hints of a similar problem in the secondary temperature channel, but the effects on salinity are less notable than those on dissolved oxygen.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 30db

±0.3 mL/L from 30db to 100db

±0.12 mL/L from 1000db to 200db

±0.04 mL/L below 200db

Both salinity channels were found to be close to bottle samples.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen, extracted chlorophyll and salinity.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. Many of the sensors had been used for only a few casts since their last factory check.

The XMLCON files did not change through the cruise.

The calibration constants were checked for all instruments and the only corrections needed were to add the date of calibration for the PAR sensor and correct spelling of the date for the transmissivity. The corrected file was saved as 2014-50-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2014-50-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. The temperature and conductivity channels are close on downcasts though there were some spikes. As usual the upcasts are noisy and the differences are larger.

The dissolved oxygen voltage has some odd shifts. These are not single-point spikes, but a shift to higher values followed by one to lower values, over about 30 scans. The shifts are on the order of ±0.0005V When DERIVE is run to get DO in concentration units, two-sided spikes are produced of size ±0.1mL/L. Running ALIGN first has no effect on the shape of those spikes, just the depth at which they appear. A test run of WILDEDIT applying it to the DO voltage was ineffective because there are many points in the excursions. A variety of filters were tested and they reduce the excursions somewhat, but smooth the data too much. There was as major problem with the DO (and other sensors) during the upcast of event #31. Titrated DO samples show the downcast data are reasonable, but not the upcast. The DO sensor pump intake was cleaned after event #31 because of obvious biological fouling. The problem looks like it is due to interrupted flow, but even after cleaning the sensor pump intake some spikes appear. The secondary temperature is noisy and may have some very small shifts similar to DO, but that is not clear.

The altimetry looks useful and fluorescence, transmissivity, pH, PAR and SPAR look ok.

The descent rate of the CTD is highly variable, generally high on average, but in some cases very steady and in others very noisy. There was some stormy weather during the cruise.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2014-50-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files to check for outliers. CTDEDIT was used to clean a few salinity points in files #8 and 36.

The ED1 files were copied to BOT files.

A preliminary header check was done and it showed that there were no station names in the files. This will be corrected later. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

A few bottles were removed from the list. For cast #7 Niskin #1 was accidentally fired during the downcast, and for cast ##8 two bottles were fired at the surface but only 1 was sampled, so the 1st (Niskin #12) was removed from the list since there were more records for the 2nd (Niskin #13). Two bottles were fired during each of events #23 and #24; they were not sampled so were removed from the addsamp file.

Casts #23 and 24 will not be processed further.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine.

Merge CSV to Headers was used to add water depths and station names to the SAM files. (SAMMRH)

Those files were then bin-averaged and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2014-50-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-50chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2014-50chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-50oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2014-50oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2014-50SAL.xls. The analysis was done within 14-19 days of collection. The files were simplified and saved as 2014-50sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-50nuts.xls. This includes a precision study.

Then the file was simplified, reordered on sample numbers and saved as 2014-50-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. The only oddity noticed was that cast #8 had an extra bottle fired beyond those noted on the rosette sheet. This was noted when the ADDSAMP file was prepared and the decision was made to treat the samples as having come from Niskin #13 rather than #12 so that there was more CTD data available to go with them. They were fired very close together. It is probably better to change this to #12 for accuracy, but use the CTD data from #13. This step was a little awkward, so if there are some inconsistencies in the final file, this will likely be the source of the problem. A note was added to the MRG file to explain what was done. The bottle data come from Niskin #12 and the CTD data come from around the time of firing of Niskin #12, just happening to get recorded in Niskin #13.

##### Compare

Salinity

Compare was run with pressure as reference channel.

When bottles were excluded from the comparison if the standard deviation in the CTD salinity was >0.001, the fits were quite flat, with just one other outlier for each of the salinity channels:

- Event #1 - The secondary salinity was found to be low by 0.024. The standard deviation in the CTD salinity was very slightly <0.001. The primary salinity was also an outlier but had standard deviation >0.001. This outlier could be a sign of poor flushing of the Niskin bottle. The only other shallow samples were from cast #34 which had CTD data much higher than the bottle but the standard deviation was >0.001 and cast #10 where a bottle at 99db had salinity that was slightly high. So poor flushing does not appear to be a significant problem.

- Event #35 – The primary salinity was found to be low by 0.011. While the standard deviation was <0.001, there was a lot of noise in the CTD data shortly before firing time. The secondary salinity was also low but was removed because the standard deviation was >0.001.

The noise level in the traces suggests that the outliers are due to noisy CTD salinity rather than problems in sampling or analysis, so no quality flags are justified.

In other recent inland cruises there have been problems that are thought to be due to poor flushing of Niskin bottles. For this cruise there were only 3 salinity samples from above 100db where higher gradients make incomplete flushing a significant problem, and the descent rate was often noisy which probably helps flushing. One shallow bottle did look like it might not have flushed well, but the CTD data were fairly noisy during the stop so it is not clear. The other 2 shallow bottles had CTD salinity higher than the bottles. Flushing may be an issue for interpretation of shallow samples, but without more shallow bottles, we have no evidence for this.

The salinity was analyzed about 2 weeks after collection, so there is no expectation of a problem with evaporation and there is no evidence of any such problem.

When the outliers are excluded there remain 12 bottles and the average differences show the primary salinity to be high by an average of 0.0003 and the secondary salinity to be low by 0.0004 with a standard deviation of 0.0008 and 0.0010, respectively. The fits were quite flat with pressure, with the secondary salinity having the flatter fit.

The difference between the primary and secondary salinity is ~0.0007 which is reasonably close to the differences described in section 9.

No recalibration of salinity appears necessary.

For full details for the COMPARE run see file 2014-50-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There are some clear outliers, but most of the data falls in a fairly tight group.

All major outliers come from either events #1 or event #31. All bottles from 200m upwards from event #31 are clear outliers. It was noted at sea that there was biological fouling of the CTD and rosette. And event #1 is from Saanich Inlet where the DO gradients are high and the comparison is usually poor.

So for the major outliers, there is no reason to suspect the analysis or sampling.

Other less severe outliers were examined, ones that differed from points on the trendline by from 0.15mL/L to 0.3mL/L. The standard deviations in the CTD dissolved oxygen data were not much higher for those bottles than for others, but they are quite high for this cruise in general. An examination of profiles shows that there were many spikes, up to ±0.1mL/L. So the CTD DO does not look as reliable as usual. The sensor was on the secondary system and the secondary temperature was also found to be very noisy.

The following 5 outliers were examined more closely, but in only one case does it look like a change of quality flag is justified:

Event #13 – Sample #123 – No problem noted by analyst. There was a lot of variability at the bottom as the CTD moved downwards through the stop, so a mismatch with the Niskin contents may be partly responsible for the outlier.

Event #14 – Sample #148 – No problem noted by analyst.

Event #33 – Sample #222 – Originally flagged 3 by analyst. Bottle is from the bottom – it is only high by 0.03mL/L, so not severe, but the analyst felt it raised the likelihood that the titration overshot - the flag was downgraded to 4.

Event #34 – Sample #238 – No problem noted by analyst.

Event #36 – Sample #271 – No problem noted by analyst.

When outliers are removed based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0281

or

CTD DO Corrected = CTD DO * 1.0191 + 0.0366

depending on whether an offset is allowed or not. Since there was only 1 bottle with DO<1.8mL/L and it comes from an area with a high vertical DO gradient, allowing an offset is not justified.

Plots of Titrated DO and CTD DO against CTD salinity were examined. The only outliers of note were the bottles from Event #31. All but the first 3 bottles are major outliers as already noted. And the plot makes it quite clear that the SBE DO data are the outliers, not the bottles.

The Oxygen:Dissolved:SBE data from above 250db were replaced with pad values for event #31.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

All CHL values were <1.5ug/L with the exception of one value of ~9.1ug/L in Saanich Inlet. As usual, fluorescence reads higher than CHL for low CHL but starts reading lower than the CHL by the time CHL=1ug/L. The drop in the ratio of FL/CHL is a little steeper than usual, falling to about 0.4 to 0.6 for CHL>1. For full details of the comparison see file 2014-50-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

An initial estimate was made that the DO leads the temperature traces by about 2.5s.  ALIGNCTD was run using a setting of +2.5s to advance the DO signal. The results look good. One cast was tested using 2s and 3s to see if either looked better. A setting of +2s is ok, but +2.5s is better and +3s is worse than both the lower settings. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. One cast with a quiet section was tested using a variety of settings and there was little difference among them, with the default setting of (α = 0.0245, β=9.5) looking slightly better than the others for both conductivity channels..

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. While none are very deep, it does allow us to check for sudden changes.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-50-0003 | 250 | +0.001 XN | -0.0001 XN | -0.0010 XN | High, Steady |
| 2014-50-0015 | 235 | +0.0005 | -0.00006 | -0.0012 | High, F.Noisy |
| 2014-50-0018 | 250 | +0.0006 N | ~-0.00007 N | -0.0011 N | High, F.Noisy |
| 2014-50-0019 | 250 | -0.0002 | -0.00010 | -0.0008 | Mod, V.Noisy |
| “ | 330 | ~0 N | -0.00008N | -0.0006 N | “ |
| 2014-50-0025 | 250 | -0.0006 | -0.00009N | -0.0005 VN | High, F.Noisy |
| 2014-50-0028 | 250 | -0.0007 | -0.00003 | +0.0004 | Mod, V.Steady |
| 2014-50-0031 | 250 | -0.0017 | -0.00006 | +0.0010 | High, F.Steady |
| “ | 330 | -0.0018 | -0.00003 | +0.0014 | “ |
| 2014-50-0037 | 250 | -0.001 XN | -0.0001 N | -0.0004 N | Mod, Steady |
| “ | 325 | ~0 VN | -0.00004 | -0.0004 | “ |

The conductivity and salinity differences are small but appear to be drifting through the cruise until the last cast checked for which there appears to be a shift back towards earlier values. The temperature differences are a little higher than usual. Many casts were examined to see if there was a sudden change and it is not obvious that there was. There is no obvious drift during the first half of the cruise, then there does seem to be a fairly steady change, but event #37 has differences that are closer to event #19. Examining some of the noisiest differences does not indicate that one channel is responsible, just that there are sudden changes and slight misalignments lead to large differences. There was some rough weather during the cruise which might be thought to account for some of the noise; however, events #3 and #37 had steady CTD descent rates but very noisy differences. Moreover, the differences also vary when the CTD is at the bottom or during stops for bottles, with temperature differences are ~-0.001 for events #3 and #18 at 250m, ~0 for cast #31 and ~+0.001 for event #37 (very noisy).

On the positive side, in quiet sections the salinity differences are mostly within 0.001 which suggests there has been little calibration drift.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

There are no station names or water depths in the headers so a spreadsheet was prepared with files names, station names and water depths as found in the log book; this was used with routine “MERGE: CSV file to Headers”.

##### Checking Headers

A cross-reference list was checked against the log book and the only problem noted was at the last station, where the log longitude entry differs significantly from the header. The header entry corresponds to the position for this station on a previous cruise and is close to the BO and EN entries for that cast. It is only the BE entry that is so different. It appears that the header is correct, not the log entry.

A header check was run. No problems were detected.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 0.7db which is lower than usual. Since its last factory visit, the pressure sensor has only been used for a few casts during 2014-19. There was a note in the log of a deck pressure of -1.25. There are a number of casts with negative pressures that are associated with “in-water” salinity with the pumps on, and the transmissivity values look reasonable for near the surface. In most cases the minimum pressures in the files were ~-0.4db but for one cast there are some readings that appear to be in the water with pressure of -1.5db at both the beginning and end of the file but it is possible the CTD came out of water briefly as conditions were bad. Clearly, the pressure is reading too low, but it is not clear how low; the values do vary from cast to cast. Cast #27 has a bottle stop at -0.9db. So we should at least add 1db to the pressure and adding 1.25db to the offset looks reasonable.

Readings at the bottom were checked by comparing the maximum pressure (converted to depth) plus the altimetry reading at the time of the maximum pressure with the bottom depth recorded in the log. Unfortunately the two are not from the same time and the altimetry can be noisy, so this is not a reliable comparison. Nonetheless, 11 casts were checked and the differences varied from the CTD pressure reading looking high by 2.3m to looking low by 2.3m, with the average showing the CTD pressure to be low by ~0.3db.  The difference did not appear to depend on depth or event number. This at least suggests that a simple offset is a reasonable correction.

The altimeter and water depth readings from the headers of the MRH and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for most casts and the headers look appropriate. The header entry was removed from the MRH file #24 since it was a surface only cast and the CTD never got near the bottom, and MRH file #26 for which the chains were clearly interfering with the altimetry. The entries for the SAM files were fine.

Casts #23 and 24 will not be processed further; these were surface casts for filming purposes only.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +30 records and +70 records and +50 looks like the best choice overall.

SHIFT was run on the pH:SBE channel using a setting of +50 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results were found with -0.6 records for the primary and -0.3 records for the secondary.

SHIFT was run on all casts using -0.6 records for the primary conductivity and -0.3 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Conductivity, pressure and dissolved oxygen sensors were all recalibrated in late 2013 or early 2014. They have been used for only a few casts during 2014-19 when there was a problem with the CTD system, but no indication that the problem was with the sensors. The 2 temperature sensors were recalibrated in late January and early February 2013. The primary temperature sensor was used for 5 cruises in 2013 and the salinity derived compared well with calibration samples. There is no record of the secondary temperature sensor being used since the last recalibration.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. There were an unusual number of excursions towards high temperature and low salinity, but most look like they are due to deeper mixing than usual. This cruise occurred during a very windy period, reaching a peak when the ship anchored for a while in the northern Strait of Georgia. Data from the central and southern Strait of Georgia were within the climatology so the excursions are unlikely to be due to instrumental problems. The 3-standard deviation climatology is too severe for this region, particularly for casts close to shore. There were also reports of anomalously high temperatures in the offshore areas in 2014 and those came closer to shore in the autumn.

Repeat Casts – The only repeat casts were at station 19 which had a full cast plus 2 very shallow casts done for a videographer, so a comparison is not useful.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Since the primary temperature and salinity channels seem a little less spikey, those channels were selected for archiving, and hence, editing. The secondary T-S plots show more unstable features.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. All files required some editing.

Profile plots were examined to see if the data were acceptable and problems were found at the bottom of cast #25. The pH and Fluorescence data look odd. The CTD hit bottom during this cast and the initial editing seems to have removed most of the bad data. CTDEDIT was rerun on that cast removing a little more data from the bottom of the cast. There was no bottle sampling for that cast.

##### Initial Recalibration

No salinity recalibration is needed.

CALIBRATE was run using file 2014-50-recal1.ccf to add 1.25db to the pressure and to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files:

CTD DO Corrected = CTD DO * 1.0281

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was -0.0025mL/L but the standard deviation is 0.02mL/L. (See file 2014-50-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.02mL/L and standard deviation of 0.06mL/L; the results vary with depth with the CTD reading slightly low below 200m. The CTD looking relatively high above 100m could be partly due to incomplete flushing of Niskin bottles. The rough nature of the comparison, the high variability and the noise in the sensor data do not support further recalibration. See 2014-50-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2014-50-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some unstable features, but that is normal for this region in near-shore and well-mixed casts. No problems were found.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal

and unedited except that some records were removed in editing

temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

There were many two-sided spikes in the data, but they largely

disappear in metre-averaging. The comparison with bottles looked

normal.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 30db

±0.3 mL/L from 30db to 100db

±0.12 mL/L from 1000db to 200db

±0.04 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2014-50-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were low ranging from 63% to 98%. The highest value was at the mouth of Juan de Fuca Strait where we would expect something ~100%. And the lowest values were in Haro Strait and eastern Juan de Fuca Strait where low saturation rates are normal due to vertical mixing; the DO vertical gradients were very low in this area, as usual. Similar conditions were also seen in the northern part of the Strait of Georgia where they are not so common, but operations were stopped briefly due to a storm before the ship sampled that area and chains were used for a few casts; vertical mixing due to high winds would lower the DO saturation. The last two sites sampled had the highest saturation rates for the Strait, likely due to stratification redeveloping, but they are still lower than usual. A few casts with low saturation rates were examined to see if the DO values near the surface were out of line with those from bottle samples from the upcast, and they were not. So these results are not indicative of problems with the DO calibration; rather, they reflect unusual conditions.

##### Final Bottle Files

During the downcast of event #31 there is a small section of water near 210db with lower transmissivity than above and below it. On the way up the section of odd data is much more extensive (~120db-240db) with transmissivity and dissolved oxygen much lower than during the downcast and fluorescence higher. (up to 0.4ug/L). Above 120db transmissivity and fluorescence are close to the downcast values, but the dissolved oxygen stays lower. After the cast biological fouling of the DO sensor pump inlet was discovered. The affected DO:SBE data were replaced with pad values in the bottle file. The other variables look reasonable during bottle stops, so were left in place.

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars (Notes from Log Book)

2. Stopped, returned to surface to align winch – new file started and full cast run.

6. Tripped bottle 1 on way down by mistake – change sample numbers to adjust.

7. Bottles 12&13 failed integrity test.

8. Bottles passed integrity check.

13. Cable damaged due to winch block slip – winches swapped.

23,24, Bottles tripped for filming – no sampling. Not to be archived.

25. Bottom contact maybe –no sign on rosette.

26. Weights on. Altimeter not working during cast – chains may be problem.

27. Station cast stopped – even with chains on, risk too high.

31. Oxygen sensor reading about 1ml/l lower than downcast from 250m and above. Oxygen sensor pump inlet blocked with obelia, bryozoan colonies, forams, wood…– syringes used to evacuate.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 1763 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 5013 | 27Feb2013 | Factory |  |  |
| Secondary Cond. | 3394 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 3Jan2014 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0691 | 29Dec2010 | Factory |  |  |
| SeaPoint Fluor. | 3642 | n/a |  |  |  |
| Pressure Sensor | 506 | 30Dec2013 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |