## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2014-51

Agency: MEAD

Location: West Coast Vancouver Island

Project: Night Pelagic Ecosystem Survey

Party Chief: Flostrand L.

Platform: W.E. Ricker

Date: 3 August 2014 – 18 August 2014

Processed by: Germaine Gatien

Date of Processing: 16 April 2015 – 17 April 2015

Number of original CTD HEX files: 19		Number of CTD files: 19 (1 without pumped channels)

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. Attached were an SBE 43 DO sensor (#1176) and a WET Labs ECO_AFL Fluorometer (#2215) on the primary pump and an unpumped SBE18 pH sensor (#0851 for event #1 and #852 for events #3-37).

The deck unit was a Seabird model 11.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log was missing a list of participants and the equipment list was incomplete. This proved to be a problem as there were notes in the log indicating that a transmissometer was cleaned, yet it was not listed in the configuration file or equipment list. It would have helped greatly if the equipment list could have been trusted, but things were missing from it that were mounted on the CTD. This led to some confusion and time wasted trying to ensure that no data had been lost. It appears that it was the fluorometer that was cleaned and there likely was no transmissometer present.

The pumps were not turned on for event #17 so only pressure, pH and fluorescence were kept in the file.

The primary salinity was recalibrated by adding 0.001 based on an assumption that half the drift implied by the post-cruise calibration had occurred at the time of this cruise. There were only 2 salinity samples and they waited a long time for analysis so are likely to have been affected by evaporation.

The fluorometry traces show little spiking so the sampling rate was presumed to have been set low. This rate is adjusted within the sensor itself and is set before the cruise. Normally we have a lot more detail in this channel and use a filter to reduce it. No filter was applied to these data.

There was dissolved oxygen sampling for only the 1st of the 8 cruises that have used this sensor between factory visits. The results of cruise 2014-01 were used to recalibrate DO. The post-cruise factory calibration suggests a drift that is reasonably close to the correction applied.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained. It had no list of participants and the equipment list was incomplete. This proved to be a problem as there were notes in the log indicating that a transmissometer was cleaned, yet it was not listed in the configuration file or equipment list. It would have helped greatly if the equipment list could have been trusted but things were missing from it that were mounted on the CTD. This led to some confusion and time wasted trying to ensure that no data had been lost. It appears that it was the fluorometer that was cleaned and there likely was no transmissometer present.

There were two configurations as the pH sensor was changed after event #1. They were saved as 2014-51-ctd1.xmlcon and 2014-51-ctd2.xmlcon

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity, fluorometer and DO sensors were obtained.

The salinity analysis spreadsheet was obtained; there were 2 samples taken at ~500m.

The calibration constants were checked for all instruments. The only change made was to the dissolved oxygen factor E (from 3.6 to 3.75) to reflect tests run using 2014-01 data to fine-tune the hysteresis setting. This will not affect this cruise since the hysteresis correction will not be needed, but it was changed to avoid confusion in the future.

##### Conversion of Full Files from Raw Data

All files were converted using con file 2014-51-ctd1.xmlcon or 2014-51-ctd2.xmlson, with hysteresis correction turned off.

A few casts were examined and all expected channels are present.

The two temperature and conductivity channels are fairly close to each other during the downcasts but are noisier in the upcasts.

The fluorometer, pH and dissolved oxygen traces look normal.

The descent rate is fairly noisy but kept high (and not TOO high).

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A setting of +4s was found appropriate to advance the DO voltage for 4 other Ricker cruises of 2014, both before and after this one. Checks of the offset between distinctive features in T and DO suggest that +4s is right for this cruise too.

ALIGNCTD was used to advance the DO Voltage by 4s relative to the pressure.

Plots were examined before and after that step and the results with +4s look reasonable.

##### CELLTM

A setting of (α = 0.03, β=9) was found appropriate for the thermal mass correction for 4 other Ricker cruises using these sensors in 2014, both before and after this one.

CELLTM was run using (α = 0.03, β=9) for both the primary and secondary conductivity.

Tests were run on a few casts but were hard to interpret because the upcast and downcast traces were  often very different, but in areas where the profiles were similar CELLTM did improve the correspondence on a T-S curve.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

When these sensors were used during 2014-01 there was so much noise in the data that no reasonable comparison could be made between channels. For cruises 2014-16, 2014-17, 2016-36, 2014-66 and this cruise, the casts are all fairly shallow, making comparison less reliable than if we were looking at very low gradient zones as found in deeper water. The deepest sampling was about 250m for most of the other cruises, so some casts from each of the other cruises were compared with 2 from this cruise at 240db. The differences were:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-16-0224 | 240 | +0.0003 | +0.00013 | +0.0011 | Fairly steady |
| 2014-16-0087 | 240 | +0.0004 | +0.0002 | +0.0017 | Fairly steady |
| 2014-17-0155 | 240 | -0.0001 | +0.0004 | +0.0006 |  |
| 2014-17-0170 | 240 | -0.0002 | -0.00005 | -0.0004 |  |
| 2014-36-0016 | 240 | -0.0003 | -0.00003 | ~0 noisy | Fairly steady |
| 2014-36-0141 | 240 | -0.0002 | -0.0001 | -0.0013 | Fairly steady |
| 2014-36-0174 | 240 | -0.0002 | -0.0002 | -0.0017 N | Fairly steady |
| 2014-51-0003 | 240 | -0.0003 | -0.00058 | -0.0061 | Noisy,moderate |
| 2014-51-0013 | 240 | -0.0006 | -0.00057 | -0.0060 | Noisy,moderate |
| 2014-66-0045 | 240 | ~0 N | -0.00043 | -0.0044 | Fairly steady |
| 2014-66-0159 | 240 | -0.0001 | -0.00058 | -0.0062 | Fairly steady |
| 2014-66-0181 | 240 | ~0 | -0.00064 | -0.0072 | Fairly noisy |

The differences from this cruise and 2014-66 look quite different from those of the earlier cruises. There is some suggestion that the differences between the conductivity and salinity are growing. It will be seen later that post-cruise calibrations show that drift was smaller than this study shows. There may be differences in the conductivity channels while moving that are due to alignment variations rather than calibration drift.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. As is typical of this type of fluorometer, there were some negative values, so Fluorescence values <0ug/L were replaced with pad values.

##### Checking Headers

The header check was run and no problems were found.

The surface check gave an average value of 3.3db which is slightly low for the Ricker but reasonable. There were a few slightly negative pressures, but they were isolated and seen only at the end of casts and are associated with very low conductivity and occur  after the pumps were turned off.

The cross-reference check was compared with the log book and no errors were found.

The cruise track was plotted and added to the end of this report. No problems were found.

##### Shift - Alignment

Fluorescence

A shift of +12 is the usual setting used to align the Wetlabs ECO fluorescence with temperature and it was used for 2014-36 and 2014-66 when the same sensor was used. Tests were run on a few casts and the results look generally satisfactory.

SHIFT was run on all casts to advance the ECO fluorometer by +12 records.

pH

Because of hysteresis in the pH signal, tests are best run on casts with distinctive features in the pH traces. This was done by finding the distance between such features in the upcast and downcast and comparing that with the temperature offset after applying a variety of shifts. Settings between +50 and +60 records produced the best results overall during 2014-36 with the same equipment, and that setting looks good for these data as well.

SHIFT was run on all casts with the setting +55 records.

Conductivity

For the 5 other 2014 Ricker cruises using these sensors that have been processed, the best results were found using -0.7records for the primary conductivity. For the secondary -0.9 was best for the first 3 cruises but -1.3records worked better for 2014-66. Tests were run to see if these settings worked well for the current cruise and the 2014-66 settings did so.

Two runs of SHIFT were used to apply advance the primary by -0.7 records and the secondary by -1.3 records.

Dissolve Oxygen

Checks were made of the alignment of DO and temperature to see if the shift applied earlier (§5) had done a good job; it worked very well for some features but was too little or too much for others. The earlier adjustment appears to be the best we can achieve.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Comparisons of CTD and bottle salinity samples

There were only 2 salinity samples taken. They were analyzed more than 2 months after collection so were given quality flags of 4 by the analyst since the liners used have been found to be prone to evaporation leading to bottle salinity values that are too high.

The samples are said to be from 500m and 490m. Unfortunately, it is not clear whether these are really depths or might they be pressures. It is also not recorded how far above the CTD the Niskin was mounted. For High Seas Salmon cruises the Niskin is usually 5m above the CTD. An examination of the pressure versus scan number does not indicate that the CTD was stopped anywhere other than at the bottom which was at about 512db for cast #7. For cast #13 the CTD rose during the stop at the bottom from 495db to 487db. If we assume that the Niskin was mounted 5m above those levels the comparison suggests:

| Event | Pressure | Sal0 | Sal1 | Salinity:Bottle | Sal0-bot | Sal1-bot |
| --- | --- | --- | --- | --- | --- | --- |
| 7 | 507 | 34.0752 | 34.0699 | 34.0857 | -0.0105 | -0.0158 |
| 13 | 495 | 34.064 | 34.0585 | 34.0575 | 0.0065 | 0.0010 |
| 13 | 487 | 34.0494 | 34.0497 | 34.0575 | -0.0081 | -0.0078 |

Event #7: The vertical salinity gradient at the bottom of event #7 was low. Both sensors are reading lower than the bottle by at least 0.01.

Event #13: The vertical salinity gradient is higher and the CTD rose during the stop at the bottom. It is uncertain at what depth the bottle was closed, even if we have made the correct assumption about how far above the CTD the Niskin was mounted. So, estimates were based on the bottle closing when the CTD was at 500db or 492db. The latter is more likely since the bottle is usually fired after a wait, so most likely the sensors are reading lower than the bottle for this event as well, but not by as much as for the first case. If the bottle was fired soon after the stop the CTD salinity channels both appear to be higher than the bottles.

There is too much uncertainty in this study to conclude anything about the calibration. We do have a factory calibration after cruise 2014-66 which likely sets an upper limit of drift. The primary salinity was likely low by about 0.002 and the secondary low by 0.004 after 2014-66.

- Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity and temperature sensors were both recalibrated in late January 2013 and they were used for 8 cruises before the next factory visit. The comparisons with bottles, where there was sampling, were not fully trusted. The best sampling was from 2014-01 when the primary salinity was found to be low by about 0.004, but even for that cruise there was some concern that there might be some evaporation of samples and some small effect from poor flushing especially in light of that being the first use of the sensors since the previous factory visit. The primary sensors were only used for one cast so were recalibrated to match the secondary which was not corrected.

Dissolved Oxygen:

The DO sensor was used for 8 cruises between factory calibrations, but only 2014-01 included DO sampling, so the fit from that cruise was used for all the others (slope/offset = 1.0502/+0.0142). During 2014-01 data hysteresis tests were done and factor E was fine-tuned.

Pressure:

The pressure sensor was recalibrated in May 2013 and the factory offset was found appropriate for the 4 cruises since then that have been processed.

Historic ranges – All temperature and salinity data fell within the local climatology except for the salinity from cast #17 for which the pumps were not operating.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

Post-cruise calibrations were obtained from March 2015 for the primary conductivity and December 2014 for the other temperature and conductivity sensors. They show that the primary and secondary temperature sensors were reading slightly low, by 0.0008C° and 0.0007°, respectively. The conductivity sensors had drifted so that the primary salinity was low by 0.0026 and the secondary low by 0.0046 salinity units. The combined effect produces primary and secondary salinity low by about 0.002 and 0.004psu. Since this was the 6th of 8 cruises between calibrations, it is assumed that half the drift occurred by the time of this cruise. So salinity is likely low by 0.001.

There are post-cruise calibrations for both pH sensors but the data provided are from after modifications to the sensors, so not appropriate for data collected earlier.

The pressure post-calibration shows the pressure may be slightly low by no more than 0.8db, but this calibration was done as part of some equipment modifications that might possibly affect the offset.

- DETAILED EDITING

The first issue is to decide which sensor pair to edit. For most casts there is little to distinguish between them, but the primary temperature and salinity have been selected for cruises 2014-36 and 2014-66 because they looked slightly less noisy, so the primary channels were selected for archiving for this cruise too.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records.

All casts required some editing except #17 for which the pumped channels will not be archived. Notes were entered in the headers about editing applied.

- Initial Recalibration

Using the post-cruise pressure calibration from December 2014 would produce pressure higher by about 0.8db, but it is possible it was affected by modifications to equipment that were carried out at that time. There were some very low salinity values associated with pressure of 0.5db suggesting the sensors were very close to the surface; 0.5db looks more reasonable than 1.3db there. No recalibration will be applied.

The comparison of 2 bottle samples with CTD salinity suggest that the CTD sensors were reading low by something like 0.01 but there were many uncertainties in the comparison, the most serious being that the samples waited more than 2 months to be analyzed and had a type of seal that renders samples prone to evaporation.  The post-cruise calibration indicates that the primary salinity was low by about 0.002 in March 2015. This was the 6th of 8 cruises between factory visits, so it is reasonable to estimate that half the drift had occurred by the time of this cruise. The primary salinity will be recalibrated by adding 0.001.

The only information available about the SBE dissolved oxygen sensor since it was last recalibrated comes from 2014-01, so the correction applied to that data set will also be used for this cruise:

DOX_CTD Corrected = DOX_BOT = 1.0502* DOX_CTD +0.0142

File 2014-51-recal1.ccf was prepared to apply the above correction to Oxygen:Dissolved:SBE and to add 0.001 to the primary salinity. CALIBRATE was run on the EDT files.

##### Fluorometer Processing

At this stage a median filter, fixed size=11, is usually applied to reduce spikiness in fluorescence channels, but examination of all casts shows little spikiness, so this step was skipped. It looks like the sampling rate was set low on the sensor just as it was for 2014-66.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except #17 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag

For event #17 the following channels were removed:

Scan_Number, Temperature:Primary, Temperature:Secondary, Salianity:T0:C0, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Status:Pump, Descent_Rate and Flag

Profile plots were produced at this point to check for errors. No problems were found.

T-S plots were produced; there are some small unstable features that likely reflect real conditions.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

----------------------------------------------------------------------------------

Data Processing Notes:

----------------------

Fluorescence and pH data are nominal and unedited except that some

records were removed in editing temperature and salinity.

Dissolved oxygen was calibrated using the results of cruise 2014-01 when the

same sensor was used. There was no calibration sampling from this cruise.

Recalibration of the salinity channel was based on a report from the factory on

drift in the temperature and conductivity sensors as of 23 December 2014 and

3 March 2015, respectively. That report shows the primary salinity had drifted

low by about 0.002psu. This was the 6th of 8 cruises between the factory

checks so it was estimated that about the CTD salinity was reading low by

about 0.001psu.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were collected.  Calibration is required for each cast to get

absolute values, although general trends within a cast are likely real.

For details on the processing see the report: 2014-51_Processing_Report.doc.

----------------------------------------------------------------------------------

A separate note was added to the header of cast #17 to explain why the pumped channels were removed.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

The track plot looks fine.

The sensor history files were updated.

##### Dissolved Oxygen Surface Saturation

Dissolved Oxygen saturation was derived and plotted. The surface saturations varied from about 70% to 160% with a median value of ~115%. These values look reasonable but we have no DO sampling to confirm them. Note that during 2014-66 saturation rates were low overall using the same sensor.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | No | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 31Jan2013 | Factory | Mar2015 | Factory |
| Conductivity | 2128 | 29Jan2013 | Factory | Dec2014 | Factory |
| Secondary Temp. | 2374 | 31Jan2013 | Factory | Dec2014 | Factory |
| Secondary Cond. | 3396 | 30Jan2013 | Factory | Dec2014 | Factory |
| SBE 43 DO sensor | 1176 | 5Mar2013 | Factory | 23dec14 | Factory |
| WetLabs ECO-AFL/FL | 2215 |  |  |  |  |
| SBE18 pH | 0851 | 29May2012 | Factory | 4Sept14 | Factory |
| SBE18 pH | 0852 | 21Apr2011 | Factory | 31dec14 | Factory |
| Pressure Sensor | 75636 | 25Mar2013 | Factory | 17Dec14 | factory |